## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | VENTILATOR TRANSPORT WITH MONITOR
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | INTUBATION DIFFICULT VIDEO ADULT
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | REGULATOR SUCTION HIGH WALL MOUNT
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | ACT MACHINE
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | VENTILATOR ICU ADULT & PEDIATRIC
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | CYLINDER O2 SIZE E WITH REGULATOR
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | PAPR RESPIRATOR COMPLETE
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | BED INTENSIVE CARE
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | PUMP FEEDING
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | VENTILATOR TRANSPORT ADULT & PEDIATRIC
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | VENTILATOR BIPAP ICU
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | PUMP INFUSION
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | MANOMETER CUFF ETT
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | CYLINDER OXYGEN SIZE K/H WITH FLOW METER
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | VENTILATOR BIPAP IN-PATIENT
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | STAND IV

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | MIC10049 | VENTILATOR TRANSPORT WITH MONITOR |  | Medical Eqpt | 300 | [10 empty] | 0
2 | MIC10009 | INTUBATION DIFFICULT VIDEO ADULT |  | Medical Eqpt | 300 | [10 empty] | 0
3 | MGE20020 | REGULATOR SUCTION HIGH WALL MOUNT |  | Medical Eqpt | 4000 | [10 empty] | 0
4 | MOR20001 | ACT MACHINE |  | Medical Eqpt | 20 | [10 empty] | 0
5 | MGE20014 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | Medical Eqpt | 1000 | [10 empty] | 0
6 | MIC10046 | VENTILATOR ICU ADULT & PEDIATRIC |  | Medical Eqpt | 2115 | [10 empty] | 0
7 | MGE20007 | CYLINDER O2 SIZE E WITH REGULATOR |  | Medical Eqpt | 2000 | [10 empty] | 0
8 | MIF50016 | PAPR RESPIRATOR COMPLETE |  | Medical Eqpt | 30000 | [10 empty] | 0
9 | MWD10006 | BED INTENSIVE CARE |  | Medical Eqpt | 1500 | [10 empty] | 0
10 | MGE50026 | PUMP FEEDING |  | Medical Eqpt | 8745 | [10 empty] | 0
11 | MAL20015 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | Medical Eqpt | 300 | [10 empty] | 0
12 | MIC10043 | VENTILATOR BIPAP ICU |  | Medical Eqpt | 2000 | [10 empty] | 0
13 | MGE50027 | PUMP INFUSION |  | Medical Eqpt | 11405 | [10 empty] | 0
14 | MIC10025 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH |  | Medical Eqpt | 2915 | [10 empty] | 0
15 | SAR10803 | MANOMETER CUFF ETT |  | Medical Eqpt | 2000 | [10 empty] | 0
16 | MGE20022 | CYLINDER OXYGEN SIZE K/H WITH FLOW METER |  | Medical Eqpt | 500 | [10 empty] | 0
17 | MWD10034 | VENTILATOR BIPAP IN-PATIENT |  | Medical Eqpt | 2000 | [10 empty] | 0
18 | MGE40004 | STAND IV |  | Medical Eqpt | 16000 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | MIC10049 | 0 | VENTILATOR TRANSPORT WITH MONITOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MONITOR INCLUDED | YES : FOR PATIENT VITAL PARAMETERS
2 | PARAMETERS | ECG , RESP , NIBP , 2X IBP , TEMP , SPO2
3 | SPO2 | MASIMO OR NELLCOR OXI-MAX ( MUST COMPLY WITH SECTOR SUPPLIES STANDARD )
4 | TEMP | YES
5 | WAVEFORM | 3 - 5 WAVFORMS
6 | SCREEN SIZE | >=12"
7 | MONITOR TYPE | VENTILATION PARAMETERS AND PATIENT VITAL SIGNS DISPLAYED ON SAME MONITOR
8 | VENTILATOR
9 | CONFIGURATION | ADULT, PEDIATRIC , NEONATAL
10 | TIDAL VOLUME, L | 2-2000 ML
11 | BREATH RATE, BR/MIN | 0 - 80
12 | INSPIRATORY TIME, SEC | 0.2 - 3 SEC
13 | INSPIRATORY FLOW, L/MIN | 3 - 100 L/MIN, SPECIFY
14 | PRESSURE LEVEL, CM H2O | SPECIFY
15 | PRESSURE RAMP | SPECIFY
16 | PRESSURE SUPPORT | 0 - 50
17 | TRIGGER MECHANISM | FLOW OR PRESSURE OR BOTH, SPECIFY.
18 | SENSITIVITY, CM H2O | FLOW OR PRESSURE TRIGGER, SPECIFY
19 | FIO2, PERCENTAGE | 21-100 %
20 | I:E RATIO | 1:9 TO 4:1
21 | ADJUSTABLE PEEP | 0 - 40 CM H2O/ HIGHER
22 | OPERATING MODES
23 | ASSIST / CONTROL MODE | YES
24 | VOLUME MODE | YES
25 | PRESSURE MODE | YES
26 | SPONTANEOUS / CPAP MODE | YES
27 | SIMV MODE | YES
28 | MONITOR PARAMETERS
29 | PRESSURE, PA , MAP, PEEP. | YES
30 | VOLUME, TIDAL, MINUTE | YES
31 | INSPIRATORY / EXPIRATORY TIME | YES
32 | I:E RATIO | YES
33 | RESPIRATORY RATE | YES
34 | PATIENT ALARMS
35 | FIO2, LOW/HIGH MINUTE VOLUME, LOW INSPIRATORY PRESSURE, HIGH PRESSURE, LOSS OF PEEP, APNEA, INVERSE IE, HIGH CONTINUOUS PRESSURE OCCLUSION, HIGH RESPIRATORY RATE, OTHERS. | ALL
36 | EQUIPMENT ALARMS
37 | GAS-SUPPLY FAILURE, POWER FAILURE, VENT INOPERATIVE, LOW BATTERY, SELF DIAGNOSTIC, OTHERS. | ALL
38 | SILENCE | 2 MINUTES
39 | CONTROL TYPES | DIAL, SOFT KEYS, TOGGLE SWITCH
40 | BREATHING CIRCUIT ADULT & PEDIATRIC REUSABLE | 3 SET ( OPTION )
41 | AIR COMP /TURBINE AS AIR BACKUP SOURCE | YES, SPECIFY
42 | BATTERY WITH CHARGER | BUILT-IN, OPERATIONAL HOURS ON BATTERY >= 6HRS
43 | WEIGHT OF ( VENTILATOR + MONITOR ) | <=12KG
44 | OTHER SPECIFICATION | FDA, CE, ISO APPROVED.
45 | PLUG TYPE | BRITISH
46 | MOBILE CART | REQUIRED , 4-5 CASTORS WITH BASKET
47 | CERTIFICATIONS | APPROVED REGULATORY BODY
48 | CIRCUIT | 100 EACH INCLUDED WITH THE UNIT
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | MIC10009 | 0 | INTUBATION DIFFICULT VIDEO ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | VIDEO LARYNGOSCOPE FOR DIFFICULT AIRWAY INTUBATION
2 | DESIGN | LIGHT WEIGHT & PORTABLE
3 | THE SYSTEM SHALL BE BASED ON THE LATEST TECHNOLOGICAL ADVANCES IN DIFFICULT INTUBATIONS | YES
4 | THE UNIT SHALL BE DESIGNED TO PROVIDE A CLEAR VIEW OF THE VOCAL CORDS DURING INTUBATION | YES
5 | THE UNIT SHALL SIMPLIFY INTUBATION OF THE DIFFICULT AIRWAY | YES
6 | TRAUMA AIRWAY | YES
7 | RANGE IN BLADE SIZES ALLOW COVERAGE PATIENT WEIGHTS | YES
8 | IT SHALL HAVE 60° ANGLED BLADES | YES
9 | IT SHOULD HAVE PRE-SHAPED STYLET FOR QUICK INTUBATION PROCESS | YES
10 | IT SHOULD UTILIZE MINIMUM FORCE FOR INTUBATION | YES
11 | IT SHOULD HAVE ANTI-FOG FEATURE | YES
12 | THE UNIT SHALL HAVE INTEGRATED HIGH RESOLUTION CAMERA AT GOOD POSITION ON BLADE TO PROTECT FROM BLOOD AND SECRETION | YES
13 | THE UNIT SHALL MOUNT ON MOBILE CARTS AND CASTORS | YES
14 | INTUBATION OF CERVICAL SPINE IMMOBILIZATION | YES
15 | DIFFICULT AIRWAY MANAGEMENT AND ROUTINE INTUBATION | YES
16 | SCREEN SIZE NOT LESS THAN 7 INCH | LCD OR TFT
17 | BLADES SIZE
18 | SIZE 3 REUSABLE | QTY:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
19 | SIZE 4 REUSABLE | QTY:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
20 | MADE OF MEDICAL GRADE PLASTIC OR STAINLESS STEEL FOR LESS TRAUMA HANDLINGO TEETH AND SOFT TISSUE AND EASY | SPECIFY
21 | BUILT IN LED LIGHT FOR ILLUMINATION | YES
22 | RECHARGABLE BATTERY ON SYSTEM | YES
23 | STYLET | QTY:3
24 | ACCESSORIES | ANY ACCESSORIES, OPTIONS AND CONSUMABLE ITEMS NECESSARY TO OPERATE THE OFFERED SYSTEM(S) MUST BE CLEARLY IDENTIFIED AND PRICED SEPARATELY
25 | SIZE 3 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 120 ITEMIZED AND MUST BE OPTIONAL PRICE
26 | SIZE 4 SINGLE USE SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 120 ITEMIZED AND MUST BE OPTIONAL PRICE
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | MGE20020 | 0 | REGULATOR SUCTION HIGH WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE(S) OF SUCTION | ALL COMMON PROCEDURES AND GI
2 | MOUNTING | WALL MOUNTED
3 | FUNCTION | CONTINUOUS AND INTERMITTENT
4 | ADJUSTABLE CYCLE | INDEPENDENT OFF/ON
5 | INCREMENTS | 3 - 30 SEC
6 | VACUUM ADJUST LOCK | YES
7 | VACUUM GAUGE | DIGITAL OR ANALOG
8 | PIN STOP | YES
9 | RANGE, MM HG | 0 - 300, FULL LINE
10 | GRADUATIONS, MM HG | FIVE
11 | DISINFECTION | SPECIFY
12 | STERILIZATION | SPECIFY
13 | SAFTY TRAPE | YES
14 | PROBE AS PER SITE | YES
15 | COLLECTION BOTTLE | DISPOSABLE SYSTEM
16 | MATERIAL OF BOTTLE | POLY CARBONATE
17 | OTHER SPECIFICATIONS | COLOR-CODED GAUGE
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | MOR20001 | 0 | ACT MACHINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEAREANCE | YES, (PROVIDE COPY WITH OFFER)
2 | CONFIGURATION | HAND HELD
3 | DETECTION PRINCIPLE | PHOTO METRIC AND OPTICAL
4 | APPLICATIONS | POINT OF CARE & CRITICAL AREAS ( ER , ICU , CCU , OR , WARD )
5 | TESTS AVAILABLE | ACT , ACT LR , PT-INR , APTT , CT , CT APTT CUVETTE
6 | TEST CAPABILITY | WHOLE BLOOD FINGER BREAK
7 | AMOUNT, ML | APPROX. 0.5
8 | THROUGHPUT RESULT/HR. | 17-49 TEST PER CYCLE
9 | DAILY QC | AUTO
10 | DISPLAY | LCD
11 | RESULTS DISPLAYED | TEST RESULT IN SEC , OPD , DATE , TIME
12 | ALARM RANGE | CARTRIDGE EXPIRED , QC REQUIRED , QC FAILED , BATTERY STATUS
13 | RECHARGEABLE BATTERY | YES LITHIUM ION
14 | SPLASH PROOF | YES
15 | PRINTER | YES
16 | COMPUTER INTERFACE | RS 323 SERIAL PORT AND NETWORK-ENABLED CAPABILITY
17 | BARCOADING | YES
18 | LIQUID QC | 2 TYPES ( HI RANGE & LOW RANGE
19 | POWER SUPPLY
20 | VOLTAGE | 220 VOLT
21 | FREQUENCY | 60 ZHZ.
22 | POWER CORD | BRITISH
23 | REGULATORY COMPLIANCES | MDMA SFDA
24 | OTHER SPECIFICATION | REAGENT COMPATIBLE WITH SECTOR STANDARD
25 | NUMBER OF TEST | 100 BOX ( 45 CUVETTE PER BOX ) TO GET FOC DEVICE
26
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | MGE20014 | 0 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FLOW METER | WALL MOUNTED
2 | FLOW RANGE | 0 - 15 L / MIN
3 | COLOR CODED | SPECIFY
4 | PRESSURE COMPENSATED | YES
5 | GLASS ENVELOP | INCLUDED WITH INDICATOR BALL
6 | METAL NIPPLE | YES
7 | PROBE | AS PER SITE
8
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | MIC10046 | 0 | VENTILATOR ICU ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SYSTEM SELF CHECK | YES
2 | CONTROLLED / ASSIST MODES | PRESSURE, VOLUME & VOLUME GUARANTEED
3 | INTERMITTENT MODES | SIMV-VC, SIMV-PC , SIMV-VG
4 | TWO LEVEL OF PRESSURE MODE | BIPAP INVASIVE OR EQUIVELENT BIPAP , NIV-ST ( WITH MASK )
5 | SPONTANEOUS MODE | CPAP/PS MODE
6 | NIV , NIV-ST | REQUIRED
7 | MODE SETTINGS & CONTROLS
8 | TIDAL VOLUME | 50 - 2000 ML
9 | RESPIRATORY RATE | 3 - 80 B/M
10 | FIO2 | 21% - 100%
11 | INSPIRATORY PRESSURE | 1 - 60 CMH2O
12 | PEEP | 0 - 40 CMH2O
13 | PRESSURE SUPPORT | 0 - 60 CMH2O
14 | I:E RATION | 1:9 TO 4:1
15 | INSPIRATORY TIME | 0.25 - 15 SEC
16 | FLOW TRIGGER | 1 - 9 L/M
17 | PRESSURE TRIGGER | -10 TO -0.25
18 | INSPIRATORY PAUSE | YES, SPECIFY
19 | PRESSURE SLOPE/RISE TIME | YES, SPECIFY
20 | BIAS FLOW | YES, SPECIFY
21 | TUBE COMPENSATION | YES
22 | LEAK COMPENSATION | YES
23 | TRIGGER COMPENSATION | YES
24 | APNEA BACKUP / RATE | YES
25 | MANUAL BREATH | YES
26 | INSPIRATORY HOLD | YES
27 | EXPIRATORY HOLD | YES
28 | AUTOPEEP MEASUREMENT | YES
29 | WEANING PARAMETERS (P0.1, NIF & RSBI) | YES, SPECIFY
30 | LUNG COMPLIANCE & AIRWAY RESISTANCE | YES
31 | MONITORED PARAMETERS
32 | PIP, PPLAT, MAP & PEEP | YES
33 | TV & MV (CONTROLLED & SPONTANEOUS) | YES
34 | RESPIRATORY RATE | YES
35 | FIO2 | YES
36 | CLINICAL ALARMS (HIGH/LOW)
37 | FIO2 | YES
38 | PEAK PRESSURE | YES
39 | TV & MV | YES
40 | RESPIRATORY RATE | YES
41 | FIO2 | YES
42 | PATIENT DISCONNECT | YES
43 | LEAK | YES
44 | TECHNICAL
45 | O2 SENSOR | BUILT IN
46 | ACTIVE EXPIRATORY VALVE (AUTOCLAVABLE IS PREFERRED) | YES
47 | ELECTRONIC MICRO PUMP NEBULIZER | OPTIONAL ITEMIZED PRICE
48 | DISPLAY & CONTROL | COLORED TOUCH SCREEN NOT LESS THAN 12 INCH
49 | PATIENT CIRCUIT SUPPORTED ARM | ADJUSTABLE
50 | BACKUP BATTERY | NOT LESS THAN 2 HRS
51 | MOBILE BASE/ CART WITH LOCKING CASTORS | YES
52 | CENTRAL GAS INLETS (O2 & AIR) | BRITISH STANDARD BOC
53 | ELECTRICITY 220V, 60 HZ OR AUTO. | YES
54 | AIR INLET FILTER | YES
55 | TECHNICAL ALARMS
56 | GAS SUPPLY OR POWER FAILURE, VENT INOPERATIVE & LOW BATTERY | SPECIFY
57 | DATA DISPLAYED | NUMERICAL, GRAPH/WAVEFORMS & LOOPS
58 | TRENDS | NOT LESS THAN 72 HRS FOR ALL PATIENT DATA, SPECIFY
59 | WEANING MODE | YES
60 | HUMEDIFIER WITH CAHMBER & HEATER WIRE : APPROVED STANDARD BY SECTOR | INCLUDED WITH UNIT
61 | MOBILE CART | INCLUDED WITH BASKET
62 | PLUG STANDARD | BRITISH
63 | LANGAUGE | ENGLISH LANGUAGE
64 | UNIT OF MEASUREMENT | INTERNATIONAL STANDARD SI
65 | APRV OR EQUIVALENT | INCLUDED
66 | ADDITIONAL MODES: PAV, ASV, PPS, NAVA & BILEVEL-VG OR EQUIVALENT & SMARTCARE WITH CAPNO | OPTIONAL
67 | P/V TOOL, LOW FLOW PV, SPIRODYNAMIC, STATIC PV LOOP, OR EQUIVALENT & OTHERS | OPTIONAL
68 | ETCO2, ETO2, VCO2, VO2, EE & RQ | OPTIONAL
69 | OPEN LUNG TOOL, LUNG RECRUITMENT, FRC WITH PEEP INVIEW OR EQUIVALENT & OTHERS | OPTIONAL
70 | BREATHING CIRCUIT | 100 WITH EACH UNIT , OPEN STANDARD
71 | DEVICE CERTIFICATION | APPROVED FROM REGULATORY BODY
72 | AIR COMPRESSOR INTEGRATED WITH THE DEVICE | ITEMIZED PRICE
73 | PROVIDE WHERE THE DEVICE & TECHNOLOGY WERE USED | PROVIDE REF HOSPITALS
74 | MODES OF VENTILATION
75 | BI-LEVEL MODE ( APRV ) | OPTIONAL ITEMIZED
76 | INVERSE RATIO MODE : IRV | OPTIONAL
77
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | MGE20007 | 0 | CYLINDER O2 SIZE E WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | OXYGEN CYLINDER E-SIZE | YES
2 | ALUMINUM CYLINDER | YES
3 | NOMINAL CONTENTS (LITRES) | 680L
4 | PIN INDEX | YES
5 | CAPACITY | SPECIFY
6 | ALUMINUM TROLLEY | INCLUDED
7 | OXYGEN REGULATOR | INCLUDED WITH FLOW METER 15 L/MIN
8 | OTHER SPECIFICATION | SPECIFY
| Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | MIF50016 | 0 | PAPR RESPIRATOR COMPLETE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PROVIDES HIGHER FILTRATION EFFICIENCY. | YES
2 | THE BLOWER MUST HAVE 2 OR 3 FILTRATION INLET. | YES
3 | THE BLOWER SHOULD BE LIGHT WEIGHT FOR EASY CARRY OR WEAR. | YES
4 | THE BLOWER SHOULD HAVE ADJUSTABLE WEST BELT. | YES
5 | THE BLOWER MUST HAVE DETECTOR FOR MEASUREMENT AIR FLOW BEFORE USE. | YES
6 | THE BLOWER SHOULD HAVE INTEGRATED BATTERY NOT LESS 4 HOUR AFTER FULL CHARGER. | YES
7 | THE DEVICE DISPLAY SCREEN MUST PROVIDE WARNING INDICATORS FOR THE SAFETY STATUS ALWAYS AND BE VISIBLE FOR AIRFLOW PERFORMANCE / PERFORMANCE FILTERS / BATTERY CHARGIN OR SIMILER WARMINGS. | YES
8 | PROVIDES HIGH AIRFLOW TO MAINTAIN POSITIVE PRESSURE. | YES, APPROX (150 - 200) LITERS PER MINUTE
9 | PROVIDES BUMP-CAP—LIKE HEAD PROTECTION; FULL-HELMET DESIGN PROVIDES MORE COMFORTABLE HEAD ROOM AND CONVENIENCE IN DONNING AND DOFFING. | YES
10 | REQUIRES LESS THAN A MINUTE TO DISASSEMBLE OR ASSEMBLE. | YES
11 | THE BLOWER MUST HAVE A LOWER NOISE-LEVEL SYSTEM COMPATIBLE WITH STETHOSCOPE USE. | YES
12 | THE SYSTEM SHOULD HAVE LONG LOOSE-FITTING HOOD W/CAPE, ( POLYETHYLENE BARRIER FILM), DOUBLE LAYER. | QTY : 6
13 | PAPR FILTER FOR BIOHAZARD | QTY : 12
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | MWD10006 | 0 | BED INTENSIVE CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | INTENSIVE CARE BED/ ICU ADULT
2 | PATIENT CONTROLS | YES
3 | TYPE | SPECIFY
4 | FUNCTIONS | HEAD HIGH/ LOW , KNEE UP / DOWN, TRENDELENBURG & REVERSE TRENDELENBURG, ELECTRICAL FLEXAFOOT
5 | NURSE CONTROLS | YES
6 | PATIENT CONTROL LOCKOUT | YES
7 | FULL LOW INDICATOR | YES
8 | CPR MODE | DUAL SIDED
9 | CPR CONTROL | YES
10 | GRAPHICAL CAREGIVER INTERFACE WITH TOUCH SCREEN FOR BED AND MATTRESS FUNCTIONS | YES
11 | AUTOMATIC CONTOUR | YES
12 | CHAIR MODE | FULL CHAIR & FULL CHAIR FRONT EGRESS POSITION
13 | ALARM SYSTEM | HEAD ANGLE ALARM 30 AND 45 DEGREES, BED EXIT ALARM, BRAKE OFF ALARM
14 | TRENDELENBURG GAUGE | YES
15 | BACKUP FUNCTION | ELECTRO-HYDRAULIC OR OTHER SPECIFY
16 | BED SURFACE TYPE | SPECIFY
17 | C-ARM APPLICATION | YES
18 | FRACTURE FRAME RECEPTACLES | YES
19 | ELECTRIC FOOT EXTENDER / RETRACTABLE | YES / ELECTRICAL
20 | DRAINAGE BAG HOLDER | YES
21 | MAX. PATIENT WEIGHT, KG | 230 KG OR MORE
22 | OVERALL DIMENSIONS
23 | L X W, CM | L210 235X W93 -102 CM APPROX.
24 | HEIGHT, CM | 45 CM – 92 CM APPROX
25 | SIDE RAILS LENGTH
26 | FRACTION OF OVERALL LENGTH | YES SPECIFY
27 | CASTERS
28 | SIZE (DIAMETER) CM | 12.5 CM OR MORE SPECIFY
29 | FUNCTION | SWIVEL, BRAKE, STEER
30 | CENTRAL BREAK SYSTEM | YES
31 | BUMPERS | CORNER/ FULL PERIMETER
32 | REMOVAL HEADBOARD & FOOTBOARD | YES
33 | IV POLE MOUNT | MINIMUM 3 PROVISIONS
34 | INTEGRATED BED SCALE WITH PATIENT WEIGHT STATISTICS FOR UP TO7 DAYS OR MORE | YES
35 | POWER SUPPLY | 220 VOLTS, 60 HZ
36 | BATTERY (BACK-UP) | OPTIONAL ITEMIZED
37 | DOUBLE INSULATION | YES
38 | ELECTRIC SHOCK PROTECTION | CLASS 1, TYPE B
39 | NUMBER OF MOTORS | 4 OR MORE SPECIFY
40 | GROUNDED | YES
41 | ISOLATED (MOTOR GROUND | YES
42 | NON-CONDUCTIVE SIDE RAILS | YES
43 | AIR MATTRESS | ANTI-BACTERIAL, ANTI-STATIC, MOISTURE PROOF, FIRE RESIST & CHEMICAL-RESIST, TOUGH & HIGH DURABLE
44 | INTEGRATED WITH THE BED | YES
45 | DIMENSIONS 20CM OR MORE HEIGHT | COMPATIBLE WITH BED DIMENSIONS
46 | CONTROLS ARE IMBEDDED WITHIN THE SIDE RAILS | YES
47 | AUTOMATIC DEFLATE | FULL INFLATION ONCE CPR MODE IS PRESSED
48 | MAX INFLATE FUNCTION | YES
49 | X-RAY SLEEVE INSIDE THE MATTRESS | YES
50 | OTHER SPECIFICATIONS | FDA APPROVED CE MARKED
51 | PRINTOUT FOR THE ITEM OFFERED WITH ALL TECHNICAL PARAMETERS EXACTLY MATCHES THE SOFT COPY IN THE CD. | YES
52 | SUPPLIER IS AN AUTHORIZED REPRESENTATIVE OF THE MANUFACTURER BY SFDA AND/OR AN OFFICIAL AGENT BY MINISTRY OF TRADE. | YES (ATTACH OFFICIAL DOCUMENTS)
53 | ATTACH MDMA CERTIFICATE | YES (GIVE THE MDMA NUMBER)
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | MGE50026 | 0 | PUMP FEEDING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | FOR LIQUID FEEDING
2 | CONFIGURATION | IV POLE ON STAND WITH 4-5 CASTORS
3 | PUMP TYPE | ROTARY PERISTALTIC OR LINEAR PERISTALTIC, SPECIFY
4 | OPERATING MODE | CONTINUOUS, VOLUME/TIME SELECTION AND BOLUS APPLICATION
5 | FLOW RANGE IN ML/HOUR | 0.5- 600 OR MORE
6 | DISPLAY | LCD
7 | ACCURACY | 5 - 10 %, SPECIFY
8 | OCCLUSION PRESSURE | 15 - 25 PSI, SPECIFY
9 | OCCLUSION ALARM | YES, AUDIBLE & VISUAL
10 | LOW BATTERY ALARM | YES, VISUAL
11 | FLOW ERROR ALARM | YES, AUDIBLE & VISUAL
12 | SET NOT FIXED ALARM | YES, AUDIBLE & VISUAL
13 | DOSE COMPLETE ALARM | YES, AUDIBLE & VISUAL
14 | PARAMETERS RETAINED IN MEMORY | RATE, DOSE & VOLUME INFUSED
15 | FUNCTIONAL BACK UP MEMORY | YES, AT LEAST 24 HOURS OR MORE
16 | FEEDING ADMINISTRATION SET | YES, OPEN SYSTEM
17 | SELF DIAGNOSTIC TEST | YES
18 | BUILT IN RECHARGEABLE BATTERY | YES
19 | POWER SUPPLY | 220 V, 60 HZ
20 | REGULATORY COMPLIANCE | FDA, CE APPROVAL
21 | OTHER SPECIFICATION | THE DEVICE MUST BE ABLE TO CONNECT WITH HIS SYSTEM
22 | PLUG | BRITISH
23 | LANGUAGE | ENGLISH
24 | NG FEEDING BAGS, 1000 ML, WITH TUBES COMPATIBLE WITH FEEDING PUMPS, CAP, SIDE PORT | QTY : 1000 ITEMIZED PRICE ALSO
25 | PROVIDE FREE PUMP WITH 1000 SETS FEEDING | MUST QUOTE
26
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | MAL20015 | 0 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PATIENT TYPE | ADULT, PEDIATRIC
2 | BUILT-IN TURBINE ( COMPRESSOR ) | YES
3 | CONTROLS
4 | MONITORING | TOTAL EXPIRATORY MINUTE VOLUME, SPONTANEOUS MINUTE VOLUME, PEAK PRESSURE, MEAN PRESSURE, PLATEAU PRESSURE, POSITIVE END EXPIRATORY PRESSURE, EXPIRATORY TIDAL VOLUME, INSPIRATION TIME, FREQUENCY, SPONTANEOUS BREATHING FREQUENCY, OXYGEN CONCENTRATION, VENTILATION TIME RATIO, GAS CONSUMPTION, BATTERY CAPACITY.
5 | TIDAL VOLUME, L | 50-2000 ML
6 | BREATH RATE, BR/MIN | 0 - 60 OR MORE
7 | INSPIRATORY TIME, SEC | 0.2-10 SEC
8 | INSPIRATORY FLOW, L/MIN | 60 UP TO 100
9 | PRESSURE LEVEL, CM H2O | 0-60 CM H2O
10 | PRESSURE RAMP | SPECIFY
11 | PRESSURE SUPPORT | 0 - 30 CM H20
12 | TRIGGER MECHANISM | FLOW OR PRESSURE OR BOTH, SPECIFY.
13 | SENSITIVITY, CM H2O | FLOW OR PRESSURE TRIGGER, SPECIFY
14 | FIO2, PERCENTAGE | 21 % - 100 %
15 | I:E RATIO | 1:4 TO 3:1
16 | ADJUSTABLE PEEP EITHER INTERNAL OR EXTERNAL, SPECIFY | 0-20 CM H2O
17 | OPERATING MODES
18 | INTERMITTENT POSITIVE PRESSURE VENTILATION (VC-CMV) | YES
19 | ASSISTED INTERMITTENT POSITIVE PRESSURE VENTILATION (VC-AC) | YES
20 | SYNCHRONIZED INTERMITTENT MANDATORY VENTILATION (VC-SIMV) | YES
21 | SPONTANEOUS VENTILATION (SPNCPAP) | YES
22 | PRESSURE SUPPORT TO ASSIST SPONTANEOUSLY BREATHING PATIENTS (SPNCPAP+PS) | YES
23 | MASK VENTILATION WITH LEAKAGE COMPENSATION IN THE MODES SPNCPAP (SPNCPAP+PS)(NIV) & BIPAP | YES
24 | DISPLAY TYPE | COLORED , BUILT-IN WITH THE UNIT
25 | MONITOR PARAMETERS
26 | PRESSURE, PIP, MAP, PEEP. | YES
27 | VOLUME, TIDAL, MINUTE | YES
28 | INSPIRATORY/EXPIRATORY TIME | YES
29 | I:E RATIO | YES
30 | RESPIRATORY RATE | YES
31 | PATIENT ALARMS
32 | FIO2, LOW/HIGH MINUTE VOLUME, LOW INSPIRATORY PRESSURE, HIGH PRESSURE, LOSS OF PEEP, APNEA, INVERSE IE, HIGH CONTINUOUS PRESSURE OCCLUSION, HIGH RESPIRATORY RATE, OTHERS. | SPECIFY
33 | EQUIPMENT ALARMS
34 | GAS-SUPPLY FAILURE, POWER FAILURE, VENT INOPERATIVE, LOW BATTERY, SELF DIAGNOSTIC, OTHERS. | SPECIFY
35 | SILENCE | 2 MINUTES
36 | CONTROL TYPES | DIAL, SOFT KEYS, TOGGLE SWITCH
37 | POWER, VAC | 220, 60HZ.
38 | BREATHING CIRCUIT ADULT & PEDIATRIC | 3 SET
39 | BUILT-IN BATTERY WITH CHARGER | 2 HOURS AT LEAST
40 | THE ABOVE UNIT IS COMPATIBLE WITH AMBULANCE CAR | YES
41 | ALL ACCESSORIES NEEDED ,INCLUDING PATIENT BREATHING CIRCUIT LOUPES & FILTERS WITH EACH VENTILATOR | YES
42 | CERTIFICATES | APPROVED FROM REGULATORY BODY
43 | LANGUAGE | ENGLISH
44 | PLUGS | BRITISH STANDARD
45 | UNIT OF MEASUREMENT | SI : INTERNATIONAL STANDARD
46 | MOBILE STAND | INCLUDED ON CASTORS WITH BASKET
47 | BED RAIL HANGER | INCLUDED
48 | HANDLE | YES
49 | WEIGHT | 7 KG APPROX
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | MIC10043 | 0 | VENTILATOR BIPAP ICU |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | THE UNIT MUST BE THE LATEST TECHNOLOGY IN NONINVASIVE VENTILATION. | YES
2 | THE UNIT MUST BE POSITIVE AIRWAY PRESSURE MACHINE TO BE USED IN HEALTHCARE FACILITIES. | YES
3 | THE UNIT MUST BE USED FOR ADULT AND PEDIATRIC PATIENTS. | YES
4 | THE UNIT MUST HAVE THE FOLLOWING CONTROLS: | YES
5 | INSPIRATORY TIME 0.5 TO 3 SEC. | YES
6 | INSPIRATORY PRESSURE 4 TO 30 CM H2O. | YES
7 | EXPIRATORY PRESSURE 4 TO 25 | YES
8 | RESPIRATORY RATE 5 TO 40 BPM. | YES
9 | THE UNIT MUST HAVE THE FOLLOWING MODES: | YES
10 | CPAP | YES
11 | BIPAP: SPONTANEOUS MODE, TIMED MODE AND SPONTANEOUS-TIMED MODE. | YES
12 | BIPAP WITH BACK UP RATE. | YES
13 | PRESSURE ASSIST CONTROL. | YES
14 | VOLUME ASSURED PRESSURE SUPPORT WITH BACKUP RATE | YES
15 | VOLUME ASSURED PRESSURE SUPPORT WITH AUTOMATIC EXPIRATORY PRESSURE ADJUSTMENT. | YES
16 | THE UNIT MUST HAVE DIFFERENT CONTROL FOR AN INSPIRATORY TIME, EXPIRATORY TIME, PIP, MAXIMUM PIP, RESPIRATORY RATE, CPAP AND EPAP, TRIGGER SENSITIVITY, AND SCREEN LOCK. | YES
17 | THE UNIT MUST BE EQUIPPED WITH AUTO PATIENT START AND STOP. | YES
18 | THE UNIT MUST HAVE AN ADJUSTABLE RAMP TIME WITH A RANGE OF 0 TO 45 MIN APPROXIMATELY. | YES
19 | THE UNIT SHOULD HAVE A BUILT-IN FILTER. | YES
20 | THE UNIT MUST HAVE LOW-PRESSURE OXYGEN INLET. | YES
21 | THE UNIT MUST HAVE LEAK COMPENSATION FOR AT LEAST 100 L/M. | YES
22 | THE UNIT MUST HAVE A BUILT-IN HEATED HUMIDIFIER WITH TEMPERATURE ADJUSTED. | YES
23 | THEY UNIT SHOULD INCLUDE HUMIDIFIER BLOCKER TO USE THE DEVICE WITHOUT HUMIDIFIER. | YES
24 | THE UNIT MUST HAVE A DIGITAL WAVEFORM DISPLAY AND WAVEFORM SOFTWARE. | YES
25 | THE UNIT SHOULD HAVE MEMORY PORT WITH MEMORY CARD INCLUDED. | YES
26 | THE UNIT SHOULD HAVE A CARRYING CASE. | YES
27 | THE UNIT MUST BE INSTALLED ON A MOVABLE STAND. | YES
28 | THE UNIT SHOULD BE EASY TO OPERATE. | YES
29 | THE UNIT SHOULD HAVE EVENT LOG. | YES
30 | THE CONTROLS, LABELS, AND DISPLAYS SHOULD BE CLEAR AND VISIBLE. | YES
31 | THE CONTROLS SHOULD BE PROTECTED AGAINST ACCIDENTAL SETTING CHANGES. | YES
32 | THE UNIT SHOULD BE SMALL IN SIZE AND LIGHTWEIGHT WITH APPROXIMATE WEIGHT LESS THAN 4KG WITH HUMIDIFIER. | YES
33 | THE UNIT SHOULD BE ABLE TO RESIST TIPPING OVER. | YES
34 | ALARMS SHOULD ALLOW QUICK ASSESSMENT AND CORRECTION OF THE ALARM CONDITION. | YES
35 | THE PRIORITY OF THE ALARM SHOULD BE INDICATED BY DIFFERENT AUDIBLE TONES AND VISUAL INDICATORS. | YES
36 | AUDIBLE ALARMS AND VISUAL INDICATORS MUST ACTIVATE WHEN SWITCHING FROM LINE TO BATTERY POWER. | YES
37 | THE UNIT MUST ALSO HAVE POWER FAILURE, OVERHEAT AND FAULT ALARMS. | YES
38 | THE FOLLOWING ALARMS SHOULD ALSO BE INCLUDED | YES
39 | LEAK ALARM | YES
40 | LOW AND HIGH PRESSURE | YES
41 | RESPIRATORY RATE | YES
42 | APNEA ALARM IS A MUST | YES
43 | BIPAP UNITS SHOULD BE ABLE TO OPERATE IN A VARIETY OF ENVIRONMENTS AND SHOULD NOT BE AFFECTED BY ELECTROMAGNETIC INTERFERENCE AND ELECTROSTATIC DISCHARGE. | YES
44 | SERVICING BY A SKILLED TECHNICIAN SHOULD BE EASY; THE OPERATOR’S MANUAL SHOULD PROVIDE ADEQUATE INFORMATION FOR CLINICIANS, USERS, AND CAREGIVERS. | YES
45 | THE UNIT SHOULD BE WELL CONSTRUCTED WITH DURABLE MATERIAL TO WITHSTAND TYPICAL ABUSE & CLEANING. | YES
46 | THE UNIT SHOULD BE FDA APPROVED AND CE MARKED. | YES
47 | THE UNIT SHOULD COMPLY WITH INTERNATIONAL RESPIRATORY STANDARDS. | YES
48 | THE UNIT SHOULD BE 220VAC AT 60HZ. | YES
49 | ALL AVAILABLE OPTIONS, CONSUMABLES, OR DISPOSABLES DEEMED NECESSARY FOR THE INTENDED FUNCTION OF THE EQUIPMENT AND NOT LISTED IN THE TECHNICAL SPECIFICATIONS SHOULD BE INCLUDED AND SPECIFIED. | PROVIDE ITEMIZED PRICE
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | MGE50027 | 0 | PUMP INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PROGRAMMING PATINET ID / NURSE ID ON THE PUMP DURING ADMINISTRATION | YES ( PLEASE SPECIFY )
2 | QUICK START OF INFUSION | YES
3 | FLOW RANGE ML/HR | 0.1 - 999 ML/HR
4 | ACCURACY | 0.05
5 | DRUG/DOSE CALC | YES
6 | VTBI, ML | 0.1 TO 9999
7 | MAX PRESSURE, PSI | SPECIFY MAX
8 | PUMPING MECHANISM | SPECIFY MECHANISM
9 | IV SET | YES
10 | FREE-FLOW PROTECTION | YES
11 | MANUAL AND PROGRAMMABLE BOLUS | PLEASE SPECIFY BOLUS RATE
12 | ALARMS & INDICATORS
13 | EMPTY BAG | YES
14 | FLOW ERROR | YES
15 | DOOR OPEN | YES
16 | CONTINUOUS PRESSURE MONITORING PSI (ON THE FRONT SCREEN ALL THE TIME) | SPECIFY
17 | DRUG DOSE CALCULATION | ML/H, µG/H, µG/MIN, µG/KG/H, µG/KG/MIN, MG/H, MG/MIN,MG/KG/MIN, MG/KG/H, IU/H, IU/MIN, IU/KG/H, IU/KG/MIN.
18 | FLEXIBLE PRESSURE ALARM SETTINGS | IN LINE PRESSER
19 | HIGH PRESSURE/OCCLUSION | YES
20 | PRE- OCCLUSIONALARM | YES
21 | CIRCUIT MALFUNCTION | YES
22 | INFUSION NEAR END | YES
23 | INFUSION COMPLETE | YES
24 | LOW BATTERY | YES
25 | DEPLETED BATTERY | YES
26 | USER PROMPT | YES
27 | AUDIBLE ALARM | YES
28 | VOLUME CONTROL | YES
29 | MOMENTARY SILENCE | YES
30 | PPM DUE DATA REMINDER | YES
31 | FLUID RESISTANT | YES
32 | NURSE CALL JACK | YES
33 | DATA LOG | YES
34 | EVENT LOG | YES
35 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED)
36 | NUMBER OF EVENTS | 1500
37 | PRINT OUT | YES
38 | DISPLAY | LED AND LCD WITH COLOR CABLITIES IS A PLUS & BIG SCREEN INTERFACE IS ALSO A PLUS
39 | INFUSED VOLUME DISPLAY | YES , DIGITAL
40 | WEIGHT WITH BATTERY KG | SPECIFY WIGHT
41 | TIME /DATE | YES
42 | POWER SOURCE, VAC | 220VOLT / 60 HZ.
43 | POWER SUPPLY | INTERNAL OR EXTERNAL (IF EXTERNAL MUST BE LIFETIME WARRANTY)
44 | POWER CABLE | DURABLE POWER CABLE
45 | BATTERY | YES, RECHARGEABLE BUILT IN
46 | BATTERY | NI-MH
47 | OPERTING TIME , HR | SPECIFY HOURS @ 10 ML/HR
48 | RECHARGE TIME , HR | PREFERRED LOWER CHARGING HOURS
49 | POLE MOUNTING | PREFERRED
50 | OTHER SPECIFICATIONS | FDA / CE
51 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED).
52 | STACKABLE IN THE SAME DOCKING STATION | INFUSION PUMPS / SYRINGE PUMPS / ENTRAL FEEDING PUMPS.
53 | DOSE ERROR REDUCTION SYSTEM (SMART TECHNOLOGY) | YES
54 | LIBRARY SIZE | YES CONFIGURABLE
55 | NO. OF DRUG ENTITIES/CARE AREA | YES PLEASE SPECIFY
56 | WIRELESS CONNECTIVITY | UPGRADABLE
57 | TIME RETAINED | FIRST IN FIRST OUT
58 | ACCESSORIES FOR THE PUMP | CLOSED SYSTEM
59 | ALL SETS HAVE SELF -SEALING NEEDLE FREE INJECTION SITES | YES
60 | ALL SETS SHOULD BE LATEX DEHP & PVC FREE | YES
61 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
62 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
63 | ACCESSORIES FOR INFUSION PUMP | CLOSE SYSTEM
64 | THE VENDOR MUST PROVIDE A FIXED PRICE FOR THE ALL ABOVE MENTIONED ACCESSORIES VALID FOR 3 YEARS | YES
65 | PUMP HL7 READY | YES ( PLEASE SPECIFY )
66 | PROVIDE CONSUMABLE DEAL ( PUMPS FREE AND PRICE ON SETS ) | OPTIONAL DEAL SECTOR MIGHT CHOOSE
67 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
68 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
69 | THE WORKSTATION SHOULD BE MOUNTABLE , WALL UNIT (FIXATION TO WALL RAILS OF DIFFERENT MANUFACTURER).MOVABLE POLE-BASED UNIT.
70 | SMART ALARM SYSTEM BY AUDIBLE, VISUAL AND DISPLAYED MESSAGE. | OCCLUSION, END OF INFUSION, EMPTY BATTERY, NO RATE SET, DEFECT ON PUMP. PRE-ALARMS: LOW BATTERY, NEAR-END OF INFUSION , PLEASE SPECIFY
71 | MOBILE STAND TO FIX THE PUMP ON INSTEAD OF PENDANT OR WALL FOR SPECIAL SITE NEEDS | ITEMIZED PRICE OPTIONAL
72 | PRIMARY INFUSION SET (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 120 SETS WITH EACH PUMP ITEMIZED
73 | LIGHT SENSITIVE PRIMARY INFUSION SET WITH NO INJECTION PORT FOR NEONATAL TPN (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 40 SETS WITH EACH PUMP ITEMIZED
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | MIC10025 | 0 | MONITOR PHYSIOLOGICAL BEDSIDE 15 INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | THE SYSTEM SHOULD BE BASED ON THE LATEST TECHNOLOGY IN MONITORS, PATIENT TRANSPORT THAT COVER ALL APPLICATIONS NEEDED IN ICU. | YES
2 | THE MONITOR SHOULD HAVE INTENSIVE CARE SOFTWARE FOR ADULT, PEDIATRIC. | YES
3 | THE MONITOR SHOULD PROVIDE A HIGH LEVEL OF ACCURACY OF MONITORED PARAMETERS. | YES
4 | THE MONITOR SHOULD BE MOBILE STAND WITH LOCKABLE WHEELS (MINIMUM OF 5), WITH BASKET, | YES
5 | THE MONITOR SHOULD PROVIDE A HIGH LEVEL OF ACCURACY (E.G. 2% APPROXIMATE FOR ECG) OF ALL MONITORED PARAMETERS. VENDORS SHOULD PROVIDE PARAMETERS ACCURACY DETAILS. | YES
6 | THE MONITOR SHOULD BE MODULAR WITH SINGLE OR MULTI-PARAMETERS PER MODULE. | YES
7 | THE DISPLAY SHOULD BE COLOR LCD AND SHOULD HAVE AN APPROXIMATE SIZE OF 15 INCHES WITH A HIGH LEVEL OF CONTRAST, BRIGHTNESS, RESOLUTION, AND VISIBILITY FROM A DISTANCE AND AT A WIDE ANGLE. MONITORED PARAMETERS VALUES SHOULD BE LARGE ENOUGH TO BE READ | YES
8 | FROM A DISTANCE. | YES
9 | THE MONITOR SHOULD BE A TOUCHSCREEN. | YES
10 | THE MONITOR SHOULD BE CAPABLE OF DISPLAYING SIX (6) PHYSIOLOGICAL WAVEFORMS SIMULTANEOUSLY AS A MINIMUM. | YES
11 | THE MONITOR SHOULD BE CAPABLE OF INTERFACING WITH THE HOSPITAL AND BEDSIDE CLINICAL INFORMATION SYSTEM (HCIS). IT SHOULD BE ABLE TO DOWNLOAD ANY RECORDED PARAMETER, WAVEFORMS AND NUMERICAL DATA, TO THE HCIS. THE VENDOR SHOULD SUBMIT A LIST OF | YES
12 | COMPATIBLE HCIS. (PLEASE PROVIDE IT AS ITEMIZED PRICE) | YES
13 | THE VENDOR SHOULD PROVIDE STATE OF THE ART TECHNOLOGY FOR INFORMATION TECHNOLOGY USED IN THE MONITOR APPLICATIONS. | YES
14 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL CLINICAL SOFTWARE THAT IS USED ON THE MONITOR (VENDOR SHOULD SUPPORT DOCUMENTS WITH EVIDENCE THAT PROVE THE BENEFIT OF THE CLINICAL PROGRAMS) | YES
15 | THE UNIT SHOULD BE WELL CONSTRUCTED WITH DURABLE MATERIALS TO WITHSTAND TYPICAL ABUSE, HEAVY USED ENVIRONMENT. THE VENDOR SHOULD PROVIDE SUPPORT DOCUMENT THAT PROVES THE DURABILITY OF HIS PRODUCT (INTERNATIONAL REFERENCE SITE SHOULD BE MENTIONED). | YES
16 | THE MONITOR SHOULD BE CAPABLE OF PROVIDING AN AUXILIARY ECG OUTPUT SIGNAL FOR DEFIBRILLATOR/IABP SYNCHRONIZATION, AND AN AUXILIARY BP OUTPUT SIGNAL FOR IABP SYNCHRONIZATION (PREFERABLE). | YES
17 | THE MONITOR SHOULD PROVIDE USER-ADJUSTABLE VISUAL AND AUDIBLE ALARMS TAILORED FOR INDIVIDUAL PARAMETERS. | YES
18 | THE MONITOR SHOULD BE CAPABLE OF PROVIDING GRAPHICAL AND NUMERICAL 24-HOURS TRENDS OF ALL MONITORED PARAMETERS. | YES
19 | THE MONITOR SHOULD BE SOFTWARE-UPGRADEABLE TO INCLUDE FUTURE UPDATES AND/OR ADDITIONS OF SOFTWARE APPLICATION PACKAGES. | YES
20 | THE VENDOR SHOULD PROVIDE DETAILED TECHNICAL INFORMATION FOR THE UNIQUE FEATURE IN THE MONITOR. | YES
21 | THE VENDOR SHOULD PROVIDE DETAILED TECHNICAL INFORMATION ABOUT THE FOLLOWING: | YES
22 | HR ACCURACY | YES
23 | ECG HR RANGE, BPM | YES
24 | INTERPRETATION | YES
25 | ARRHYTHMIA DETECT TECHNIQUE AND NUMBER OF LEAD ANALYZED | YES
26 | ST TECHNIQUE AND NO. OF LEADS ANALYZED | YES
27 | RESPIRATION METHOD AND WAVEFORM DISPLAYED FEATURE | YES
28 | DISPLAY SIZE, (CM) IN | YES
29 | DISPLAY USER INTERFACE FEATURE | YES
30 | REMOTE DISPLAY AVAILABILITY | YES
31 | TRENDING PARAMETERS | YES
32 | NETWORKING AND CENTRALIZATION. | YES
33 | HARDWIRED/WIRELESS | YES
34 | COMMUNICATION PROTOCOLS | YES
35 | ALARM TYPES AND FEATURE FOR THE MACHINE | YES
36 | THE MONITOR SHOULD BE CAPABLE OF RECORDING AND DISPLAY IN WAVEFORMS AND NUMERICAL VALUES THE FOLLOWING PARAMETERS (PLEASE PROVIDE IT AS ITEMIZED PRICE): | YES
37 | ECG, WITH 5 LEADS DISPLAY, DETECTION OF PACEMAKER SPIKES, AND ARRHYTHMIA. | YES
38 | RESPIRATION, WITH WAVEFORM DISPLAY AND ALARMS. | YES
39 | INVASIVE PRESSURE (IP). | YES
40 | PULSE OXIMETRY (SPO2), WITH NELLCOR-OXYMAX OR MASSIMO RD SET CABLES AND FINGER PROBES. | YES
41 | NON-INVASIVE BLOOD PRESSURE (NIBP), AND QUICK-CONNECT HOSES. | YES
42 | TEMPERATURE, IN TWO CHANNELS FOR TWO SIMULTANEOUS PROBES, SKIN AND RECTAL. | YES
43 | END-TIDAL CO2 (ETCO2), WITH WAVEFORM DISPLAY, RAPID WARM-UP AND CALIBRATION TIMES, AND LOW DEAD-SPACE CONNECTORS. | YES
44 | CARDIAC OUTPUT. | YES
45 | LEVEL OF CONSCIOUSNESS. | YES
46 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR DIFFERENT MODULES. | YES
47 | THE MONITOR SHOULD HAVE A MODULE THAT CAN MEASURE AT THE SAME TIME: | YES
48 | ECG, RESP. | YES, QTY: 1
49 | PULSE OXIMETRY (SPO2), NELLCOR-OXYMAX OR MASSIMO. | YES, QTY: 1
50 | NON-INVASIVE BLOOD PRESSURE (NIBP). | YES, QTY: 1
51 | INVASIVE BLOOD PRESSURE (IBP). | YES, QTY: 2
52 | TEMPERATURE. | YES, QTY: 1
53 | ALL PATIENT INPUT SHOULD BE DEFIBRILLATOR-PROTECTED. | YES
54 | ALL CABLES, PROBES, AND TRANSDUCERS (EXCEPT FOR DISPOSABLES) FOR ALL REQUIRED PARAMETERS SHOULD BE PROVIDED WITH THE MONITOR, FOR ADULTS AND PEDIATRIC. | YES
55 | ALL HARDWARE AND ACCESSORIES FOR EACH MODULE SHOULD BE DESIGNED FOR ALL TYPE OF PATIENT USE (PROVIDE ITEMIZED PRICE). | YES
56 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL OPTION, ACCESSORIES ON THE MACHINE. | YES
57 | THE MANUFACTURER SHOULD GUARANTY THAT SPARE PARTS AND TECHNICAL SUPPORT WILL BE PROVIDED FOR AT LEAST TEN YEARS. | YES
58 | HARDWARE AND SOFTWARE SHOULD BE INDUSTRY STANDARD WITH THE LATEST SPECIFICATION; THE VENDOR SHOULD PROVIDE A SPECIFICATION OF BOTH HARDWARE AND SOFTWARE USED IN THE SYSTEM. | YES
59 | THE VENDOR SHOULD PROVIDE AN ITEMIZED PRICE FOR ALL OPTIONS IN THE SYSTEM. | YES
60 | THE UNIT SHOULD BE PROVIDED WITH A LINE (POWER) CORD OF ACCEPTABLE DURABILITY, QUALITY, LENGTH, AND SHOULD BE SECURED WITH ADEQUATE STRAIN RELIEFS. | YES
61 | THE CHASSIS SHOULD BE GROUNDED AND GROUNDING RESISTANCE SHOULD NOT EXCEED 0.15 OHM. | YES
62 | IF THE UNIT IS DOUBLE INSULATED, IT SHOULD BE SO LABELED. | YES
63 | ELECTRICAL LEAKAGE CURRENT FROM THE CHASSIS OF THE SYSTEM SHOULD NOT EXCEED [0.5 MA PER IEC 601-1 OR 0.3 MA IN THE U.S. PER NFPA 99-1993].OPERATOR SAFETY AND SYSTEM PERFORMANCE SHOULD NOT BE ADVERSELY AFFECTED BY FLUID SPILLS. | YES
64 | THE SYSTEM PERFORMANCE SHOULD NOT BE AFFECTED BY EMI RADIATED OR CONDUCTED THROUGH THE POWER LINES FROM ANOTHER DEVICE. | YES
65 | THE POWER CORD SHOULD BE AUTOMATIC RETRACTABLE. | YES
66 | THE VENDOR SHOULD PROVIDE A FULL WARRANTY OF AT LEAST THREE (3) YEARS. | YES
67 | GENERAL SPECIFICATIONS AND REQUIREMENTS: | YES
68 | POWER SUPPLY: 230 VAC 60 HZ. | YES
69 | HOSPITAL GRADE BRITISH TYPE POWER PLUGS, LINE CORDS, AND STRAIN RELIEF. | YES
70 | THE UNIT MUST MEET THE APPLICABLE REQUIREMENTS AND STANDARDS OF THE UNDERWRITERS LABORATORIES (UL), CE MARK, IEC, ISO, AND/OR FDA APPROVAL. | YES
71 | THE UNIT SHOULD BE MANUFACTURED ACCORDING TO GMP GUIDELINES. | YES
72 | LABELED WITH ALL APPROPRIATE OPERATION AND SAFETY TAGS AND SYMBOLS. | YES
73 | SWITCHES, KNOBS, AND OTHER CONTROLS SHOULD BE VISIBLE AND CLEARLY IDENTIFIED, AND THEIR FUNCTIONS SHOULD BE SELF-EVIDENT. | YES
74 | SWITCHES, KNOBS, AND OTHER CONTROLS SHOULD BE PROTECTED AGAINST ACCIDENTAL SETTING CHANGES AND SHOULD BE SEALED AGAINST PENETRATION OF LIQUIDS. | YES
75 | THE UNIT SHOULD BE SIMPLE TO OPERATE AND EASY TO CLEAN AND DISINFECT WITH ISOPROPANOL ALCOHOL 70%. | YES
76 | THE EXTERIOR SHOULD BE WELL CONSTRUCTED WITH DURABLE FLUID/SHOCK RESISTANT MATERIALS TO WITHSTAND TYPICAL ABUSE AND CLEANING. | YES
77 | RUGGED TO HANDLE IN THE ROUGH TRANSPORT ENVIRONMENT. | YES
78 | THE UNIT SHOULD HAVE NO SHARP EDGES. ALL EXTERNAL COMPONENTS SHOULD BE SECURELY MOUNTED. | YES
79 | ALL AVAILABLE OPTIONS, CONSUMABLES, OR DISPOSABLES DEEMED NECESSARY FOR THE INTENDED FUNCTION OF THE EQUIPMENT AND NOT LISTED IN THE TECHNICAL SPECIFICATIONS SHOULD BE INCLUDED AND SPECIFIED. | PROVIDE ITEMIZED PRICE
80 | WALL MOUNT BRACKETS AS ITEMIZED PRICE | PROVIDE ITEMIZED PRICE
81 | ETCO2 ITEMIZED | PROVIDE ITEMIZED PRICE
82 | CENTRAL STATION ITEMIZED (LICENSE OF 20 BEDS) | PROVIDE ITEMIZED PRICE
83 | ITEMIZED: THE MONITOR SHOULD HAVE A MODULE THAT CAN MEASURE 8 WAVES AT THE SAME TIME AS FOLLOWS : | PROVIDE ITEMIZED PRICE
84 | QTY: 1 ECG, RESP. | PROVIDE ITEMIZED PRICE
85 | QTY: 1 PULSE OXIMETRY (SPO2), NELLCOR-OXYMAX OR MASSIMO. | PROVIDE ITEMIZED PRICE
86 | QTY: 1 NON-INVASIVE BLOOD PRESSURE (NIBP). | PROVIDE ITEMIZED PRICE
87 | QTY: 3 INVASIVE BLOOD PRESSURE (IBP). | PROVIDE ITEMIZED PRICE
88 | QTY: 2 TEMPERATURE. | PROVIDE ITEMIZED PRICE
Company Stamp |  |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | SAR10803 | 0 | MANOMETER CUFF ETT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MANOMETER CUFF ETT | FOR TRACHEAL & LARYNGEAL TUBES
2 | HANGER HOOK | INCLUDED WITH DEVICE
3 | ANALOG GAUGE | YES
4 | GAUGE SCALE | 0 - 100 CM H2O
5 | GREEN ZONE INDICATOR | YES
Company Stamp |  |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | MGE20022 | 0 | CYLINDER OXYGEN SIZE K/H WITH FLOW METER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CYLINDER SIZE H | FOR MEDICAL OXYGEN
2 | CONSTRUCTION | STEEL
3 | CAPACITY | 7,700 LITERS APPROX
4 | REGULATORS / VALVE | CGA 540
5 | FLOW METER 15L/MIN | INCLUDED WITH CYLINDER
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | MWD10034 | 0 | VENTILATOR BIPAP IN-PATIENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MODES | NIV, NIV-ST , HIFLOW
2 | CPAP | 4 TO 20 CM H2O
3 | S/T | 4 TO 40 CM H2O
4 | PARAMETERS CONTROL RANGES
5 | IPAP | 4 TO 40 CM H2O
6 | EPAP | 4 TO 20 CM H2O
7 | CPAP | 4 TO 20 CM H2O
8 | RATE | 4 TO 40 BPM
9 | TIMED INSPIRATION | 0.5 TO 3.0 SEC
10 | IPAP >RISE TIME | 0.05 TO 0.40 SEC
11 | OXYGEN CONCENTRATION (%O2) | 21 TO 100%
12 | DISPLAY RANGES
13 | IPAP | 0 TO 50 CM H2O
14 | EPAP | 0 TO 50 CM H2O
15 | CPAP | 0 TO 50 CM H2O
16 | RATE | 0 TO 120 BPM
17 | EXHALED TIDAL VOLUME (VT) | 20 TO 2000 ML
18 | MINUTE VENTILATION (MIN VENT) | 0 TO 99 L/MIN
19 | TOTAL LEAK (TOT LEAK) | 0 TO 80%
20 | PATIENT LEAK(PT. LEAK) | 0 TO 100 L/MIN
21 | PEAK INSPIRATORY PRESSURE (PIP) | 0 TO 50 CM H2O
22 | PERCENT OF PATIENT TRIGGERED BREATHS (PT. TRIG) | 0 TO 100%
23 | IT/TOT | 0 TO 100%
24 | ALARM CONTROL RANGES
25 | HIGH PRESSURE | 5 TO 60 CM H2O
26 | LOW PRESSURE | DISABLED;0 TO 40 CM H2O
27 | LOW PRESSURE | DELAY 0 TO 60 SEC
28 | APNEA | DISABLED; 20 TO 40 SEC
29 | LOW MINUTE VENTILATION | DISABLED; 1 TO 99 L/MIN
30 | HIGH RATE | 4 TO 120 BPM
31 | LOW RATE | 4 TO 120 BPM
32 | POWER SUPPLY | 220 VAC - 60 HZ.
33 | INSPIRATORY FLOW ON NIV, NIV-ST MODE | 2-220 L/MIN
34 | FLOW ON HIFLOW MODE | 2-70 L/MIN
35 | HI FLOW CANULA, OPEN SYSTEM | YES
36 | SCREEN , AND MONITORING | NOT LESS THAN 7' COLOR , TOUCH
37 | HUMIDIFIER | REQUIRED
38 | BUILT IN TURABIAN | YES
39 | OTHER SPECIFICATION | FDA, CE, SPECIFY OTHERS
40
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | MGE40004 | 0 | STAND IV |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FRAME | STAINLESS STEEL 304 PREFERRED
2 | TYPE | TELESCOPIC
3 | HOOKS | 4 - 5 APPROX
4 | HEIGHT | ADJUSTABLE
5 | BASE | MOBILE HEAVY BASE TO WITHSTAND THE HANGED ITEMS WEIGHTS
6 | CASTORS | QTY : 5
7 | CAPACITY | 5 KG PER HOOK
8
Company Stamp |  |  |  | Signature