## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | COMPRESSION DEVICE CHEST
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | AED SYSTEM
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | RESUSCITATION KIT AMBULATORY
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | VENTILATOR TRANSPORT ADULT & PEDIATRIC
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | BRONCHOSCOPE WORKSTATION VIDEO
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | ENDOSCOPY WORKSTATION
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | CABINETS STORAGE INSTRUMENT
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | CART CHART PATIENT
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | CART CRASH
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | CART DRESSING
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | CYLINDER O2 SIZE D WITH REGULATOR
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | FLOWMETER AIR
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | FLOWMETER O2 15 L/Min WALL MOUNTED
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | REGULATOR O2
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | REGULATOR SUCTION HIGH WALL MOUNT
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | WHEELCHAIR PEDIATRIC
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | HAMPER LINEN DOUBLE
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | HAMPER LINEN SINGLE
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | STAND IV
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | STOOL FOOT 2 STEP
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | TABLE INSTRUMENT LARGE
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | TROLLEY CYLINDER SIZE F/G WITH STABILISERS
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | TRUCK CYLINDER SIZE E
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | DEFIBRILLATOR UNIT
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | DIAGNOSTIC SET WALL MOUNTED
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | ECG 12 LEAD WITH CART
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | GLUCOMETER
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | LARYNGOSCOPE ADULT
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | LARYNGOSCOPE ADULT & PEDIATRIC
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | LARYNGOSCOPE NEONATE
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | LARYNGOSCOPE PEDIATRIC
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | LIGHT EXAM MOBILE
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | MATTRESS AIR MED RISK
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | MATTRESS ANTIDECUBITUS
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | NEBULIZER ULTRASONIC
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | OTOSCOPE TABLETOP
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | OXIMETER PULSE HAND HELD
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | PUMP FEEDING
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | PUMP INFUSION
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | PUMP INSULINE
ITEM 42 | This sheet contains the details of the bid with each technical specification for | ITEM 42 | PUMP SYRINGE
ITEM 43 | This sheet contains the details of the bid with each technical specification for | ITEM 43 | SPHYGMOMANOMETER ANEROID MOBILE
ITEM 44 | This sheet contains the details of the bid with each technical specification for | ITEM 44 | SPHYGMOMANOMETER ANEROID WALL
ITEM 45 | This sheet contains the details of the bid with each technical specification for | ITEM 45 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN
ITEM 46 | This sheet contains the details of the bid with each technical specification for | ITEM 46 | STETHOSCOPE ADULT
ITEM 47 | This sheet contains the details of the bid with each technical specification for | ITEM 47 | THERMOMETER DIGITAL DOCKING STATION
ITEM 48 | This sheet contains the details of the bid with each technical specification for | ITEM 48 | THERMOMETER DIGITAL HANDHELD
ITEM 49 | This sheet contains the details of the bid with each technical specification for | ITEM 49 | THERMOMETER DIGITAL WALL MOUNT
ITEM 50 | This sheet contains the details of the bid with each technical specification for | ITEM 50 | THERMOMETER INFRARED
ITEM 51 | This sheet contains the details of the bid with each technical specification for | ITEM 51 | THERMOMETER TYMPANIC
ITEM 52 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | TRANSILLUMINATOR VEIN HANDHELD
ITEM 53 | This sheet contains the details of the bid with each technical specification for | ITEM 53 | OTO OPTHALMOSCOPE WALL MOUNT
ITEM 54 | This sheet contains the details of the bid with each technical specification for | ITEM 54 | SCALE CHAIR
ITEM 55 | This sheet contains the details of the bid with each technical specification for | ITEM 55 | SCALE INFANT
ITEM 56 | This sheet contains the details of the bid with each technical specification for | ITEM 56 | SCALE PATIENT WITH HEIGHT
ITEM 57 | This sheet contains the details of the bid with each technical specification for | ITEM 57 | BAG AMBO ADULT & PEDIATRIC
ITEM 58 | This sheet contains the details of the bid with each technical specification for | ITEM 58 | RESUSCITATOR KIT ADULT & PEDIATRIC
ITEM 59 | This sheet contains the details of the bid with each technical specification for | ITEM 59 | SUCTION UNIT ELECTRICAL MOBILE
ITEM 60 | This sheet contains the details of the bid with each technical specification for | ITEM 60 | WHEELCHAIR STANDARD SIZE 24
ITEM 61 | This sheet contains the details of the bid with each technical specification for | ITEM 61 | CHAIR BATH LARGE
ITEM 62 | This sheet contains the details of the bid with each technical specification for | ITEM 62 | CRRT UNIT
ITEM 63 | This sheet contains the details of the bid with each technical specification for | ITEM 63 | DIALYSIS HAEMOFILTRATION
ITEM 64 | This sheet contains the details of the bid with each technical specification for | ITEM 64 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO
ITEM 65 | This sheet contains the details of the bid with each technical specification for | ITEM 65 | ANALYZER BLOOD GAS POC
ITEM 66 | This sheet contains the details of the bid with each technical specification for | ITEM 66 | DRAINAGE UNIT CHEST
ITEM 67 | This sheet contains the details of the bid with each technical specification for | ITEM 67 | INTUBATION DIFFICULT VIDEO ADULT
ITEM 68 | This sheet contains the details of the bid with each technical specification for | ITEM 68 | INTUBATION DIFFICULT VIDEO PEDIATRIC
ITEM 69 | This sheet contains the details of the bid with each technical specification for | ITEM 69 | MONITOR CAPNOGRAPHY
ITEM 70 | This sheet contains the details of the bid with each technical specification for | ITEM 70 | MONITOR PHYSIOLOGICAL BEDSIDE 17 INCH
ITEM 71 | This sheet contains the details of the bid with each technical specification for | ITEM 71 | MONITOR PHYSIOLOGICAL BEDSIDE 21 INCH
ITEM 72 | This sheet contains the details of the bid with each technical specification for | ITEM 72 | MONITOR PHYSIOLOGICAL MRI COMPATIBLE
ITEM 73 | This sheet contains the details of the bid with each technical specification for | ITEM 73 | MONITOR PHYSIOLOGICAL TRANSPORT
ITEM 74 | This sheet contains the details of the bid with each technical specification for | ITEM 74 | NEGATIVE AIR SYSTEM WITH HEPAFILTER
ITEM 75 | This sheet contains the details of the bid with each technical specification for | ITEM 75 | PACEMAKER CARDIAC
ITEM 76 | This sheet contains the details of the bid with each technical specification for | ITEM 76 | PHOTOTHERAPY DOUBLE PORTABLE
ITEM 77 | This sheet contains the details of the bid with each technical specification for | ITEM 77 | THROMBOSIS PREVENTION VIEN DEEP
ITEM 78 | This sheet contains the details of the bid with each technical specification for | ITEM 78 | VENTILATOR BIPAP ICU
ITEM 79 | This sheet contains the details of the bid with each technical specification for | ITEM 79 | VENTILATOR CPAP ADULT
ITEM 80 | This sheet contains the details of the bid with each technical specification for | ITEM 80 | VENTILATOR HFO ADULT & PEDIATRIC
ITEM 81 | This sheet contains the details of the bid with each technical specification for | ITEM 81 | VENTILATOR ICU ADULT & PEDIATRIC
ITEM 82 | This sheet contains the details of the bid with each technical specification for | ITEM 82 | VENTILATOR ICU NEONATAL
ITEM 83 | This sheet contains the details of the bid with each technical specification for | ITEM 83 | VENTILATOR TRANSPORT MRI COMPATABLE
ITEM 84 | This sheet contains the details of the bid with each technical specification for | ITEM 84 | VENTILATOR TRANSPORT WITH MONITOR
ITEM 85 | This sheet contains the details of the bid with each technical specification for | ITEM 85 | ULTRASOUND POC ADVANCE PORTABLE
ITEM 86 | This sheet contains the details of the bid with each technical specification for | ITEM 86 | CENTRAL STATION PHYSIOLOGICAL 4 BEDS LICENCE
ITEM 87 | This sheet contains the details of the bid with each technical specification for | ITEM 87 | CENTRAL STATION PHYSIOLOGICAL 8 BEDS LICENCE
ITEM 88 | This sheet contains the details of the bid with each technical specification for | ITEM 88 | CENTRAL STATION PHYSIOLOGICAL 12 BEDS LICENCE
ITEM 89 | This sheet contains the details of the bid with each technical specification for | ITEM 89 | CENTRAL STATION PHYSIOLOGICAL 20 BEDS LICENCE
ITEM 90 | This sheet contains the details of the bid with each technical specification for | ITEM 90 | BEDSIDE PHYSIOLOGICAL MONITOR 14 BED STATION
ITEM 91 | This sheet contains the details of the bid with each technical specification for | ITEM 91 | ANALYZER NITRIC OXIDE NEONATE
ITEM 92 | This sheet contains the details of the bid with each technical specification for | ITEM 92 | INCUBATOR PREMIUM NEONATE
ITEM 93 | This sheet contains the details of the bid with each technical specification for | ITEM 93 | INCUBATOR STANDARD NEONATAL
ITEM 94 | This sheet contains the details of the bid with each technical specification for | ITEM 94 | INCUBATORS TRANSPORT INFANT WITH VENTILATOR
ITEM 95 | This sheet contains the details of the bid with each technical specification for | ITEM 95 | INTUBATION DIFFICULT VIDEO NEONATE
ITEM 96 | This sheet contains the details of the bid with each technical specification for | ITEM 96 | RESUSCITATOR & WARMER INFANT
ITEM 97 | This sheet contains the details of the bid with each technical specification for | ITEM 97 | THERAPY HIGH FLOW NEONATE
ITEM 98 | This sheet contains the details of the bid with each technical specification for | ITEM 98 | VENTILATOR CPAP NASAL NEONATE
ITEM 99 | This sheet contains the details of the bid with each technical specification for | ITEM 99 | VENTILATOR HFO NEONATE
ITEM 100 | This sheet contains the details of the bid with each technical specification for | ITEM 100 | WARMER INFANT WITH BASSINET
ITEM 101 | This sheet contains the details of the bid with each technical specification for | ITEM 101 | OZONE AND HYDROGEN PEROXIDE DISINFECTION SYSTEM
ITEM 102 | This sheet contains the details of the bid with each technical specification for | ITEM 102 | SURFACE DECONTAMINATION SYSTEM H2O2
ITEM 103 | This sheet contains the details of the bid with each technical specification for | ITEM 103 | CHAMBER ISOLATION PATIENT PORTABLE
ITEM 104 | This sheet contains the details of the bid with each technical specification for | ITEM 104 | APHERESIS UNIT
ITEM 105 | This sheet contains the details of the bid with each technical specification for | ITEM 105 | CABINET FLAMMABLE STORAGE
ITEM 106 | This sheet contains the details of the bid with each technical specification for | ITEM 106 | CART UTILITY
ITEM 107 | This sheet contains the details of the bid with each technical specification for | ITEM 107 | CHAIR BLOOD DRAWING
ITEM 108 | This sheet contains the details of the bid with each technical specification for | ITEM 108 | ANALYZER AUTOMATED BLOOD GAS PH
ITEM 109 | This sheet contains the details of the bid with each technical specification for | ITEM 109 | X-RAY DIGITAL MOBILE
ITEM 110 | This sheet contains the details of the bid with each technical specification for | ITEM 110 | PROTECTION X-RAY APRON
ITEM 111 | This sheet contains the details of the bid with each technical specification for | ITEM 111 | PUMP INFUSION & SYRINGE MRI COMPATIBLE
ITEM 112 | This sheet contains the details of the bid with each technical specification for | ITEM 112 | ULTRASOUND 4D GENERAL STATION
ITEM 113 | This sheet contains the details of the bid with each technical specification for | ITEM 113 | ULTRASOUND 4D GENERAL PORTABLE
ITEM 114 | This sheet contains the details of the bid with each technical specification for | ITEM 114 | ULTRASOUND 4D OBGYN PORTABLE
ITEM 115 | This sheet contains the details of the bid with each technical specification for | ITEM 115 | ULTRASOUND CARDIAC ADVANCE STATION
ITEM 116 | This sheet contains the details of the bid with each technical specification for | ITEM 116 | ULTRASOUND CARDIAC PORTABLE
ITEM 117 | This sheet contains the details of the bid with each technical specification for | ITEM 117 | ULTRASOUND GENERAL PORTABLE
ITEM 118 | This sheet contains the details of the bid with each technical specification for | ITEM 118 | ULTRASOUND VASCULAR PORTABLE
ITEM 119 | This sheet contains the details of the bid with each technical specification for | ITEM 119 | EEG UNIT
ITEM 120 | This sheet contains the details of the bid with each technical specification for | ITEM 120 | CABINET WARMER BLANKET
ITEM 121 | This sheet contains the details of the bid with each technical specification for | ITEM 121 | FIXATOR SKULL
ITEM 122 | This sheet contains the details of the bid with each technical specification for | ITEM 122 | SUCTION UNIT HIGH VACUUM
ITEM 123 | This sheet contains the details of the bid with each technical specification for | ITEM 123 | TABLE MAYO SMALL
ITEM 124 | This sheet contains the details of the bid with each technical specification for | ITEM 124 | TABLE OPERATING ENT
ITEM 125 | This sheet contains the details of the bid with each technical specification for | ITEM 125 | TABLE OPERATING GENERAL
ITEM 126 | This sheet contains the details of the bid with each technical specification for | ITEM 126 | TELESCOPE 90 DEG
ITEM 127 | This sheet contains the details of the bid with each technical specification for | ITEM 127 | ULTRASOUND VASCULAR DOPPLER
ITEM 128 | This sheet contains the details of the bid with each technical specification for | ITEM 128 | ANESTHESIA MACHINE MRI COMPATIBLE
ITEM 129 | This sheet contains the details of the bid with each technical specification for | ITEM 129 | BLANKET WARMING SYSTEM AIR
ITEM 130 | This sheet contains the details of the bid with each technical specification for | ITEM 130 | CART ANESTHESIA SUPPLY
ITEM 131 | This sheet contains the details of the bid with each technical specification for | ITEM 131 | COMPRESSION SYSTEM SEQUENTIAL
ITEM 132 | This sheet contains the details of the bid with each technical specification for | ITEM 132 | INTUBATION SET COMPLETE CASE
ITEM 133 | This sheet contains the details of the bid with each technical specification for | ITEM 133 | INTUBATION DIFFICULT VIDEO
ITEM 134 | This sheet contains the details of the bid with each technical specification for | ITEM 134 | LARYNGOSCOPE SET COMPLETE
ITEM 135 | This sheet contains the details of the bid with each technical specification for | ITEM 135 | PUMP PCA
ITEM 136 | This sheet contains the details of the bid with each technical specification for | ITEM 136 | TOURNIQUET SYSTEM ANESTHESIA
ITEM 137 | This sheet contains the details of the bid with each technical specification for | ITEM 137 | WARMER BLOOD
ITEM 138 | This sheet contains the details of the bid with each technical specification for | ITEM 138 | WARMER BLOOD RAPID INFUSION
ITEM 139 | This sheet contains the details of the bid with each technical specification for | ITEM 139 | ASPIRATION SYSTEM THROMBUS
ITEM 140 | This sheet contains the details of the bid with each technical specification for | ITEM 140 | MONITOR CARDIAC OUTPUT THROUGH ESOPHAGUS
ITEM 141 | This sheet contains the details of the bid with each technical specification for | ITEM 141 | PUMP BALLON
ITEM 142 | This sheet contains the details of the bid with each technical specification for | ITEM 142 | REFRIGERATOR MEDICATION 700L
ITEM 143 | This sheet contains the details of the bid with each technical specification for | ITEM 143 | REFRIGERATOR MEDICATION 160L
ITEM 144 | This sheet contains the details of the bid with each technical specification for | ITEM 144 | CABINET MEDICATION
ITEM 145 | This sheet contains the details of the bid with each technical specification for | ITEM 145 | CABINET NARCOTIC
ITEM 146 | This sheet contains the details of the bid with each technical specification for | ITEM 146 | CART MEDICATION
ITEM 147 | This sheet contains the details of the bid with each technical specification for | ITEM 147 | CART DISPENSING
ITEM 148 | This sheet contains the details of the bid with each technical specification for | ITEM 148 | SAFE NARCOTIC LARGE
ITEM 149 | This sheet contains the details of the bid with each technical specification for | ITEM 149 | SAFE NARCOTIC SMALL
ITEM 150 | This sheet contains the details of the bid with each technical specification for | ITEM 150 | BALANCE APPARATUS
ITEM 151 | This sheet contains the details of the bid with each technical specification for | ITEM 151 | GAUGE ROOM PRESSURE
ITEM 152 | This sheet contains the details of the bid with each technical specification for | ITEM 152 | BED BURN
ITEM 153 | This sheet contains the details of the bid with each technical specification for | ITEM 153 | BED ELECTRICAL PEDIATRIC
ITEM 154 | This sheet contains the details of the bid with each technical specification for | ITEM 154 | BED INTENSIVE CARE
ITEM 155 | This sheet contains the details of the bid with each technical specification for | ITEM 155 | BED INTENSIVE CARE PEDIATRIC
ITEM 156 | This sheet contains the details of the bid with each technical specification for | ITEM 156 | CABINET BEDSIDE
ITEM 157 | This sheet contains the details of the bid with each technical specification for | ITEM 157 | CRIB NEWBORN
ITEM 158 | This sheet contains the details of the bid with each technical specification for | ITEM 158 | STRETCHER EMERGENCY
ITEM 159 | This sheet contains the details of the bid with each technical specification for | ITEM 159 | STRETCHER RECOVERY
ITEM 160 | This sheet contains the details of the bid with each technical specification for | ITEM 160 | STRETCHER TRANSPORT
ITEM 161 | This sheet contains the details of the bid with each technical specification for | ITEM 161 | TABLE OVERBED
ITEM 162 | This sheet contains the details of the bid with each technical specification for | ITEM 162 | BED INTENSIVE CARE BARIATRIC
ITEM 163 | This sheet contains the details of the bid with each technical specification for | ITEM 163 | ‫FREEZER 21 BODY
ITEM 164 | This sheet contains the details of the bid with each technical specification for | ITEM 164 | STRETCHER MORTUARY

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | MAL20002 | COMPRESSION DEVICE CHEST |  | Medical Eqpt | 5 | [10 empty] | 0
2 | MAL20011 | AED SYSTEM |  | Medical Eqpt | 3 | [10 empty] | 0
3 | MAL20013 | RESUSCITATION KIT AMBULATORY |  | Medical Eqpt | 1 | [10 empty] | 0
4 | MAL20015 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | Medical Eqpt | 117 | [10 empty] | 0
5 | MED10001 | BRONCHOSCOPE WORKSTATION VIDEO |  | Medical Eqpt | 3 | [10 empty] | 0
6 | MED10006 | ENDOSCOPY WORKSTATION |  | Medical Eqpt | 1 | [10 empty] | 0
7 | MGE10006 | CABINETS STORAGE INSTRUMENT |  | Medical Eqpt | 10 | [10 empty] | 0
8 | MGE10007 | CART CHART PATIENT |  | Medical Eqpt | 8 | [10 empty] | 0
9 | MGE10009 | CART CRASH |  | Medical Eqpt | 127 | [10 empty] | 0
10 | MGE10011 | CART DRESSING |  | Medical Eqpt | 31 | [10 empty] | 0
11 | MGE20006 | CYLINDER O2 SIZE D WITH REGULATOR |  | Medical Eqpt | 4 | [10 empty] | 0
12 | MGE20011 | FLOWMETER AIR |  | Medical Eqpt | 1090 | [10 empty] | 0
13 | MGE20014 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | Medical Eqpt | 411 | [10 empty] | 0
14 | MGE20017 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | Medical Eqpt | 240 | [10 empty] | 0
15 | MGE20019 | REGULATOR O2 |  | Medical Eqpt | 699 | [10 empty] | 0
16 | MGE20020 | REGULATOR SUCTION HIGH WALL MOUNT |  | Medical Eqpt | 591 | [10 empty] | 0
17 | MGE30007 | WHEELCHAIR PEDIATRIC |  | Medical Eqpt | 12 | [10 empty] | 0
18 | MGE40001 | HAMPER LINEN DOUBLE |  | Medical Eqpt | 66 | [10 empty] | 0
19 | MGE40002 | HAMPER LINEN SINGLE |  | Medical Eqpt | 17 | [10 empty] | 0
20 | MGE40004 | STAND IV |  | Medical Eqpt | 706 | [10 empty] | 0
21 | MGE40006 | STOOL FOOT 2 STEP |  | Medical Eqpt | 10 | [10 empty] | 0
22 | MGE40007 | TABLE INSTRUMENT LARGE |  | Medical Eqpt | 10 | [10 empty] | 0
23 | MGE40009 | TROLLEY CYLINDER SIZE F/G WITH STABILISERS |  | Medical Eqpt | 2 | [10 empty] | 0
24 | MGE40010 | TRUCK CYLINDER SIZE E |  | Medical Eqpt | 2 | [10 empty] | 0
25 | MGE50001 | DEFIBRILLATOR UNIT |  | Medical Eqpt | 114 | [10 empty] | 0
26 | MGE50002 | DIAGNOSTIC SET WALL MOUNTED |  | Medical Eqpt | 35 | [10 empty] | 0
27 | MGE50005 | ECG 12 LEAD WITH CART |  | Medical Eqpt | 96 | [10 empty] | 0
28 | MGE50006 | GLUCOMETER |  | Medical Eqpt | 80 | [10 empty] | 0
29 | MGE50009 | LARYNGOSCOPE ADULT |  | Medical Eqpt | 99 | [10 empty] | 0
30 | MGE50010 | LARYNGOSCOPE ADULT & PEDIATRIC |  | Medical Eqpt | 13 | [10 empty] | 0
31 | MGE50011 | LARYNGOSCOPE NEONATE |  | Medical Eqpt | 10 | [10 empty] | 0
32 | MGE50012 | LARYNGOSCOPE PEDIATRIC |  | Medical Eqpt | 30 | [10 empty] | 0
33 | MGE50014 | LIGHT EXAM MOBILE |  | Medical Eqpt | 17 | [10 empty] | 0
34 | MGE50017 | MATTRESS AIR MED RISK |  | Medical Eqpt | 166 | [10 empty] | 0
35 | MGE50018 | MATTRESS ANTIDECUBITUS |  | Medical Eqpt | 159 | [10 empty] | 0
36 | MGE50020 | NEBULIZER ULTRASONIC |  | Medical Eqpt | 37 | [10 empty] | 0
37 | MGE50021 | OTOSCOPE TABLETOP |  | Medical Eqpt | 6 | [10 empty] | 0
38 | MGE50024 | OXIMETER PULSE HAND HELD |  | Medical Eqpt | 199 | [10 empty] | 0
39 | MGE50026 | PUMP FEEDING |  | Medical Eqpt | 194 | [10 empty] | 0
40 | MGE50027 | PUMP INFUSION |  | Medical Eqpt | 216 | [10 empty] | 0
41 | MGE50028 | PUMP INSULINE |  | Medical Eqpt | 12 | [10 empty] | 0
42 | MGE50029 | PUMP SYRINGE |  | Medical Eqpt | 1210 | [10 empty] | 0
43 | MGE50035 | SPHYGMOMANOMETER ANEROID MOBILE |  | Medical Eqpt | 5 | [10 empty] | 0
44 | MGE50036 | SPHYGMOMANOMETER ANEROID WALL |  | Medical Eqpt | 10 | [10 empty] | 0
45 | MGE50037 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | Medical Eqpt | 654 | [10 empty] | 0
46 | MGE50038 | STETHOSCOPE ADULT |  | Medical Eqpt | 45 | [10 empty] | 0
47 | MGE50042 | THERMOMETER DIGITAL DOCKING STATION |  | Medical Eqpt | 35 | [10 empty] | 0
48 | MGE50043 | THERMOMETER DIGITAL HANDHELD |  | Medical Eqpt | 174 | [10 empty] | 0
49 | MGE50044 | THERMOMETER DIGITAL WALL MOUNT |  | Medical Eqpt | 55 | [10 empty] | 0
50 | MGE50045 | THERMOMETER INFRARED |  | Medical Eqpt | 180 | [10 empty] | 0
51 | MGE50046 | THERMOMETER TYMPANIC |  | Medical Eqpt | 4 | [10 empty] | 0
52 | MGE50048 | TRANSILLUMINATOR VEIN HANDHELD |  | Medical Eqpt | 36 | [10 empty] | 0
53 | MGE50058 | OTO OPTHALMOSCOPE WALL MOUNT |  | Medical Eqpt | 109 | [10 empty] | 0
54 | MGE60001 | SCALE CHAIR |  | Medical Eqpt | 2 | [10 empty] | 0
55 | MGE60002 | SCALE INFANT |  | Medical Eqpt | 16 | [10 empty] | 0
56 | MGE60003 | SCALE PATIENT WITH HEIGHT |  | Medical Eqpt | 22 | [10 empty] | 0
57 | MGE70001 | BAG AMBO ADULT & PEDIATRIC |  | Medical Eqpt | 4 | [10 empty] | 0
58 | MGE70003 | RESUSCITATOR KIT ADULT & PEDIATRIC |  | Medical Eqpt | 23 | [10 empty] | 0
59 | MGE80001 | SUCTION UNIT ELECTRICAL MOBILE |  | Medical Eqpt | 40 | [10 empty] | 0
60 | MHC10052 | WHEELCHAIR STANDARD SIZE 24 |  | Medical Eqpt | 51 | [10 empty] | 0
61 | MHC30008 | CHAIR BATH LARGE |  | Medical Eqpt | 5 | [10 empty] | 0
62 | MHD10001 | CRRT UNIT |  | Medical Eqpt | 27 | [10 empty] | 0
63 | MHD10004 | DIALYSIS HAEMOFILTRATION |  | Medical Eqpt | 2 | [10 empty] | 0
64 | MHD10005 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO |  | Medical Eqpt | 15 | [10 empty] | 0
65 | MIC10003 | ANALYZER BLOOD GAS POC |  | Medical Eqpt | 6 | [10 empty] | 0
66 | MIC10006 | DRAINAGE UNIT CHEST |  | Medical Eqpt | 27 | [10 empty] | 0
67 | MIC10009 | INTUBATION DIFFICULT VIDEO ADULT |  | Medical Eqpt | 6 | [10 empty] | 0
68 | MIC10010 | INTUBATION DIFFICULT VIDEO PEDIATRIC |  | Medical Eqpt | 4 | [10 empty] | 0
69 | MIC10018 | MONITOR CAPNOGRAPHY |  | Medical Eqpt | 41 | [10 empty] | 0
70 | MIC10026 | MONITOR PHYSIOLOGICAL BEDSIDE 17 INCH |  | Medical Eqpt | 61 | [10 empty] | 0
71 | MIC10027 | MONITOR PHYSIOLOGICAL BEDSIDE 21 INCH |  | Medical Eqpt | 27 | [10 empty] | 0
72 | MIC10028 | MONITOR PHYSIOLOGICAL MRI COMPATIBLE |  | Medical Eqpt | 3 | [10 empty] | 0
73 | MIC10029 | MONITOR PHYSIOLOGICAL TRANSPORT |  | Medical Eqpt | 378 | [10 empty] | 0
74 | MIC10032 | NEGATIVE AIR SYSTEM WITH HEPAFILTER |  | Medical Eqpt | 46 | [10 empty] | 0
75 | MIC10033 | PACEMAKER CARDIAC |  | Medical Eqpt | 6 | [10 empty] | 0
76 | MIC10037 | PHOTOTHERAPY DOUBLE PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
77 | MIC10042 | THROMBOSIS PREVENTION VIEN DEEP |  | Medical Eqpt | 15 | [10 empty] | 0
78 | MIC10043 | VENTILATOR BIPAP ICU |  | Medical Eqpt | 3 | [10 empty] | 0
79 | MIC10044 | VENTILATOR CPAP ADULT |  | Medical Eqpt | 13 | [10 empty] | 0
80 | MIC10045 | VENTILATOR HFO ADULT & PEDIATRIC |  | Medical Eqpt | 25 | [10 empty] | 0
81 | MIC10046 | VENTILATOR ICU ADULT & PEDIATRIC |  | Medical Eqpt | 25 | [10 empty] | 0
82 | MIC10047 | VENTILATOR ICU NEONATAL |  | Medical Eqpt | 4 | [10 empty] | 0
83 | MIC10048 | VENTILATOR TRANSPORT MRI COMPATABLE |  | Medical Eqpt | 4 | [10 empty] | 0
84 | MIC10049 | VENTILATOR TRANSPORT WITH MONITOR |  | Medical Eqpt | 14 | [10 empty] | 0
85 | MIC10054 | ULTRASOUND POC ADVANCE PORTABLE |  | Medical Eqpt | 6 | [10 empty] | 0
86 | MIC10056 | CENTRAL STATION PHYSIOLOGICAL 4 BEDS LICENCE |  | Medical Eqpt | 2 | [10 empty] | 0
87 | MIC10057 | CENTRAL STATION PHYSIOLOGICAL 8 BEDS LICENCE |  | Medical Eqpt | 7 | [10 empty] | 0
88 | MIC10058 | CENTRAL STATION PHYSIOLOGICAL 12 BEDS LICENCE |  | Medical Eqpt | 3 | [10 empty] | 0
89 | MIC10060 | CENTRAL STATION PHYSIOLOGICAL 20 BEDS LICENCE |  | Medical Eqpt | 4 | [10 empty] | 0
90 | MIC10069 | BEDSIDE PHYSIOLOGICAL MONITOR 14 BED STATION |  | Medical Eqpt | 11 | [10 empty] | 0
91 | MIC20001 | ANALYZER NITRIC OXIDE NEONATE |  | Medical Eqpt | 12 | [10 empty] | 0
92 | MIC20008 | INCUBATOR PREMIUM NEONATE |  | Medical Eqpt | 5 | [10 empty] | 0
93 | MIC20009 | INCUBATOR STANDARD NEONATAL |  | Medical Eqpt | 10 | [10 empty] | 0
94 | MIC20010 | INCUBATORS TRANSPORT INFANT WITH VENTILATOR |  | Medical Eqpt | 2 | [10 empty] | 0
95 | MIC20011 | INTUBATION DIFFICULT VIDEO NEONATE |  | Medical Eqpt | 4 | [10 empty] | 0
96 | MIC20017 | RESUSCITATOR & WARMER INFANT |  | Medical Eqpt | 1 | [10 empty] | 0
97 | MIC20020 | THERAPY HIGH FLOW NEONATE |  | Medical Eqpt | 550 | [10 empty] | 0
98 | MIC20023 | VENTILATOR CPAP NASAL NEONATE |  | Medical Eqpt | 5 | [10 empty] | 0
99 | MIC20024 | VENTILATOR HFO NEONATE |  | Medical Eqpt | 2 | [10 empty] | 0
100 | MIC20027 | WARMER INFANT WITH BASSINET |  | Medical Eqpt | 1 | [10 empty] | 0
101 | MIF40003 | OZONE AND HYDROGEN PEROXIDE DISINFECTION SYSTEM |  | Medical Eqpt | 11 | [10 empty] | 0
102 | MIF40004 | SURFACE DECONTAMINATION SYSTEM H2O2 |  | Medical Eqpt | 22 | [10 empty] | 0
103 | MIF50013 | CHAMBER ISOLATION PATIENT PORTABLE |  | Medical Eqpt | 3 | [10 empty] | 0
104 | MLB10003 | APHERESIS UNIT |  | Medical Eqpt | 4 | [10 empty] | 0
105 | MLB10027 | CABINET FLAMMABLE STORAGE |  | Medical Eqpt | 3 | [10 empty] | 0
106 | MLB10035 | CART UTILITY |  | Medical Eqpt | 19 | [10 empty] | 0
107 | MLB10045 | CHAIR BLOOD DRAWING |  | Medical Eqpt | 1 | [10 empty] | 0
108 | MLB20002 | ANALYZER AUTOMATED BLOOD GAS PH |  | Medical Eqpt | 2 | [10 empty] | 0
109 | MMI110012 | X-RAY DIGITAL MOBILE |  | Medical Eqpt | 28 | [10 empty] | 0
110 | MMI120001 | PROTECTION X-RAY APRON |  | Medical Eqpt | 10 | [10 empty] | 0
111 | MMI140003 | PUMP INFUSION & SYRINGE MRI COMPATIBLE |  | Medical Eqpt | 4 | [10 empty] | 0
112 | MMI90003 | ULTRASOUND 4D GENERAL STATION |  | Medical Eqpt | 2 | [10 empty] | 0
113 | MMI90004 | ULTRASOUND 4D GENERAL PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
114 | MMI90006 | ULTRASOUND 4D OBGYN PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
115 | MMI90008 | ULTRASOUND CARDIAC ADVANCE STATION |  | Medical Eqpt | 2 | [10 empty] | 0
116 | MMI90009 | ULTRASOUND CARDIAC PORTABLE |  | Medical Eqpt | 3 | [10 empty] | 0
117 | MMI90010 | ULTRASOUND GENERAL PORTABLE |  | Medical Eqpt | 4 | [10 empty] | 0
118 | MMI90014 | ULTRASOUND VASCULAR PORTABLE |  | Medical Eqpt | 13 | [10 empty] | 0
119 | MOP20002 | EEG UNIT |  | Medical Eqpt | 8 | [10 empty] | 0
120 | MOR10007 | CABINET WARMER BLANKET |  | Medical Eqpt | 22 | [10 empty] | 0
121 | MOR10029 | FIXATOR SKULL |  | Medical Eqpt | 15 | [10 empty] | 0
122 | MOR10087 | SUCTION UNIT HIGH VACUUM |  | Medical Eqpt | 12 | [10 empty] | 0
123 | MOR10093 | TABLE MAYO SMALL |  | Medical Eqpt | 71 | [10 empty] | 0
124 | MOR10095 | TABLE OPERATING ENT |  | Medical Eqpt | 1 | [10 empty] | 0
125 | MOR10096 | TABLE OPERATING GENERAL |  | Medical Eqpt | 5 | [10 empty] | 0
126 | MOR10114 | TELESCOPE 90 DEG |  | Medical Eqpt | 4 | [10 empty] | 0
127 | MOR10122 | ULTRASOUND VASCULAR DOPPLER |  | Medical Eqpt | 7 | [10 empty] | 0
128 | MOR20006 | ANESTHESIA MACHINE MRI COMPATIBLE |  | Medical Eqpt | 4 | [10 empty] | 0
129 | MOR20009 | BLANKET WARMING SYSTEM AIR |  | Medical Eqpt | 11 | [10 empty] | 0
130 | MOR20010 | CART ANESTHESIA SUPPLY |  | Medical Eqpt | 4 | [10 empty] | 0
131 | MOR20011 | COMPRESSION SYSTEM SEQUENTIAL |  | Medical Eqpt | 26 | [10 empty] | 0
132 | MOR20015 | INTUBATION SET COMPLETE CASE |  | Medical Eqpt | 25 | [10 empty] | 0
133 | MOR20017 | INTUBATION DIFFICULT VIDEO |  | Medical Eqpt | 37 | [10 empty] | 0
134 | MOR20019 | LARYNGOSCOPE SET COMPLETE |  | Medical Eqpt | 44 | [10 empty] | 0
135 | MOR20022 | PUMP PCA |  | Medical Eqpt | 4 | [10 empty] | 0
136 | MOR20026 | TOURNIQUET SYSTEM ANESTHESIA |  | Medical Eqpt | 5 | [10 empty] | 0
137 | MOR20029 | WARMER BLOOD |  | Medical Eqpt | 36 | [10 empty] | 0
138 | MOR20030 | WARMER BLOOD RAPID INFUSION |  | Medical Eqpt | 19 | [10 empty] | 0
139 | MOR30001 | ASPIRATION SYSTEM THROMBUS |  | Medical Eqpt | 4 | [10 empty] | 0
140 | MOR30013 | MONITOR CARDIAC OUTPUT THROUGH ESOPHAGUS |  | Medical Eqpt | 216 | [10 empty] | 0
141 | MOR30015 | PUMP BALLON |  | Medical Eqpt | 5 | [10 empty] | 0
142 | MPH40001 | REFRIGERATOR MEDICATION 700L |  | Medical Eqpt | 24 | [10 empty] | 0
143 | MPH40005 | REFRIGERATOR MEDICATION 160L |  | Medical Eqpt | 2 | [10 empty] | 0
144 | MPH50001 | CABINET MEDICATION |  | Medical Eqpt | 10 | [10 empty] | 0
145 | MPH50002 | CABINET NARCOTIC |  | Medical Eqpt | 2 | [10 empty] | 0
146 | MPH50005 | CART MEDICATION |  | Medical Eqpt | 33 | [10 empty] | 0
147 | MPH50006 | CART DISPENSING |  | Medical Eqpt | 6 | [10 empty] | 0
148 | MPH50007 | SAFE NARCOTIC LARGE |  | Medical Eqpt | 2 | [10 empty] | 0
149 | MPH50008 | SAFE NARCOTIC SMALL |  | Medical Eqpt | 4 | [10 empty] | 0
150 | MRH10006 | BALANCE APPARATUS |  | Medical Eqpt | 2 | [10 empty] | 0
151 | MSE10043 | GAUGE ROOM PRESSURE |  | Medical Eqpt | 22 | [10 empty] | 0
152 | MWD10002 | BED BURN |  | Medical Eqpt | 5 | [10 empty] | 0
153 | MWD10004 | BED ELECTRICAL PEDIATRIC |  | Medical Eqpt | 20 | [10 empty] | 0
154 | MWD10006 | BED INTENSIVE CARE |  | Medical Eqpt | 50 | [10 empty] | 0
155 | MWD10007 | BED INTENSIVE CARE PEDIATRIC |  | Medical Eqpt | 5 | [10 empty] | 0
156 | MWD10013 | CABINET BEDSIDE |  | Medical Eqpt | 71 | [10 empty] | 0
157 | MWD10016 | CRIB NEWBORN |  | Medical Eqpt | 7 | [10 empty] | 0
158 | MWD10021 | STRETCHER EMERGENCY |  | Medical Eqpt | 48 | [10 empty] | 0
159 | MWD10022 | STRETCHER RECOVERY |  | Medical Eqpt | 10 | [10 empty] | 0
160 | MWD10023 | STRETCHER TRANSPORT |  | Medical Eqpt | 26 | [10 empty] | 0
161 | MWD10024 | TABLE OVERBED |  | Medical Eqpt | 115 | [10 empty] | 0
162 | MWD20002 | BED INTENSIVE CARE BARIATRIC |  | Medical Eqpt | 1 | [10 empty] | 0
163 | NMG10012 | ‫FREEZER 21 BODY |  | Medical Eqpt | 1 | [10 empty] | 0
164 | NMG10028 | STRETCHER MORTUARY |  | Medical Eqpt | 1 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | MAL20002 | 0 | COMPRESSION DEVICE CHEST |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Configurable to 102 – 120 compressions per minute, fixed, or variable during use
2 | Pressure pad during ventilation
3 | Safety system controls
4 | No patient weight limitation
5 | Power Source: Battery – Rechargeable Lithium-ion Polymer (LiPo) with external power supply or car cable
6 | External Power supply: 220VAC, 60Hz
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | MAL20011 | 0 | AED SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Automatic External Defibrillator (AED)
2 | Waveform | biphasic
3 | Defibrillator Charge Hold Time | not less than 20 sec
4 | Energy Selecion | 120, 150, 200 J Adult - 50, 75, 85J Pediatric
5 | Charge Time | less than 10 sec
6 | Electrode with CPR sensor | Prefereed
7 | Automatic self test | Yes
8 | Defibrillator advisory | Yes
9 | Languages | Arabic or English
10 | Patient Impedance | 0 - 300 Ohms
11 | Display
12 | Display Format | LCD with CPR Feedback
13 | Depth Display | Yes
14 | Size | less than 4''
15 | ECG Viewing Time | specify
16 | Display sweep speed | 20 or 25 mm/sec
17 | Battery
18 | Capacity | 300 shock or 13 hours ECG monitoring
19 | Low Battery indication | not less than 30 discharge
20 | shelf life | 5 years
21 | Electrodes
22 | Adult with CPR sensor | Yes
23 | Pediatric | Yes
24 | Shelf life | ≥ 18 month
25 | General
26 | Size (HxWxD) cm | specify
27 | Weight KG | less than 4KGs
28 | Ingress Protection (IP) | water and dust proof, specify
29 | Carry Case | soft carry case
30 | OTHER SPECIFICATIONS | FDA, CE, ISO approved
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | MAL20013 | 0 | RESUSCITATION KIT AMBULATORY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | A portable manually operated resuscitator (Ambu bag) for adult use.
2 | To incorporate a pressure relieving valve and rubber bag of ~ 2000 ml
3 | One-way valve with 45cm H2O pop off with override.
4 | Tube for introducing oxygen: the oxygen concentration must be able to be as high as 55%.
5 | The resuscitator must be steam sterilizable
6 | Offer shall include one set of all available adult size masks
7 | Storage and carrying case shall be included
8 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | MAL20015 | 0 | VENTILATOR TRANSPORT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PATIENT TYPE | Adult, Pediatric
2 | Built-in Turbine ( Compressor ) | Yes
3 | CONTROLS
4 | Monitoring | Total expiratory minute volume, Spontaneous minute volume, Peak pressure, Mean pressure, Plateau pressure, Positive end expiratory pressure, Expiratory tidal volume, Inspiration time, Frequency, Spontaneous breathing frequency, Oxygen concentration
5 | Tidal volume, L | 50-2000 mL
6 | Breath rate, br/min | 0 - 60 or more
7 | Inspiratory time, sec | 0.2-10 SEC
8 | Inspiratory flow, L/min | 60 up to 100
9 | Pressure level, cm H2O | 0-60 CM H2O
10 | Pressure ramp | Specify
11 | Pressure support | 0 - 30 CM H20
12 | Trigger Mechanism | Flow or pressure or both, Specify.
13 | Sensitivity, cm H2O | Flow or pressure trigger, specify
14 | FiO2, Percentage | 21 % - 100 %
15 | I:E ratio | 1:4 to 3:1
16 | Adjustable PEEP either Internal or External, specify | 0-20 cm H2O
17 | OPERATING MODES
18 | Intermittent Positive Pressure Ventilation (VC-CMV) | Yes
19 | Assisted Intermittent Positive Pressure Ventilation (VC-AC) | Yes
20 | Synchronized Intermittent Mandatory Ventilation (VC-SIMV) | Yes
21 | Spontaneous ventilation (SpnCPAP) | Yes
22 | Pressure Support to assist spontaneously breathing patients (SpnCPAP+PS) | Yes
23 | Mask ventilation with leakage compensation in the modes SpnCPAP (SpnCPAP+PS)(NIV) & BIPAP | Yes
24 | Display Type | Colored , Built-in with the unit
25 | MONITOR PARAMETERS
26 | Pressure, PIP, MAP, PEEP. | Yes
27 | Volume, Tidal, Minute | Yes
28 | Inspiratory/Expiratory Time | Yes
29 | I:E Ratio | Yes
30 | Respiratory Rate | Yes
31 | PATIENT ALARMS
32 | FiO2, Low/High minute volume, Low Inspiratory pressure, High pressure, Loss of PEEP, Apnea, Inverse IE, High continuous pressure occlusion, High respiratory rate, Others. | Specify
33 | EQUIPMENT ALARMS
34 | Gas-supply failure, Power failure, Vent inoperative, Low battery, Self Diagnostic, Others. | Specify
35 | Silence | 2 minutes
36 | CONTROL TYPES | Dial, soft keys, toggle switch
37 | POWER, VAC | 220, 60Hz.
38 | Breathing circuit Adult & Pediatric | 3 set
39 | BUILT-IN BATTERY WITH CHARGER | 2 HOURS AT LEAST
40 | THE ABOVE UNIT IS COMPATIBLE WITH AMBULANCE CAR | Yes
41 | ALL ACCESSORIES NEEDED ,INCLUDING PATIENT BREATHING CIRCUIT LOUPES & FILTERS with each VENTILATOR | Yes
42 | CERTIFICATES | APPROVED FROM REGULATORY BODY
43 | LANGUAGE | ENGLISH
44 | PLUGS | BRITISH STANDARD
45 | UNIT OF MEASUREMENT | SI : INTERNATIONAL STANDARD
46 | MOBILE STAND | INCLUDED ON CASTORS WITH BASKET
47 | BED RAIL HANGER | INCLUDED
48 | HANDLE | YES
49 | WEIGHT | 7 Kg APPROX
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | MED10001 | 0 | BRONCHOSCOPE WORKSTATION VIDEO |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LCD or TFT MEDICAL MONITOR FULL (HD)
2 | The System shall Have the Highest – Resolution 24- 26 Flat Panel Monitor | Yes
3 | Color tones are completely calibrated for use with flexible endoscopes | Yes
4 | Adjustable image size | Yes
5 | Digital inputs: SDI / DVI-D. Analog inputs: RGBS / VGA / S-Video / Composite | Yes
6 | Widescreen aspect ratio . | specify
7 | DIGITAL VIDEO PROCESSOR (HD)
8 | HD signal output | Yes
9 | User interface | Control panel , keyboard
10 | Standard video outputs RGB, Y/C, Composite | Yes
11 | Electronic magnification | Specify
12 | Picture in picture display for any combination of endoscopic images | Yes
13 | Auto white balance | Yes
14 | Color adjustment function | Yes
15 | Light source integrated or separate | Xenon 300 watt
16 | BRONCHOSCOPE | Adult
17 | HD scope | Yes
18 | Diagnostic, Therapeutic | Yes
19 | Single channel | Yes
20 | Angle of view (°) | 120 degree approx.
21 | Focal range | 3–100 approx.
22 | Tip deflection (°) up/down | 180/130 approx.
23 | Insertion tube diameter (mm) | 180/130 approx.
24 | Instrument channel diameter (mm) | 3 approx.
25 | Available working length (mm) | 600mm
26 | Compatible with ESU | Yes
27 | Chip at the tip of scope | Yes
28 | BRONCHOSCOPE | Pediatric
29 | HD scope | Specify
30 | Angle of view | 120 degree
31 | Focal range | 3-100 approx.
32 | Insertion tube diameter | 4.0 mm
33 | Instrument channel diameter | 2 mm
34 | Available working length | 600mm
35 | Color video printer | Yes
36 | Compatible with processor | Yes
37 | Slim and compact design . | Yes
38 | USB host port . | Yes
39 | Convenient remote control capability 1 & 2 . | Yes
40 | 4 frame memory and 2 or 4-split image print mode . | Yes
41 | The printer can print one image while simultaneously storing additional images . | Yes
42 | TROLLEY | Yes
43 | Mobile base with lockable front casters. | Yes
44 | Shelf weight capacity, Kg. | Yes
45 | Isolation Transformer | Yes
46 | Central On/Off switch | Yes
47 | Outlet strip with circuit breaker | Yes
48 | Standard instrument Set (Itemized Price, compatible with Endoscopy system). | Yes
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | MED10006 | 0 | ENDOSCOPY WORKSTATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LCD or TFT Medical Monitor Full HD
2 | The System shall Have the Highest – Resolution 24 Flat Panel Monitor . | Yes
3 | Color tones are completely calibrated for use with flexible endoscopes . | Yes
4 | Adjustable image size . | Yes
5 | Multi modality display capability , including Picture in Picture (PIP) and Picture out Picture ( POP) with various image size ombination . | Specify
6 | 180 degree rotation display and mirror display modes . | Yes
7 | Digital inputs: SDI / DVI-D. Analog inputs: RGBS / VGA / S-Video / Composite. | Yes
8 | Widescreen aspect ratio . | Yes
9 | Video Processor
10 | Digital Color Video Processor utilizes Digital Signal Processing (DSP) technology | Yes
11 | which automatically adjusts and optimizes the color and illumination of the video | Yes
12 | Signal to provide maximum image clarity and resolution . | Yes
13 | White balance adjustment is possible using the white balance switch on the front panel . | Yes
14 | Standard color chart output can be display . | Yes
15 | Automatic gain control (AGC) . | Yes
16 | The Image contrast have three modes ( N,H,L) . | Yes
17 | Automatic Iris mode and Automatic focusing . | Yes
18 | Fine Patterns or edges in the endoscopies images can be enhanced electrically to increase the image sharpness . | Yes
19 | The Enhancement level can be selected from 4 levels ( off,1,2, and 3) . | Yes
20 | HDTV signal output . | Yes
21 | SDTV signal output . | Yes
22 | Endoscopic Light Source
23 | Separate or built in the controller . | Yes
24 | Day light illumination ( Xenon or LED ) . | Yes
25 | Automatically adjust light intensity to achieve ideal illumination of the gastrointestinal tract . | Yes
26 | Air feeding pump have 4 level ( off , low , mid and high ) . | Yes
27 | Emergency lamp . | Yes
28 | IMAGE MANAGEMENT SYSTEM | Option
29 | Computer based unit that records ,manages and archives digital images and videos of Endoscopy procedures . | Yes
30 | User friendly software designed specifically for Endoscopy purposes . | Yes
31 | Write multi – session and multi patient CDs/DVDs . | Yes
32 | Controllable via Keyboard , Mouse , foot switch and buttons on the scope . | Yes
33 | USB support for storage on file on USB drivers . | Yes
34 | Customizable Print-outs for the documented information . | Yes
35 | Dicom / HL7 | Yes
36 | Easy connect to thermal printer . | Specify
37 | Hard disc memory | Yes
38 | HD Gastro scope for Adult | QTY:2
39 | For Diagnostic and Therapeutic . | Yes
40 | Single channel Gastro scope . | Yes
41 | Bending Up/down (degrees) 210/90 approx. | Yes
42 | Bending Left/Right (degrees) 100/100 approx. | Yes
43 | Working Length (mm) 1050 mm approx. | Yes
44 | Field of View 140 degree . | Yes
45 | Direct of view forward . | Yes
46 | Outer Diameter 9.8mm . | Yes
47 | Instrument Channel 2.8 mm approx . | Yes
48 | Depth of View ( 3 – 100 ) mm approx . | Yes
49 | Standard Instrument Set . | Yes
50 | High frequency Compatibility . | Yes
51 | Gastro scope for Pediatric | Qty:1
52 | Field of View 140 degree approx. . | Yes
53 | Outer Diameter 8 mm approx. | Yes
54 | HD Colon scope for Adult | Qty:2
55 | Single channel Colon scope . | Yes
56 | Bending Up/down (degrees) 180/180 . | Yes
57 | Bending Left/Right (degrees) 160/160 . | Yes
58 | Working Length (mm) 1680 mm approx. | Yes
59 | Outer Diameter 13.2 mm approx. | Yes
60 | Instrument Channel 3.2 mm approx . | Yes
61 | Colon Scope for Pediatric | Qty:1
62 | Sigmoid scope for Adult | Qty:2
63 | Single channel Sigmoid scope . | Yes
64 | Working Length (mm) 700 mm approx. | Yes
65 | Outer Diameter 11.6 mm approx. | Yes
66 | Irrigation Unit
67 | Separate or built in the controller or light source . | Yes
68 | Special design for Endoscopy system . | Yes
69 | Light weight . | Yes
70 | Disposable kit . | Yes
71 | Pressure 0 - 48 psi . | Yes
72 | Easy to install on the trolley . | Yes
73 | Suction Unit
74 | fluid management system provides reusable/disposable suction jar/liner options. | Yes
75 | workstation mountable . | Yes
76 | hydrophobic microbial filter . | Yes
77 | variable suction with vacuum gauge . | Yes
78 | compact ergonomic design for easy operation . | Yes
79 | suitable for all endoscopic procedures . | Yes
80 | easy change autoclavable bottle cap/float assembly . | Yes
81 | CD/DVD | Option
82 | COLOR VIDEO Printer : | Yes
83 | Slim and compact design . | Yes
84 | USB host port . | Yes
85 | Analogue video inputs and outputs (RGB, Composite Video, S-Video) . | Specify
86 | Convenient remote control capability 1 & 2 . | Yes
87 | 4 frame memory and 2 or 4-split image print mode . | Yes
88 | The printer can print one image while simultaneously storing additional images . | Yes
89 | Trolley
90 | Compact, ergonomic trolley ideal for any endoscopic requirement . | Yes
91 | Rides on 4 antistatic dual wheels, 2 equipped locking brakes . | Yes
92 | Three or more fixed shelves . | Yes
93 | Cable Conduit vertical or horizontal . | Yes
94 | Keyboard tray . | Yes
95 | Video endoscopic holder . | Yes
96 | Pump holder | Specify
97 | Video endoscopic in take . | Yes
98 | Foot switch pedal holder . | Yes
99 | Irrigation rod with holder . | Yes
100 | Drawer with lock . | Yes
101 | Power box more than 7 sockets . | Yes
102 | Original trolley | Yes
103 | Scope storage cabinet | Yes
104 | for 8 scopes | Yes
105 | Removable plastic drip tray | Yes
106 | Hinged Glass doors | Specify
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | MGE10006 | 0 | CABINETS STORAGE INSTRUMENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Heavy duty, all-welded 20 gauge stainless steel construction
2 | Five removable, adjustable (1/2" increments) stainless steel shelves
3 | Full height, stainless steel doors with glass panes
4 | Convenient T-handle door latches with lock
5 | Adjustable leveling glides
6 | Dimensions: 77.25" H x 16" Deep x (24" W Single model) or (36 , 48" TwinWide model)
7 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | MGE10007 | 0 | CART CHART PATIENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | trolley free standing with wheels
2 | Capacity | 40 charts
3 | no. of shelfs | specify
4 | no. of drowers | 6 or more
5 | material | made of plastic
6 | high-quality swivel casters two locking | yes
7 | soft grip handles | yes
8 | stabilizing frame with bumber | yes
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | MGE10009 | 0 | CART CRASH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CRASH CART
2 | CARDIO-RESPIRATORY RESUSCITATION SYSTEM | Mobile cart with lockable drawers and defibrillator transport tray ( bracket )
3 | MATERIAL | HEAVY DUTY QULAITY TYPE
4 | ACCESSORIES
5 | Oxygen supply sys. | Yes, 1 cylinder
6 | Suction unit portable, battery. | Yes
7 | Oxygen flow meter | Yes
8 | Laryngoscope blades set | Yes, Adult/Pedia/Neonate
9 | Resuscitation bag | Yes with mask for adult/pedia/neonate
10 | ALL ACCESS NEEDED | SPECIFY
11 | Tourniquet | Yes, Velcro fastener
12 | Forceps | Yes, 2 sets
13 | Magill | Yes
14 | Artery | Yes
15 | Dressing | Yes
16 | Suture | Yes
17 | Penlight | Yes
18 | Sphygmomanometer ( aneroid ) | Yes, with cuffs, (adult, & pedia.)
19 | Infusion pole | Yes, Telescopic
20 | CART
21 | Wheels | 4 anti-static wheel, > 12.5 cm dia.
22 | Brakes | > 2
23 | Waste container | Yes
24 | Heart board | Yes
25 | Oxygen tank holder | Yes
26 | Medication tray | Yes
27 | Infusion pole receptacle | Yes
28 | Push handle | Yes
29 | Bumpers | Yes
30 | Top guard rail | Yes
31 | Drawers | Yes
32 | OTHER SPECIFICATION | Work surface and size shall accommodate the Defibrillator and all accessories mentioned.
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | MGE10011 | 0 | CART DRESSING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety standard | FDA approval or CE Marking
2 | Design | Heavy duty, compact design & high quality finishing
3 | CONSTRUCTION | Rectangular tube frame, Made from S/S or Equiv. that is corrosion resistance & easy to clean. Rugged construction to with stand mechanical shocks caused by movement over uneven surfaces
4 | COMPARTMENT | Please Specify .
5 | WORKING TRAY | pull-out & made from S/S
6 | DRAWERS | 5 utility drawers with partitions to store dressing instruments, gauze etc?.
7 | CUPBOARD | One included
8 | ADJUSTABLE SHELF | Please Specify
9 | WASTE RECEPTACLE | Removable with locking mechanism
10 | TRAY | Swivel mounted & Made from S/S or Equiv
11 | BUCKET WITH LID | Available within fixture & made from S/S or equiv.
12 | AMPOULE OPENER | Included
13 | CASTORS | 4-antistatic castors with 2 efficient braking mechanism
14 | Table Top surface | Not less than 900 mm x 600 mm approx. Made from S/S or Equiv.
15 | Table top height | Not less than 1500 mm approx.
16 | BUMPERS | Revolving Bumpers
17 | Push Handles | Full-width hand rail on both sides
18 | Guard rail | 3-sided made from S/S or equiv.
19 | Accessories & Options | Any available options shall be quoted separately
20 | Catalogues | Original/colored detailed catalogue stating all the above mentioned specifications should be submitted with offer
21 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | MGE20006 | 0 | CYLINDER O2 SIZE D WITH REGULATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | CYLINDER O2
2 | SIZE | D
3 | Regulator | Yes
4 | Humidifier & mask | Yes
5 | Connections | Yes
6 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | MGE20011 | 0 | FLOWMETER AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Oxygen flowmeter shall deliver a set flow rate for a variety of medical devices such as tents, masks, endotracheal tubes, nasal cannulae and catheters.
2 | The flowmeter shall operate at a flow rate of 0 to 15 L/min, with an accuracy of ± 0.5 L/min or 10%, whichever is larger.
3 | Accuracy of a back-pressure-compensated flowmeter should not be affected by partial obstruction of the outlet line.
4 | The flowmeter gauge should be float in tube or aneroid type
5 | The flowmeter should be structurally sound and able to support the weight of an “E” cylinder.
6 | Flowmeter performance should not be affected by temperature and humidity extremes encountered in operation, storage, and transport.
7 | Fittings attached to flowmeters shall allow connection only to oxygen sources.
8 | Fully opening the needle valve should provide a “flood” or “flush” rate at least several times the 15 L/min maximum calibrated level.
9 | The flowmeter should be electrically conductive from inlet to outlet to prevent arcing and possible fire from the accumulation of static electricity generated by nebulizers.
10 | The flowmeter should be able to withstand a 200-psi inlet pressure without damage in the event of a faulty regulator valve.
11 | The flowmeter shall be clearly labeled and color-coded to conform to existing standards. Each unit should be permanently marked with the manufacturer, model number, calibration conditions, and specific point on the float at which the readings
12 | should be m
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | MGE20014 | 0 | FLOWMETER O2 15 L/Min WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Flow Meter | Wall Mounted
2 | Flow Range | 0 - 15 L / Min
3 | Color Coded | specify
4 | Pressure Compensated | Yes
5 | Glass Envelop | Included with indicator ball
6 | Metal Nipple | Yes
7 | Probe | As Per Site
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | MGE20017 | 0 | FLOWMETER O2 WITH HUMIDIFIER WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Pressure compensated
2 | Flow tube | Impact resistant
3 | Flow indication tube | Single
4 | Lower scale | Expanded
5 | Flow range | Up to 15 L/min, approx
6 | Humidifier cup volume | Reusable
7 | Flowmeter compatibility to all hospital wall outlets | Yes
Company Stamp |  |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | MGE20019 | 0 | REGULATOR O2 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Medical oxygen regulator | Pin index
2 | Regular outlet pressure | 50 psi approx.
3 | Adjustable flow rate | 0 - 15 L/min approx.
4 | Pressure guage scale | 0 - 3000 psi
5 | Aluminum oxygen regulators. | Yes
6 | The offered regulator should be passed the ignition test | Yes, certificate should be supplied
7 | ASTM G175 | Yes
8 | Sintered bronze filter | Yes
Company Stamp |  |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | MGE20020 | 0 | REGULATOR SUCTION HIGH WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE(S) OF SUCTION | All common procedures and GI
2 | MOUNTING | Wall mounted
3 | FUNCTION | Continuous and intermittent
4 | ADJUSTABLE CYCLE | Independent off/on
5 | Increments | 3 - 30 sec
6 | VACUUM ADJUST LOCK | Yes
7 | VACUUM GAUGE | Digital or analog
8 | Pin stop | Yes
9 | Range, mm Hg | 0 - 300, full line
10 | Graduations, mm Hg | Five
11 | Disinfection | Specify
12 | Sterilization | Specify
13 | Safty trape | Yes
14 | Probe as per site | Yes
15 | Collection bottle | Disposable System
16 | Material of bottle | Poly carbonate
17 | OTHER SPECIFICATIONS | Color-coded gauge
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | MGE30007 | 0 | WHEELCHAIR PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Country Of origin | Specify
2 | Model | Specify
3 | Manufacture | Specify
4 | APPLICATION | To carry pediatric patient
5 | FRAME | Chrome plated heavy duty steel
6 | TYPE | Manual
7 | ARM REST | Yes
8 | SEAT WIDTH in inch | 18 x 17 approx.
9 | overall height in inch | 36 approx.
10 | overall length in inch | 38 approx.
11 | FOOT REST | Yes
12 | BRAKE | Yes
13 | SWING AWAY FRONT WHEEL | Yes
14 | REAR WHEEL | Yes
15 | FOLDABLE | Yes
16 | Castor diameter
17 | Front Casters diameter in inch | Specify
18 | Rear Casters diameter in inch | 20 approx.
19 | EASILY CLEANABLE SEAT & BACK REST | Yes
20 | WEIGHT | Approx 100 kg
21 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | MGE40001 | 0 | HAMPER LINEN DOUBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Double.
2 | Two stainless steel bag holders and two removable blocking rubber frames on which is possible to apply laundry bags. The rubber is 100? resistant, jointless, washable and disinfectable. | Yes
3 | On 4 swivelling solid rubber casters, ? 80 mm, complete with PVC non-marking rubber bumpers. | Yes
4 | Dimensions: 740 x 470 x 950 h mm | Yes.
5 | stainless steel tubular frame; it is hygienically safe and can be washed and | Yes.
6 | CE approved. | Yes.
Company Stamp |  |  |  | Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | MGE40002 | 0 | HAMPER LINEN SINGLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Epoxy Painted Heavy Tubular Frame Linen Hamper
2 | BAG | Yes
3 | CASTORS | Yes
4 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | MGE40004 | 0 | STAND IV |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FRAME | Stainless steel 304 preferred
2 | TYPE | Telescopic
3 | HOOKS | 4 - 5 Approx
4 | HEIGHT | Adjustable
5 | BASE | Mobile heavy base to withstand the hanged Items weights
6 | Castors | Qty : 5
7 | CAPACITY | 5 kg per hook
Company Stamp |  |  |  | Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | MGE40006 | 0 | STOOL FOOT 2 STEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FOOT STOOL - two step
2 | MATERIAL | Specify the material type
3 | DIMENSION | ALMOST ( 14x19x6 ) WxLxH
4 | TOP SURFACE | SLIP RESISTANCE
5 | WEIGHT - ALONE | Minimum 2.5 Kg
6 | WEIGHT - STACKED | Maximum 200 Kg
7 | OTHERS | EASY CLEAN - INFECTION CONTROL REGULATION
Company Stamp |  |  |  | Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | MGE40007 | 0 | TABLE INSTRUMENT LARGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | STAINLESS STEEL INSTRUMENT TABLE
2 | DIMENSION | SPECIFY
3 | CASTORS | 4 MEDIUM SIZE CASTORS
4 | NUMBER OF SHELVES
5 | TABLE TOP | FLUSHED
6 | OTHER SPECS | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | MGE40009 | 0 | TROLLEY CYLINDER SIZE F/G WITH STABILISERS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Powder coated steel frame
2 | Compatible with F & G cylinders, max diameter 200mm, height 1350mm
3 | One piece construction
4 | Drop over chains secure the cylinder
5 | Stabilising castors at rear of trolley
6 | 1600mm anti-static wheels, 100mm swivel castors
7 | Dimensions (w x d x h): 500 x 700 x 1160mm
8 | Weight: 14 Kg
Company Stamp |  |  |  | Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | MGE40010 | 0 | TRUCK CYLINDER SIZE E |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CYLINDER TRUCK - E SIZE MEDICAL GAS | Yes
2 | Polypropylene strap and cinch buckle | Yes
3 | Fully welded construction | Yes
4 | Quality powder paint finish for long service life | Yes
5 | Cylinder capacity | 12
6 | Height in cm | 100 approx.
7 | Depth in cm | 67 approx.
8 | Width in cm | 50 approx.
9 | Weight in kg | Specify
10 | Wheel external diameter | Specify
11 | No. of Wheel | 4
Company Stamp |  |  |  | Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | MGE50001 | 0 | DEFIBRILLATOR UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DEFIBRILLATOR ENERGY SELECTION, Joules. | Yes
2 | Internal with Lead | 5-50 Joules
3 | External Biphasic , Adult | 2-200 Joules
4 | External Biphasic , Pediatric / Neonatal | 2-200Joules
5 | PADDLE CONTROLS | Charge, Discharge
6 | Synchronizer | Yes
7 | Pediatric Paddles / Neonatal | Yes
8 | Disposable PADDLE | Yes , 25 pairs , validity two years
9 | ECG Monitor | Yes
10 | Type | LCD colored
11 | Screen display size, | 5 to 7 Inch
12 | Sweep Speed, mm/sec. | 20-25, Other specify.
13 | Lead Configuration | Yes 3 - 5 -12 leads
14 | THROUGH THE PADDLE MONITORING | Yes
15 | HR Display | Yes
16 | SPO2 | Yes
17 | HR Alarms | Yes
18 | Frequency Response, Hz. | Specify
19 | Lead fault indicator | Yes
20 | EXTERNAL PACEMAKER | YES
21 | Pacing Mode | Demand, fixed rate, Specify.
22 | Pacing Rate, ppm | 50-150, Specify
23 | Output current, mA | 0-140 or wider
24 | Pulse width, m sec. | > 20
25 | ECG RECORDER | Yes
26 | Paper Speed, mm/sec. | 25, Others specify.
27 | Auto/Manual Print | Specify
28 | Annotation | Time, date, lead, gain, heart rate, operating mode, other specify.
29 | BATTERY/LINE POWER OPERATION | Both
30 | Integral or Removable | Specify
31 | Operating Time, Hr. | > OR = Two Hours continuous ECG monitoring - OR > / = 20 discharges.
32 | Charging Time, Hr. | Specify
33 | TO BE PLACED ON CRASH CART | YES
34 | AC POWER | 220Volt - 60Hz
35 | OTHER SPECIFICATIONS | FDA or CE, ISO approved.
36 | Options Itemized Prices | Yes
37 | CPR quality monitoring | Yes
38 | Trace freeze | specify
39 | CO2 MODULE | YES
Company Stamp |  |  |  | Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | MGE50002 | 0 | DIAGNOSTIC SET WALL MOUNTED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SPHYGMOMANOMETER ANEROID | INCLUDED
2 | Reusable tips set | Four pcs.
3 | Disposble tips set | Yes
4 | Handle 2.5 volts | Yes
5 | Magnification | 2.5 X approximately
6 | OPHTHALMOSCOPE | Yes
7 | Diopter corrections, D | -25 to + 40 approx.
8 | Filters | Red free
9 | Diagnostic beams | Two
10 | BATTERY OPERATED | Yes, rechargeable
11 | HARD CASE | Yes
12 | CHARGER | Yes
13 | Infrared thermometer | Yes
14 | Sphygmomanometer | Yes
15 | OTHER SPECIFICATIONS
16 | INFLATION PRESSURE, mm Hg | 300 mm Hg
17 | CUFF SIZES | Adult,Pediateric, Velcro cuff
18 | SYSTOLIC, mm Hg | 0-240 approx.
19 | DIASTOLIC, mm Hg | 0-130 approx.
20 | THERMOMETER DIGITAL
21 | CONFIGURATION | Handheld with plug in probe, table top/wall mounted holder
22 | CONTINUOUS CAPABLE | Yes
23 | OPERATIONAL MODE | Predective/monitor
24 | TEMPERATURE
25 | Screen Display type | LCD , TFT or other, Specify.
26 | Range, C | SPECIFY
27 | Accuracy, C | SPECIFY
28 | Response time, sec | SPECIFY
29 | CONTROLS | On/Off and automatic shut off
30 | ALARMS/INDICATORS
31 | Visual | Final temp., Lo & Hi temp., calibration.
32 | Audible | Final temp., Lo & Hi temp., calibration.
33 | PROBE ASSEMBLY
34 | Number | 2 Oral/rectal
35 | Type Material | Stainless steel or plastic, specify.
36 | Storage well | Yes
37 | PROBE COVERS
38 | Type | Oral, rectal or both together, specify
39 | Material | low density polyethelyne
40 | Removal | Eject button.
41 | Storage on unit | Yes
42 | POWER SUPPLY | Battery Operated
43 | Battery Type | Rechargeable or single use battery, Specify.
44 | Low Battery Indicator | Visual or Audible, Specify.
45 | Operating time | > 300 hrs
46 | For rechargeable battery, charger power requirement | 220Volt, 60 Hz.
47 | OTOSCOPE | Yes
Company Stamp |  |  |  | Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | MGE50005 | 0 | ECG 12 LEAD WITH CART |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile
2 | Cart | Yes
3 | LEADS
4 | Lead switching | Automatic/Manual
5 | Sensitivity, mm/mV | 5, 10, 20., other specify.
6 | Calibration signal | Automatic
7 | FREQUENCY RANGE, HZ
8 | Diagnostic | 0.05-150
9 | Filtered | Notch, EMG, Baseline, Wander, filters
10 | INPUT IMPEDANCE,
11 | Mega ohms | one hundred
12 | CMRR @ 60 Hz, dB | 100-140 approx
13 | Leads-off indicator | Yes
14 | RECORDER
15 | Channels | twelve
16 | Recording | YES A4 PAPER
17 | Lead marker | Automatic
18 | Timing marker | Yes
19 | Event marker | Yes
20 | Chart speed, mm/sec | 5, 10, 25, 50.
21 | LCD MONITOR 10 OR MORE | Yes ADJSTABLE
22 | No. of traces | 3, 6, 12
23 | PREVIEW SCREEN | Yes
24 | No. Waveforms stored | Min. 30
25 | ECG transmission | Yes
26 | Interpretation | Yes
27 | ECG measurements | Yes
28 | Auxiliary output | Yes
29 | Auxiliary input | Yes
30 | DEFIBRILLATOR SYRORNAIZ | YES
31 | SINGLE AVREGING | YSE
32 | Overload protection | Yes
33 | BATTERY OPERATION | Yes
34 | Battery type | Built in rechargeable
35 | Operating time, hr | Specify
36 | ACCESSORIES | All standard accessories
37 | PRINTER | Built-in
38 | A4 paper | Yes
39 | SAFTY AND PERFORMANCE | IEC 60601-2-51
40 | POWER REQUIREMENTS | 220V, 60Hz
41 | OTHER SPECIFICATIONS | Spare patient cable, chest electrodes, clamp electrodes for limb leads, with ICU Analysis. SOFT WARE Patient cable holder.FDA, CE, ISO, Approved. HL7
Company Stamp |  |  |  | Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | MGE50006 | 0 | GLUCOMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor-controlled, hand held glucose meter to measure blood glucose levels.
2 | Should be IVD marked (In Vitro Diagnostic Medical Device).
3 | Measuring range: 20-500mg/dl Approx.
4 | Must give accurate results within Approx. 15 seconds.
5 | Must store at least the 20 results.
6 | Must be supplied with 200 strips.
7 | Must be supplied with a virtually pain-free blood sampling tool (lancing device).
8 | Specify the shelf life of strips (vial closed, after opening the vial)
9 | Must be supplied with long-life batteries.
10 | Unit must have an automatic shut off to protect battery life
11 | Unit must have low battery warning
12 | Must have a case for storage.
13 | Compliance with standards & legislation:
14 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | MGE50009 | 0 | LARYNGOSCOPE ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Laryngoscope Set Adult | Yes
2 | BLADES | Yes
3 | Adult (Different Sizes) | 4 Different Sizes 1,2,3,4
4 | ILLUMINATION | With Incorporated Fiber Optic Light Carrier inside the Blade (Xenon Light)
5 | Rechargeable Battery Shell | With High Power LED Technology, more than 50,000 lux, Lithium-ion batteries
6 | Handle Sleeve | Yes
7 | Charging Unit | For 2 Rechargeable batteries with power adaptor (110-240 VAC, 60 Hz)
8 | Bag for All Laryngoscopes | Splash - Protected
9 | Other Specifications | FDA, ISO Approved.
Company Stamp |  |  |  | Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | MGE50010 | 0 | LARYNGOSCOPE ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE, CONVENTIONAL OR FIBER OPTIC | MacIntosh
2 | BLADES
3 | Adult | 4 Different
4 | Pediatric | 4 Different
5 | ILLUMINATION | xenon
6 | Spare lamps | 5
7 | CHARGER, | Yes
8 | Rechargeable Battery | Yes
9 | Power, vac, hz. | 220, 60
10 | Wall mounted | to be improvised
11 | ALL ACCESSORIES (ITEMIZEDPRICE) | Yes
12 | OTHER SPECIFICATION | FDA, CE, ISO Approved
Company Stamp |  |  |  | Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | MGE50011 | 0 | LARYNGOSCOPE NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LARYNGOSCOPE NEONATE
2 | SMALL HANDLE | YES
3 | BLADES | NEONATE BLADE SET sizes(1,0,00)
4 | BLADES TYPE | straight
5 | Easy to clean | yes
6 | LAMP | 2 EXTRA LAMP
7 | BATTERY | RECHARGABLE
8 | CASE | INCLUDED
9 | OTHER SPECS | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | MGE50012 | 0 | LARYNGOSCOPE PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Laryngoscope Set : | Neonatal & Infant
2 | BLADES
3 | Pediatric (Different Sizes) | 2 Different Sizes (0,0 / 0,1 / 1,1 / 1,2 )
4 | Pediatric (Different Sizes) | 2 Different Sizes 0,1,2,3
5 | ILLUMINATION | With Incorporated Fiber Optic Light Carrier inside the Blade (Xenon Light)
6 | Spare lamps | 6
7 | Rechargeable Battery Shell | With High Power LED Technology, more than 50,000 lux, Lithium-ion batteries
8 | Reduction Sleeve | For rechargeable batteries
9 | Handle Sleeve | Yes
10 | Charging Unit | For 2 Rechargeable batterie (110-240 VAC, 60 Hz)
11 | Bag for All Laryngoscopes | Splash - Protected
12 | Other Specifications | FDA, ISO Approved.
Company Stamp |  |  |  | Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | MGE50014 | 0 | LIGHT EXAM MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mobile examination light on sturdy and stable base
2 | On/Off switch, easily accessible on the light stand or light head
3 | Stability and precise positioning in every desired position.
4 | Weight balancing system in the rotating joint
5 | Smooth surfaces for perfect hygiene
6 | Light intensity: approx. 40,000 lux at 80 cm
7 | Quick lamp replacement
8 | Color temperature ≥ 3500K
9 | Temperature increase in the head area ≤ 2° C
10 | Angle of rotation: unlimited
11 | Minimal heat dissipation
12 | 3 white LED lights
13 | Sterilizable hand grip
14 | Main and battery-operate
15 | Shadowless light
16 | Swiveling castors with locking feature
17 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | MGE50017 | 0 | MATTRESS AIR MED RISK |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA approved and/or CE Marked | Yes, Please specify
2 | Type | Powered Air Mattress
3 | Indications | Medium Risk pressure relief therapy surface (Prevention and Treatment)
4 | Bed Type | Adult Bed
5 | Alternating Therapy | Yes
6 | Cycle Time | Yes Specify
7 | Static therapy mode | Yes
8 | Adjustable wieght capcity | Yes
9 | Integrated Pressure sensor | Yes, which monitors internal pressure 24 hours.
10 | Minimum of 16 air cylinders individual replaceable cells | Yes, please specify
11 | Mattress Dimensions (LxWxD) cm | 200 x 90 x 18 (Inflated)
12 | Cover Material , Water Proof , Bacterial Barrier , Stain Resistant , Air breathable& moist , Soft for friction elmination , Shear resistant , Washable | yes
13 | Maximum weight capacity | 150Kg
14 | Manual CPR function from the mattress | Yes
15 | Transport mode | Yes , specify
16 | Alarms , Visual , Audiable | Yes
17 | Low Pressure Alert | Yes
18 | Voltage | 220 – 240 V 50/60 Hz
19 | Flexible Hangers in the rear side | Yes
20 | Whisper quite | Yes
Company Stamp |  |  |  | Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | MGE50018 | 0 | MATTRESS ANTIDECUBITUS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | alternating pressure
2 | AIR COMPRESSOR | Yes (Low noise)
3 | ILLUMINATED ON/OFF SWITCH | Yes
4 | AIR PRODUCTION VOLUME | 0 - 10 Lpm approx
5 | PRESSURE RANGE | 20 ? 50 mmHg approx
6 | ADJUSTABLE CYCLE TIME | 0 ? 30 mins, approx
7 | MODE OF OPERATION | Yes, static and alternating
8 | MAXIMUM PATIENT WEIGHT | 200 kgs or more
9 | MATTRESS SPECIFICITY | Hypo allergenic & easily cleanable
10 | MATTRESS DIMENSION IN CMS | 90(w) X 200(l) approx.
11 | POWER SUPPLY | 220 v, 60 Hz
12 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | MGE50020 | 0 | NEBULIZER ULTRASONIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | NEBULIZER
2 | Warm-up time, min
3 | Mist temperature, 0C | 33-35 aprox.
4 | Mist output, mL/min | 0.2-3.15
5 | Noise level, dbA,
6 | max @ 1 m distance
7 | O2 conc settings, % | 40-100 % approx.
8 | HEATING UNIT
9 | Location | Tube
10 | Type | Processor controlled/heater block
11 | Power-on indicator | Yes
12 | Temperature control | Yes
13 | Safety shutoff | Yes
14 | Line voltage, VAC | 220-240 V, 60 Hz
15 | RESERVOIR
16 | Refillable | YES
17 | Usable volume, mL | >200
18 | Water ml | >200
19 | Saline (0.45/0.9%) | >150
20 | STERILIZATION METHOD
21 | Nebulizer head and Heating unit | Chemical and autoclaveable
22 | Reservoir | Chemical and autoclaveable
23 | OTHER SPECIFICATIONS | ALL ACCESS. NEEDED TABLE TOP
Company Stamp |  |  |  | Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | MGE50021 | 0 | OTOSCOPE TABLETOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Fiber optics Otoscope | Yes.
2 | 3.5V Bulb type (Xenon) | Yes.
3 | Magnification more than 2.5x | Yes.
4 | Insuflation port | Yes.
5 | Insufflation Bulb | Yes.
6 | instrument housing | Specify.
7 | Smart design | Yes.
8 | Diameter of ear-reusable specula | 2.5, 3, 4, and 5mm approx.
9 | Disposable Tip for adult and pediatric | 200 tips for each approx.
10 | Three Spare bulb | Yes.
11 | Battery rechargeable | LI-ION
12 | Charger table top | Yes.
13 | Hard case | Specify.
14 | Power Supply 220 VAC / 60 HZ | Yes.
15 | FDA or CE approved | Yes.
Company Stamp |  |  |  | Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | MGE50024 | 0 | OXIMETER PULSE HAND HELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Portable / Stand-alone/transport/HAND HELD FOR ADULT/PEDIATRIC/INFANT
2 | DISPLAYS | SpO2, pulse rate, pulse strength and/or signal, low battery, Volumes, alarm limits.
3 | Type | LCD/ with backlit LED SHOWING PARAMETERS & READINGS
4 | Sp02 RANGE, % | 0-100
5 | Accuracy, ,% | < ± 2 %
6 | PULSE RATE, bpm | 30-240
7 | PERFUSION INDEX AND/OR SIGNAL STRENGTH INDICATOR | YES
8 | RESPONSE TIME, sec | Pulse to Pulse or time ( Specify)
9 | SETTLING TIME | Pulses or time ( Specify)
10 | START-UP TIME, sec or pulses
11 | ALARMS:
12 | Audiovisual | Yes
13 | Visual | Yes, Low battery
14 | ALARM OVERRIDE | Yes
15 | Reactivation method | Yes, after 2 mins.
16 | Volume control | Yes
17 | SELF-TEST MODE | Yes
18 | Memory: up to 72 hours of data storage | YES
19 | PROBE TYPES | YES REUSABLE (1 ADULT, 1 PEDIATRIC, AND 1 NEONATE)
20 | Patient range | Adult / Neonate/ Pediatric
21 | Cable length, m | Not less than1.5 m
22 | BATTERY BACKUP :
23 | Rechargeable battery | Yes, Specify the type
24 | Low battery notice | Yes
25 | Rechargeable time, hr | Specify
26 | Battery life, hrs | Specify
27 | The unit should operate on batteries as well as AC electricity, with automatic battery charging while in use. | YES
28 | The unit shall operate on batteries | YES
29 | battery type | Specify
30 | VAC | 220 V
31 | CURRENT | 13 A
32 | FREQUENCY | 60 Hz.
33 | PLUG TYPE | 3 Pin British
34 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
35 | IMPORTANT NOTE: | Please specify whether accessories are standard or optional, otherwise they will be considered standard.
36 | MDMA | Yes
Company Stamp |  |  |  | Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | MGE50026 | 0 | PUMP FEEDING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | For Liquid feeding
2 | CONFIGURATION | IV pole on stand with 4-5 castors
3 | PUMP TYPE | Rotary peristaltic or Linear peristaltic, specify
4 | OPERATING MODE | Continuous, volume/time selection and Bolus application
5 | FLOW RANGE IN ML/HOUR | 0.5- 600 or more
6 | DISPLAY | LCD
7 | ACCURACY | 5 - 10 %, specify
8 | OCCLUSION PRESSURE | 15 - 25 psi, specify
9 | OCCLUSION ALARM | Yes, audible & visual
10 | LOW BATTERY ALARM | Yes, visual
11 | FLOW ERROR ALARM | Yes, audible & visual
12 | SET NOT FIXED ALARM | Yes, audible & visual
13 | DOSE COMPLETE ALARM | Yes, audible & visual
14 | PARAMETERS RETAINED IN MEMORY | Rate, dose & volume infused
15 | FUNCTIONAL BACK UP MEMORY | Yes, at least 24 hours or more
16 | FEEDING ADMINISTRATION SET | Yes, open system
17 | SELF DIAGNOSTIC TEST | Yes
18 | BUILT IN RECHARGEABLE BATTERY | Yes
19 | POWER SUPPLY | 220 V, 60 Hz
20 | REGULATORY COMPLIANCE | FDA, CE approval
21 | OTHER SPECIFICATION | the device must be able to connect with HIS system
22 | PLUG | BRITISH
23 | LANGUAGE | ENGLISH
24 | NG feeding bags, 1000 ml, with tubes compatible with feeding pumps, cap, side port | QTY : 1000 itemized price also
25 | provide free pump with 1000 sets feeding | must quote
Company Stamp |  |  |  | Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | MGE50027 | 0 | PUMP INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PROGRAMMING PATINET ID / NURSE ID ON THE PUMP DURING ADMINISTRATION | YES ( PLEASE SPECIFY )
2 | QUICK START OF INFUSION | YES
3 | FLOW RANGE ML/HR | 0.1 - 999 ML/HR
4 | ACCURACY | 0.05
5 | DRUG/DOSE CALC | YES
6 | VTBI, ML | 0.1 TO 9999
7 | MAX PRESSURE, PSI | SPECIFY MAX
8 | PUMPING MECHANISM | SPECIFY MECHANISM
9 | IV SET | YES
10 | FREE-FLOW PROTECTION | YES
11 | MANUAL AND PROGRAMMABLE BOLUS | PLEASE SPECIFY BOLUS RATE
12 | ALARMS & INDICATORS
13 | EMPTY BAG | YES
14 | FLOW ERROR | YES
15 | DOOR OPEN | YES
16 | CONTINUOUS PRESSURE MONITORING PSI (ON THE FRONT SCREEN ALL THE TIME) | SPECIFY
17 | DRUG DOSE CALCULATION | ML/H, µG/H, µG/MIN, µG/KG/H, µG/KG/MIN, MG/H, MG/MIN,MG/KG/MIN, MG/KG/H, IU/H, IU/MIN, IU/KG/H, IU/KG/MIN.
18 | FLEXIBLE PRESSURE ALARM SETTINGS | IN LINE PRESSER
19 | HIGH PRESSURE/OCCLUSION | YES
20 | PRE- OCCLUSIONALARM | YES
21 | CIRCUIT MALFUNCTION | YES
22 | INFUSION NEAR END | YES
23 | INFUSION COMPLETE | YES
24 | LOW BATTERY | YES
25 | DEPLETED BATTERY | YES
26 | USER PROMPT | YES
27 | AUDIBLE ALARM | YES
28 | VOLUME CONTROL | YES
29 | MOMENTARY SILENCE | YES
30 | PPM DUE DATA REMINDER | YES
31 | FLUID RESISTANT | YES
32 | NURSE CALL JACK | YES
33 | DATA LOG | YES
34 | EVENT LOG | YES
35 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED)
36 | NUMBER OF EVENTS | 1500
37 | PRINT OUT | YES
38 | DISPLAY | LED AND LCD WITH COLOR CABLITIES IS A PLUS & BIG SCREEN INTERFACE IS ALSO A PLUS
39 | INFUSED VOLUME DISPLAY | YES , DIGITAL
40 | WEIGHT WITH BATTERY KG | SPECIFY WIGHT
41 | TIME /DATE | YES
42 | POWER SOURCE, VAC | 220VOLT / 60 HZ.
43 | POWER SUPPLY | INTERNAL OR EXTERNAL (IF EXTERNAL MUST BE LIFETIME WARRANTY)
44 | POWER CABLE | DURABLE POWER CABLE
45 | BATTERY | YES, RECHARGEABLE BUILT IN
46 | BATTERY | NI-MH
47 | OPERTING TIME , HR | SPECIFY HOURS @ 10 ML/HR
48 | RECHARGE TIME , HR | PREFERRED LOWER CHARGING HOURS
49 | POLE MOUNTING | PREFERRED
50 | OTHER SPECIFICATIONS | FDA / CE
51 | EVENTS STORAGE | YES ( ALARMS, SETTINGS, ERRORS, AMOUNT INFUSED).
52 | STACKABLE IN THE SAME DOCKING STATION | INFUSION PUMPS / SYRINGE PUMPS / ENTRAL FEEDING PUMPS.
53 | DOSE ERROR REDUCTION SYSTEM (SMART TECHNOLOGY) | YES
54 | LIBRARY SIZE | YES CONFIGURABLE
55 | NO. OF DRUG ENTITIES/CARE AREA | YES PLEASE SPECIFY
56 | WIRELESS CONNECTIVITY | UPGRADABLE
57 | TIME RETAINED | FIRST IN FIRST OUT
58 | ACCESSORIES FOR THE PUMP | CLOSED SYSTEM
59 | ALL SETS HAVE SELF -SEALING NEEDLE FREE INJECTION SITES | YES
60 | ALL SETS SHOULD BE LATEX DEHP & PVC FREE | YES
61 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
62 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
63 | ACCESSORIES FOR INFUSION PUMP | CLOSE SYSTEM
64 | THE VENDOR MUST PROVIDE A FIXED PRICE FOR THE ALL ABOVE MENTIONED ACCESSORIES VALID FOR 3 YEARS | YES
65 | PUMP HL7 READY | YES ( PLEASE SPECIFY )
66 | PROVIDE CONSUMABLE DEAL ( PUMPS FREE AND PRICE ON SETS ) | OPTIONAL DEAL SECTOR MIGHT CHOOSE
67 | FIXATION OF THE PUMPS BY SIMPLE CLICK IN MECHANISM INTO THE SYSTEM. | YES
68 | ALL PUMPS CAN BE REMOVED INDIVIDUALLY WITHOUT DISTURBING THE FUNCTIONALITY AND THE DATA TRANSFER OF THE REMAINING PUMPS. | YES
69 | THE WORKSTATION SHOULD BE MOUNTABLE , WALL UNIT (FIXATION TO WALL RAILS OF DIFFERENT MANUFACTURER).MOVABLE POLE-BASED UNIT.
70 | SMART ALARM SYSTEM BY AUDIBLE, VISUAL AND DISPLAYED MESSAGE. | OCCLUSION, END OF INFUSION, EMPTY BATTERY, NO RATE SET, DEFECT ON PUMP. PRE-ALARMS: LOW BATTERY, NEAR-END OF INFUSION , PLEASE SPECIFY
71 | MOBILE STAND TO FIX THE PUMP ON INSTEAD OF PENDANT OR WALL FOR SPECIAL SITE NEEDS | ITEMIZED PRICE OPTIONAL
72 | PRIMARY INFUSION SET (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 120 SETS WITH EACH PUMP ITEMIZED
73 | LIGHT SENSITIVE PRIMARY INFUSION SET WITH NO INJECTION PORT FOR NEONATAL TPN (CLOSED) (PLEASE ITEMIZED PRICE FIXED FOR 5 YEARS) | 40 SETS WITH EACH PUMP ITEMIZED
Company Stamp |  |  |  | Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | MGE50028 | 0 | PUMP INSULINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Pump size: 3.13D x 2H x 0.84W in. with insulin cartridge
2 | Pump weight: 4.05 oz. with battery and full reservoir
3 | Battery operated: Rechargeable lithium polymer
4 | Reservoir capacity: 480-unit cartridge
5 | Infusion Set: Compatible with all standard Luer-lock infusion sets
6 | Basal Range: From 0.5 to 15 units per hour in 0.001-unit increments
7 | Bolus Range: From 0.5 to 60 units in 0.01-unit increments
8 | Rechargeable battery with micro USB.
The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
Company Stamp |  |  |  | Signature

## ITEM 42
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
42 | MGE50029 | 0 | PUMP SYRINGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled with digital LCD alphanumeric display of parameters and alarms
2 | Variable rate ranging from 0.1 to 500 mL/hr or better, with 0.1 mL/hr increments
3 | 3 % accuracy or better
4 | Variable volume-to-be-infused from 1 to 1,000 mL or similar
5 | Digitally displayed parameters to include:
6 | Infusion rate
7 | Battery / AC operation
8 | Running indicator
9 | Alarming condition when active, with indication of alarm type or code
10 | Back pressure monitor / indicator
11 | Event history
12 | Capability to accept different syringe types and sizes with automatic syringe detection and identification
13 | Syringe compatibility and auto detection shall include but not be limited to all sizes of the following (3 to 60 mL):
14 | BD
15 | Terumo
16 | Monoject
17 | Braun
18 | Fresenius
19 | Variable bolus rate
20 | Up to maximum flow rate
21 | Bolus infused volume indicator during bolus activation
22 | Protected access
23 | Audiovisual alarms shall include but not be limited to the following:
24 | Syringe installation and integrity (detection)
25 | Line disconnection (sudden drop in back pressure)
26 | Occlusion pressure pre-alarm
27 | Occlusion pressure
28 | Near end of infusion alarm
29 | End of infusion
30 | Volume limit pre-alarm
31 | Volume limit
32 | KVO (1 ml/hr; if other, specify)
33 | Low battery pre-alarm
34 | Discharged battery
35 | Internal malfunction
36 | Syringe unlocked
37 | Plunger disengaged
38 | Empty syringe
39 | Data log capability and data port for data transmission, display and printing. Any required software for such function shall be included.
40 | Logged data to include ( ≥1 year) :
41 | Settings
42 | Alarms
43 | Errors
44 | Key presses
45 | Error codes
46 | Amount infused
47 | Safety features shall include but not be limited to:
48 | Self test at start-up
49 | Nurse call interfacing capability
50 | Splash proof design
51 | Auto priming
52 | Adjustable alarm volume. No permanent silencing shall be possible.
53 | Keypad lock
54 | Impossibility to improperly install infusion set
55 | Free flow prevention system
56 | Last parameter setting retention
57 | IV stand mounting accessory shall be included
58 | Battery autonomy of 3 hrs or more when fully charged. specify:
59 | Battery type and characteristics (voltage and current capacity)
60 | Autonomy at 10 mL/hr
61 | Recharging time from depleted to 90%
62 | Compliance with standards & legislation:
63 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
64 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
65 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
66 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
67 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 43
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
43 | MGE50035 | 0 | SPHYGMOMANOMETER ANEROID MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | METHOD | ANEROID
2 | INFLATION PRESSURE, mm Hg | 300 mm Hg
3 | CUFF SIZES | Adult, Pediatric, Velcro cuff
4 | SYSTOLIC, mm Hg | 0-240 approx.
5 | DIASTOLIC, mm Hg | 0-130 approx.
6 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Signature

## ITEM 44
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
44 | MGE50036 | 0 | SPHYGMOMANOMETER ANEROID WALL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | METHOD | ANEROID
2 | INFLATION PRESSURE, mm Hg | 300 mm Hg
3 | CUFF SIZES | Adult, Pediatric, Velcro cuff
4 | SYSTOLIC, mm Hg | 0-240 approx.
5 | DIASTOLIC, mm Hg | 0-130 approx.
6 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Signature

## ITEM 45
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
45 | MGE50037 | 0 | SPHYGMOMANOMETER ELECTRONIC VITAL SIGN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DIASTOLIC PRESSURE | Yes
2 | PULSE RATE | Yes
3 | SPO2 | Included
4 | ALARMS, AUDIBLE & VISIBLE | Yes
5 | POWER SUPPLY | 220v, 60 Hz
6 | BATTERY | Yes, Built in re-chargeable batteries with integral charger.
7 | ACCESSORIES | Basket for cuffs
8 | OTHER SPECIFICATION | yes.
9 | SPO2 Probe | For Adult, Pediatric and Neonatal ( Disposable Box )
10 | Temp module | Included
11 | Oral Temp. Probe | Included
12 | Rectal Temp. Probe | Optioanl
Company Stamp |  |  |  | Signature

## ITEM 46
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
46 | MGE50038 | 0 | STETHOSCOPE ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Double head
2 | MATERIAL
3 | EAR TUBE | Stainless steel
4 | CHEST PIECE | Stainless steel
5 | BINAURAL
6 | ADJUSTMENT | To angle of ear
7 | CHEST PIECE | Adult and pediateric & Neonatal
8 | ADDITIONAL DIAPHARGM | 1 EACH
9 | EAR TIPS, TYPE | COMFORT SEALING
10 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 47
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
47 | MGE50042 | 0 | THERMOMETER DIGITAL DOCKING STATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Digital Thermometer safe for all ages
2 | Shall be for oral, underarm, or rectal use.
3 | Should use disposable probes for measurement
4 | Should have Fast and accurate readings
5 | Temperature accuracy: 0.1°C
6 | Automatic shut-off
7 | Large display for reading
8 | Should signal after temperature taken
9 | Safe to use, with no glass and mercury-free
10 | Should have an auto memory to show a minimum of last 10 temperatures taken
11 | Should include rechargeable batteries
12 | Docking station with Rechargeable Battery Pack
13 | Designed with plastics compatible with common medical-grade cleaning products
14 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 48
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
48 | MGE50043 | 0 | THERMOMETER DIGITAL HANDHELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Digital Thermometer safe for all ages and shall be for oral, underarm, or rectal use.
2 | Should use disposable probes for measurement
3 | Fast, accurate readings
4 | Temperature accuracy: 0.1°C
5 | Automatic shut-off feature
6 | Large LCD display for reading
7 | Should signal after temperature taken
8 | Safe to use, with no glass and mercury-free
9 | Should have an auto memory to show last 10 temperatures taken
10 | Should include rechargeable batteries with charging desk unit
11 | Unit should operate >1,000 uses on full battery
12 | Unit should have a self test capability
13 | Unit shall have visual indicators
14 | Final temperature stops flashing
15 | low-battery
16 | low and high temperatures
17 | Should have a suitable case
18 | Compliance with standards & legislation:
19 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
20 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 49
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
49 | MGE50044 | 0 | THERMOMETER DIGITAL WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Digital Thermometer safe for all ages
2 | Shall be for oral, underarm, or rectal use.
3 | Should use disposable probes for measurement
4 | Fast and accurate readings
5 | Readings in Fahrenheit and Celsius
6 | TEMPERATURE:
7 | °F RANGE 90 - 110
8 | °C RANGE 30 - 44
9 | Temperature accuracy: 0.1°C
10 | Automatic shut-off
11 | Large LCD display
12 | Should signal after temperature taken
13 | Safe to use, with no glass and mercury-free
14 | Should have an auto memory to show a minimum of last 10 temperatures taken
15 | Should include rechargeable batteries with charging desk unit
16 | Should have a holder for wall mounting
17 | Visual / Audible alarm indicators
18 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 50
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
50 | MGE50045 | 0 | THERMOMETER INFRARED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Non-contact thermometer uses infrared technology to take temperature
2 | Effective distance: 5-8 cm
3 | Range: 35-43˚C
4 | Accuracy: ± 0.2 ˚C
5 | Resolution: 1˚C
6 | Battery operated
7 | Indication for low battery power
8 | Auto recall of last reading
9 | LCD back lit display
10 | Automatic shut off after 10-15 second
11 | Approx. Dimension: 150 H x 40 W x 40 D mm
12 | Approx. Weight: 100 g
13 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 51
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
51 | MGE50046 | 0 | THERMOMETER TYMPANIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Quick-response ear thermometer features one-touch operation
2 | Non invasive reading from the tympanic membrane
3 | Safe for use on patients of any age
4 | Measurement Time: 2 to 3 seconds
5 | Displayed Temperature Range : 20.0 °C to 42.2 °C
6 | Display Type: Backlit liquid crystal display, four digits plus special icons
7 | Readings in Fahrenheit and Celsius
8 | High display resolution
9 | Operating Temperature : 50 °F to 104 °F (10.0 °C to 40.0 °C)
10 | Automatic Power Off
11 | Battery operated
12 | Approx. Weight : 100 g without batteries
13 | Dimensions : 15 H cm × 5 W cm × 4 D cm
14 | Standard accessories should be listed in details with part number and quantities.
15 | The offered equipment shall have an approved international certificate (CE, FDA, TUV, etc.)
Company Stamp |  |  |  | Signature

## ITEM 52
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
52 | MGE50048 | 0 | TRANSILLUMINATOR VEIN HANDHELD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Wave Length | Specify
2 | Projection of Image | Real Time Image
3 | Penetration Depth | Up to 10 mm deep
4 | Brightness | 6 lumens
5 | Weight | Specify
6 | Control Panel | Soft Touch
7 | Detection Method | Specify
8 | Modes:
9 | Universal | Yes
10 | Fine Detail | Yes
11 | Inverse | Yes
12 | Re-Size | Yes
13 | Active Vein Imaging Navigation | Yes
14 | Non-Invasive Procedure | Yes
15 | No-Heat Generation | Yes
16 | No-Cross Contamination | Yes
17 | No-Disposable & Consumable parts required | Yes
18 | Range of Use | All Ages ( Pre-mature to Adult )
19 | Skin Pigmentation | All Skin Tone
20 | Completely Hands Free for the Caregiver with Eye On Patient (EOP) | Optional hands-free S-Mount (Stability Mount) for Eyes On Patient technique
21 | Power Source | Battery or AC power
22 | Battery Life | Up to 2 hours continuous run time (4 hours with additional battery)
23 | Certifications | FDA , CE.
Company Stamp |  |  |  | Signature

## ITEM 53
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
53 | MGE50058 | 0 | OTO OPTHALMOSCOPE WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Wall Mounted
2 | TECHNICAL SPECIFICATIONS
3 | Lightweight | Yes
4 | Line Powered with transformer | Yes
5 | Handle Power Line Cord Length | More than 2.5 m
6 | Retinoscope Halogen or Xenon or LED Lamp (LED Preferred) | 3.5 V
7 | Effect | Plane and concave mirrors
8 | Polarization Filter | Integral
9 | Rotation control | Yes
10 | Ophthalmoscope Halogen or Xenon or LED Lamp voltage | 3.5 V
11 | Curved End Probes | Yes
12 | Place to store in units | Yes
13 | Lamp Spare as chosen | 1 For each
14 | PHYSICAL SPECIFICATIONS
15 | Dimensions | Specify
16 | Weight | Specify
17 | LATEST MODEL TO BE INSTALLED AS PER THE AVAILABILTIY FROM MANUFACTURER | Yes
18 | INSTALLATION & PRE - INSTALLATION, IF REQUIRED, SHOULD BE DONE BY THE COMPANY, Like civil, electrical, plumbing works etc….) | Yes, attach separate scope with details
19 | HOSPITAL LINE POWER
20 | AC Voltage / Phase | 220 V / Single Phase
21 | Current | 13 A for power socket
22 | Frequency | 60 Hz.
23 | Plug Type | 3 Pin British
24 | UNIT COMPLETE WITH FULL ACCESSORIES | Disposable / Reusable
Company Stamp |  |  |  | Signature

## ITEM 54
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
54 | MGE60001 | 0 | SCALE CHAIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Clinical, chair
2 | TYPE | Electronic
3 | MEASUREMENT UNITS
4 | Weight | Metric
5 | Height | Metric
6 | CAPACITY | 250 kg. Approx.
7 | ACCURACY | 0.1 g
8 | CALIBRATION ADJUST | Yes
9 | Power | 220V 60Hz
10 | TARE WEIGHT ADJUST | Yes
11 | OTHER SPECIFICATIONS. | Chargeable Battery with charger
Company Stamp |  |  |  | Signature

## ITEM 55
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
55 | MGE60002 | 0 | SCALE INFANT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Electronic
2 | CONFIGURATION | Platform with integral tray
3 | Tray shape | Contoured
4 | Tray material | Molded plastic
5 | WEIGHT UNITS | Metric
6 | DISPLAY | LCD
7 | CAPACITY | 20 kg
8 | ACCURACY | 2 g
9 | CALIBRATION ADJUST | Yes
10 | TARE WEIGHT ADJUST | Yes
11 | LINE POWER, VAC | 220V 60 Hz
12 | BATTERIES | Yes
13 | Low-battery indicator | Yes
14 | Tray | Yes
15 | other specifications
Company Stamp |  |  |  | Signature

## ITEM 56
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
56 | MGE60003 | 0 | SCALE PATIENT WITH HEIGHT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Floor standing
2 | Application | For weight & Height measurement
3 | Result are easy to read on LCD display | 7 high-contrast LCD
4 | Scale Type | Digital
5 | Weight measurement | precise probe
6 | Height measurement | ultrasonic sensor
7 | Height measurement unit | Metric
8 | Weight measurement unit | KG and bound
9 | Body composition analysis | Specify.
10 | Capacity | 200kg approx.
11 | Thermal printer | Specify.
12 | Platform size | 250 X 300mm approx.
13 | Calibration Adjustment | Yes
14 | Connectivity Ports | Serial, USB, and Ethernet
15 | Transport castors for mobile use Qty: 2 | Yes.
16 | Accuracy | Height±0.1cm Weight±0.1kg
17 | Wheels | Yes
18 | Other Specification
Company Stamp |  |  |  | Signature

## ITEM 57
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
57 | MGE70001 | 0 | BAG AMBO ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | A portable manually operated Ambu bag for adult and pediatric.
2 | shall have a pressure relieving valve and rubber bag (each suitable for the proposed age group)
3 | shall have One-way valves with 45cm H2O pop off with override
4 | shall have Tube for introducing oxygen: the oxygen concentration must be able to be as high as 45%
5 | shall be steam sterilizable
6 | Storage and carrying case shall be included
7 | Compliance with standards & legislation:
8 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
9 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 58
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
58 | MGE70003 | 0 | RESUSCITATOR KIT ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Portable / ambulatory
2 | TYPE
3 | Exhaled air/manual | Manual
4 | Reusable/disposable Mask | Reusable
5 | Adult & pediatric neonatil Mask | Sizes 0,1,2,3 for Adult and pediatric and neonatil
6 | TRANSPARENT MASK | Yes
7 | Cuff type | Open flange
8 | Isolation mechanism | Transparent valve
9 | SUPPLEMNTL O2 INLET | Yes
10 | 15 mm/22 mm ADAPTER | Yes
11 | MAX STROKE VOL, mL | Specify
12 | MAXIMUM CYCLING
13 | RATE, cycles/min | Specify
14 | PRESSURE LIMITING | Yes
15 | STERILIZATION | Autoclavable
16 | Storage box | Yes
17 | Suction unit | Yes
18 | Airways | Sizes 0,1,2, & 3
19 | Laryngoscope | Adult, pediatric and infant size blades, battery included
20 | Bag | Tranparent folding bag with intake valve , extension tube
21 | OTHER SPECIFICATIONS | SILECON TYPE بواقع 20 كبار + 10 اطفال + 10 مواليد
Company Stamp |  |  |  | Signature

## ITEM 59
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
59 | MGE80001 | 0 | SUCTION UNIT ELECTRICAL MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Portable aspirator (suction) device used in or outside the hospital.
2 | VACUUM SYSTEM | ELECTRONIC
3 | PRESSURE RANGE,
4 | GUAGE , mm Hg | 0-700 approx
5 | Indicator type | Specify
6 | Safety cutoff pressure | Yes
7 | Airflow, L/min | Approx 75 L/MIN
8 | COLLECTION SYSTEM
9 | Container | Poly carbonate
10 | Number
11 | Capacity, L | 3 lit.
12 | Overflow protection | Audio & visual
13 | TUBING
14 | Number of inlets | 1
15 | Length, m (ft) | Specify
16 | All Cannula’s reusable and autoclavable. | Yes
17 | FOOT SWITCH | OPTIONAL ITEMIZED
18 | MOBILE CART WITH 5 WHEELS 2 LOCKABLE | INCLUDED
19 | LINE POWER, | 220 Volt – 60 Hz.
20 | DC VOLTAGE
21 | SHOULD BE COMPATIBLE WITH AMBULANCE CAR | Yes
22 | NON RETURN VALVE | The unit must be equipped with the back flow prevent or, for the prevention of infection.
23 | OTHER SPECIFICATIONS
24 | The unit should be AC and rechargeable battery-powered, with (45) minutes minimum battery operation, and built-in charger. | YES
25 | The unit should be portable with a handle, and lightweight approximately 5kg or less. | YES
26 | overflow safety valve, and a sealed bacterial or hydrophobic filter | YES
27 | The unit should be simple to use, operate, and maintain. | yes
28 | The unit should be well constructed with durable materials to withstand typical abuse and cleaning. | yes
29 | The unit should be easy to clean, disinfect, and/or sterilize, as appropriate. | YES
Company Stamp |  |  |  | Signature

## ITEM 60
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
60 | MHC10052 | 0 | WHEELCHAIR STANDARD SIZE 24 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Aluminium wheelchair | Yes
2 | Powder coating finish black | Yes
3 | Foldable backrest, aluminium folding mechanism | Yes
4 | Flip back armrest | Yes
5 | Detachable footrest, heel loop | Yes
6 | Nylon black upholstery | Yes
7 | Front castor 8"x1" solid tire | Yes
8 | Steel brake | Yes
9 | Rear wheel 24" x 1 3/8" PU tire , , | Yes
10 | Quick release rear axle | Yes
11 | aluminium hand rim | Yes
12 | brakes | Yes
13 | Anti tipper | Yes
14 | 5 Cm seat cushion | Yes
15 | Capacity 150 Kg | Yes
16 | Wieght 14 Kg | Yes
Company Stamp |  |  |  | Signature

## ITEM 61
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
61 | MHC30008 | 0 | CHAIR BATH LARGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Bath chair with back and arms
2 | Comfortable padded arms provide support and stability
3 | Arms are easily removable
4 | Durable and lightweight
5 | made of blow molded plastic and rust proof aluminum
6 | Skid resistant rubber tips have a wide base and drainage holes
7 | Minimum weight capacity: 150kg
Company Stamp |  |  |  | Signature

## ITEM 62
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
62 | MHD10001 | 0 | CRRT UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CRRT Machine
2 | Continuous Hemodialysis (CRRT) with the following treatment modalities
3 | SCUF: Slow Continuous Ultra-Filtration | YES
4 | CVVH: Continuous Veno-Venous Hemofiltration | YES
5 | CVVHD: Continuous Veno-Venous Hemodialysis | YES
6 | CVVHDF: Continuous Veno-Venous Hemodiafiltration | YES
7 | HV-CVVH: High Volume Continuous Veno-Venous Hemofiltration | YES
8 | TPE: Therapeutic Plasma Exchange for adult and pediatric applications including consumables availability in KSA market | YES
9 | HP: Hemoperfusion program and consumables ( coated charcoal filters) | YES
10 | Pediatric CRRT / Neonatal CRRT | YES
11 | The Unit should be able to change therapy (between SCUF, CVVH, CVVHD & CVVHDF) during treatment without line disconnection/addition nor changing the set | YES
12 | The Unit should use one treatment set which can perform all CRRT therapies (SCUF, CVVH, CVVHD & CVVHDF) | YES
13 | The Unit should contain at least the following 5 operating & software controlled pumps: | YES
14 | Blood Pump | YES
15 | Substitution Pre Pump | YES
16 | Substitution Post Pump | YES
17 | Dialysate Pump | YES
18 | Drain Pump | YES
19 | The Unit should contain Heparin (syringe) pump with the ability to perform both Continuous & Bolus delivery types | YES
20 | The Unit should have the Flexibility of performing Pre or Post / Pre and Post fluid substitution at the same time with all CRRT therapies (SCUF & CVVH & CVVHD and CVVHDF) | YES
21 | The Unit must also contain the following integrated parts:
22 | Deareation chamber | YES
23 | Air Detector | YES
24 | Clamp system to stop blood return when detecting air. | YES
25 | Blood leak detector. | YES
26 | Pressure pods for measuring:
27 | Arterial line pressure ( -250 to +200 mmHg) | YES
28 | Venous line pressure ( 0 to + 300 mmHg) | YES
29 | Pre filter pressure ( 0 to +450 mmHg) | YES
30 | Drain line pressure ( -350 to +350 mmHg) | YES
31 | Fluid Control units (Scales): | YES
32 | At least 4 independent scales & scale range from 0 to 11kg at least | YES
33 | Measuring principle: | Gravimetric
34 | Screen with the following specifications:
35 | On-screen service menu preferably via touch screen control | YES
36 | Screen size not less than 10.4” Color, TFT-LCD | YES
37 | User Interface with graphical treatment parameters (Curves), Continuous TMP and filter pressure drop monitoring
38 | The unit preferably should have the ability to recognize the type of the set, allows traceability & Automatic set parameters setting to avoid any human mistakes | YES
39 | Anti electrostatic device (ASD) to avoid ECG interferences | YES
40 | The unit must have citrate anti-coagulation module with the capability of performing Citrate Anticoagulation in all CRRT therapies (CVVH & CVVHD and CVVHDF) | YES
41 | The unit should have heating system to heat the blood or fluid to maintain patient temperature | YES
42 | The unit should have back up battery, not less than 15 min | YES
43 | Additional specifications :
44 | Flexibility of extracorporeal circuit flow-path management | YES
45 | Combined simultaneous Pre & Post fluid Replacement | YES
46 | To be used for citrate anticoagulation
47 | The unit should have Fully Pre-connected complete treatment set to allow fast installation & avoid human mistakes | YES
48 | The unit should have Pump protection faceplates to protect both operator and pumps and ensure safety operation | YES
49 | The unit preferably shall have the ability to record all the treatment sessions & events on the Unit for almost complete one week. It can be loaded on computer to build analysis & statistics | YES
50 | The unit should be able to be upgradeable for any future applications. | YES
Company Stamp |  |  |  | Signature

## ITEM 63
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
63 | MHD10004 | 0 | DIALYSIS HAEMOFILTRATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Haemodialysis unit to be used with chronic patients Adult and Pediatric and Neonatal in the kidney unit, and shall be capable of performing the following procedures: | YES
2 | Haemodialysis (HD) | YES
3 | Isolated UF | YES
4 | Sodium profiling & Bicarbonate Profiling and UF profiling | YES
5 | Unit shall have One blood pumps , and provide a continuous range of flow from 50 up to 500 ml/min. | YES
6 | Unit should have Single Needle System ( SN/SP). | YES
7 | Unit shall be connected to a centralised water treatment plant supply or a portable reverse Osmosis machine. | YES
8 | The unit shall have an integrated heparin syringe pump able to supply a continuous flow | YES
9 | range from 0.5 up to 10 ml/hour and give measurable stat doses. | YES
10 | Unit shall be equipped with BP monitor to measure systolic & diastolic & mean and heart rate, that can be controlled within the software of the unit and the measurement intervals Shall be adjustable. | YES
11 | BP cuffs must be provided in sizes for pediatric up to extra large adult, which are compatible with the standard of cuff used by .
12 | Unit shall use powder / dry bicarbonate, and provide a connection to all electrolyte solutions Requested. | YES
13 | Unit shall be able to perform a true sodium profiling without affecting minor electrolytes. | YES
14 | Unit should have ionic clearance measurement device which allow the operator to enter the Target KT/V and give a forecast based on the current clearance parameters. | YES
15 | Unit shall be provided with large display : specify size provided | YES
16 | All parameters has to be clearly displayed providing operator with an easy interact with. | YES
17 | Unit shall be capable to store dialysis records. | YES
18 | Unit extracorporeal circuit shall be equipped with blood sensor that will trigger in case of blood leakage rate exceeds 0.50 ml/min | YES
19 | Unit shall have a sophisticated alarming mechanism for all machine’s parameters, such as conductivity, temperature, blood volume, blood leakage, water supply flow and pressure etc, | YES
20 | the unit shall block the patient circuit in case of a potential risk alarm is triggered.
21 | Machine shall be provided with alarming even if the power is cut or power supply fails, the alarm shall be audiovisual and the machine shall go to a safe condition and with certain High risk alarms, clamp the venous return line. | YES
22 | Machine shall be equipped with back-up battery that can operate the machine for at least 15 - 30 minutes in case of power failure | YES
23 | The blood pump should be equipped with emergency hand crank to return the blood to the patient in case of power failure | YES
24 | Unit shall have procedure timer on the front panel; the timer shall show the elapsed time Since the procedure starts or remaining time till finish. | YES
25 | Unit shall be easy to operate and shall have intuitive user interface. | YES
26 | Unit shall have integrated spillage proof key pad | YES
27 | Unit should be equipped with the following sensors for the blood circuit | YES
28 | Dialysate Delivery | YES
29 | Comfort Control ˚C : ( 35 – 39 ). | YES
30 | Conductivity Range : 13 – 15 mS/cm | YES
31 | UF Removal Rate : 0.1 – 3 L/h. | YES
32 | Temp. Alarm Limits : 34 - 40 | YES
33 | Unit should be equipped with the following sensors for the blood circuit
34 | Artrial pressure sensor, range ( -350 to at least +150 mmHg) | YES
35 | Venous pressure sensor (-100 to at least 350 mmHg) | YES
36 | Ultrasonic air sensor for the dripping chamber | YES
37 | A blood volume monitor to measure the change in blood volume DBv to allow early detection of volume depletion. | Preferred
38 | A mechanism to stop the blood returned to patient in case of alarm. | YES
39 | The dialysis machines have ability to provide adjustable dialysis fluid flow rates i.e. bypass (off) and up to 700 ml/min : Specify your range & Steps. | YES
40 | Unit should provide Ultrapure Conventional dialysis by using built in Ultra dialyzer filter for helping achieve sterile substitution and disinfection fluids. | YES
41 | The machines blood pump being easily adjustable by operator to provide blood flow rates Down to 50ml/min via pediatric blood lines. | YES
42 | Unit should have chemical and heat disinfections cycles (preferably open system) with temperature range 90-95 C degrees in all internal parts of the machine to destroy any organism at this temperature | YES
43 | Unit shall be equipped with audiovisual alarm, alarm should not be defeated indefinitely, And the volume shouldn’t be adjustable below the hearing level. | YES
44 | Unit shall have service routines for checking various machine parameters. | YES
45 | Unit shall have service menu for monitoring and controlling parameters and to perform calibrations and service for the various machine’s functions | YES
46 | Unit shall be compatible with all consumables stated. | YES
47 | Unit shall have integrated network connectivity and serial communication protocols. | YES
48 | Unit shall have a safety class I. | YES
49 | Vendor machine be compatible with hospital information system and connectable to Lab / dialysis/ dr. clinic. | YES
50 | Vendor shall conduct comprehensive training for all operators; training shall be conducted by application specialist approved by the manufacture. | YES
51 | Mechanism that can be unlocked by machine operator. | YES
Company Stamp |  |  |  | Signature

## ITEM 64
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
64 | MHD10005 | 0 | DIALYSIS HAEMOFILTRATION PORTABLE WITH RO |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Haemodialysis unit to be used with chronic patients in the kidney unit, and shall be capable of performing the following procedures: | YES
2 | Haemodialysis | YES
3 | Isolated UF | YES
4 | Sodium profiling & Bicarbonate Profiling and UF profiling | YES
5 | Unit shall have One blood pumps , and provide a continuous range of flow from 50 up to 500 ml/min. | YES
6 | Unit should have Single Needle System ( SN/SP). | YES
7 | Unit shall be connected to a centralised water treatment plant supply or a portable reverse Osmosis machine. | YES
8 | The unit shall have an integrated heparin syringe pump able to supply a continuous flow | YES
9 | range from 0.5 up to 10 ml/hour and give measurable stat doses. | YES
10 | Unit shall be equipped with BP monitor to measure systolic & diastolic & mean and heart rate, that can be controlled within the software of the unit and the measurement intervals Shall be adjustable. BP cuffs must be provided in sizes for pediatric | YES
11 | up to
12 | Unit shall use powder bicarbonate, and provide a connection to all electrolyte solutions Requested. | YES
13 | Unit shall be able to perform a true sodium profiling without affecting minor electrolytes. | YES
14 | Unit should have ionic clearance measurement device which allow the operator to enter the Target KT/V and give a forecast based on the current clearance parameters. | YES
15 | Unit shall be provided with large display : specify size provided | YES
16 | All parameters has to be clearly displayed providing operator with an easy interact with. | YES
17 | Unit shall be capable of automatically adjusting UF & Na levels to prevent IDH episodes : Specify the IDH prevention mechanism. | YES
18 | Unit shall be capable to store dialysis records. | YES
19 | Unit extracorporeal circuit shall be equipped with blood sensor that will trigger in case of blood leakage rate exceeds 0.50 ml/min | YES
20 | Unit shall have a sophisticated alarming mechanism for all machine’s parameters, such as conductivity, temperature, blood volume, blood leakage, water supply flow and pressure etc, the unit shall block the patient circuit in case of a potential ris | YES
21 | k alarm
22 | Machine shall be provided with alarming even if the power is cut or power supply fails, the alarm shall be audiovisual and the machine shall go to a safe condition and with certain High risk alarms, clamp the venous return line. | YES
23 | Machine shall be equipped with back-up battery that can operate the machine for at least 15 - 30 minutes in case of power failure | YES
24 | The blood pump should be equipped with emergency hand crank to return the blood to the patient in case of power failure | YES
25 | Unit shall have procedure timer on the front panel; the timer shall show the elapsed time Since the procedure starts or remaining time till finish. | YES
26 | Unit shall be easy to operate and shall have intuitive user interface. | YES
27 | Unit shall have integrated spillage proof key pad | YES
28 | Unit should be equipped with the following sensors for the blood circuit | YES
29 | Dialysate Delivery | YES
30 | Comfort Control ˚C : ( 35 – 39 ). | YES
31 | Conductivity Range : 13 – 15 mS/cm | YES
32 | UF Removal Rate : 0.1 – 3 L/h. | YES
33 | Temp. Alarm Limits : 34 - 40 | YES
34 | Unit should be equipped with the following sensors for the blood circuit
35 | Artrial pressure sensor, range ( -350 to at least +150 mmHg) | YES
36 | Venous pressure sensor (-100 to at least 350 mmHg) | YES
37 | Ultrasonic air sensor for the dripping chamber | YES
38 | A blood volume monitor to measure the change in blood volume DBv to allow early detection of volume depletion. : Specify | YES
39 | A mechanism to stop the blood returned to patient in case of alarm. | YES
40 | Preferably Access flow meter to show patient access flow rate (Qa in ml/min). | YES
41 | The dialysis machines have ability to provide adjustable dialysis fluid flow rates i.e. bypass (off) and up to 700 ml/min : Specify your range & Steps. | YES
42 | Unit should provide Ultrapure Conventional dialysis ( preferably sterile ) by using built in Ultra dialyzer filter. | YES
43 | The machines blood pump being easily adjustable by operator to provide blood flow rates Down to 50ml/min via pediatric blood lines. | YES
44 | Unit should have chemical and heat disinfections cycles (preferably open system) with temperature range 90-95 C degrees in all internal parts of the machine to destroy any organism at this temperature | YES
45 | Unit shall be equipped with audiovisual alarm, alarm should not be defeated indefinitely, And the volume shouldn’t be adjustable below the hearing level. | YES
46 | Unit shall have service routines for checking various machine parameters. | YES
47 | Unit shall have service menu for monitoring and controlling parameters and to perform calibrations and service for the various machine’s functions | YES
48 | Unit shall be compatible with all consumables stated. | YES
49 | Unit shall have integrated network connectivity and serial communication protocols. | YES
50 | Unit shall have a safety class I. | YES
51 | Vendor machine be compatible with hospital information system and connectable to Lab / dialysis/ dr. clinic. | YES
52 | Vendor shall conduct comprehensive training for all operators; training shall be conducted by application specialist approved by the manufacture. | YES
53 | Mechanism that can be unlocked by machine operator. | YES
54 | Vendor must provide an itemized price for the following : | YES
55 | Qty : 1000 Blood Line Tubes Sets | YES
56 | Qty : 1000 Cartridges Bicarbonate 650 g. | YES
57 | PORTABLE RO WATER | YES
58 | WATER SAVING FUNCTION | Yes
59 | PROGRAMMABLE REMINDERS FOR CLEANING, FILLING OF SALT TABLETS CONTAINER IN THE SOFTENER | Yes
60 | PRODUCT WATER CONDUCTIVITY DISPLAY | Yes
61 | INLET WATER CONDUCTIVITY DISPLAY | Yes
62 | REJECTION RATE % DISPLAY | Yes
63 | DATE, TIME, TIME SINCE LAST DISINFECTION AND TOTAL RUN TIME DISPLAY | Yes
64 | AUO /MANUAL START UP FOR HEAT DISINFECTION | Yes
65 | CUSTOMIZED CLEANING PROGRAM FOR DIFFERENT NEEDS | Yes
66 | MACHINE REJECTION RATE IS MORE THAN 95% FOR TOTAL DISSOLVED SALTS | Yes
67 | MACHINE REJECTION RATE IS MORE THAN 99% FOR BACTERIA AND PYROGEN | Yes
68 | PC PRESET AND SOFTWARE UPGRADING | Yes
69 | PROGRAMMABLE TUBING FLUSH DURING STANDBY TO PREVENT BACTERIA GROWTH | Yes
70 | PC LOGGING ATO MONITOR THE FLOW DIAGRAM DURING OPERATION FOR TROUBLESHOORING | Yes
71 | PRODUCT WATER
72 | out put water minimum 1.11 l/min | Yes
73 | total dissolved salt more than 96% | Specify
74 | bacteria and pyrogens more than 99% | Specify
75 | FEED WATER SUPPLY
76 | input : 3.0l/min | Specify
77 | pressure : 1 - 8 bar | Specify
78 | TDS below than 1500mg/l | Specify
79 | chlorine total below 0.1mg/l | Specify
80 | TEMPERATURE RANGE : 1-100 C | Specify
81 | Size | Specify
82 | PERFERABLY CONTAIN INTERNAL RESERVIOR | Specify
83 | POWER SUPPLY 220 VAC / 60 HZ | Yes
84 | FDA and CE approved | Yes
Company Stamp |  |  |  | Signature

## ITEM 65
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
65 | MIC10003 | 0 | ANALYZER BLOOD GAS POC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA approval or CE marked | yes
2 | Type | Blood System Analyzer POC
3 | Applications | Measure Blood Parameters at patient bed side; Ph,pCO2,pO2,Na+,K+,Caa++,Lac,Glu,Hct, Crea, and CL-
4 | Calculated Parameters at patient side; cHgb,Bicharbonate,cTCO2,BE(ecf),BE(b),cSO2,eGFR,eGFR-a, | Yes
5 | Utilze up to date user interface technology | Yes, PDA
6 | Blood parameters can be measured using Test cards, no need for reagents | Yes
7 | Test Card Bar Coded | Yes
8 | Measurement can be done within 30 sec or less, without need to priming or calibration | Yes, specify
9 | Lactate measurement can be done without reagent “fluid” | Yes
10 | Test cards can be stored at room temp and does not require fridge for storage | Yes
11 | Possibility to connect multi beds into on PDA | Yes
12 | Connectivity of multi bed to one central SW for results documentations | Yes
13 | Compact size | Yes, approx “20X6X3” cm
14 | Power | 220V/60Hz
Company Stamp |  |  |  | Signature

## ITEM 66
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
66 | MIC10006 | 0 | DRAINAGE UNIT CHEST |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Quick and easy set up | yes
2 | suction regulation | yes
3 | Suction Control Chamber | yes
4 | detection and monitoring of patient air leaks | yes
5 | Instant patient pressure verification | yes
6 | High Negativity Float Valve | yes
7 | filter protection | yes
8 | High Negative Pressure Protection | yes
9 | positive pressure relief valve | yes
10 | Autotransfusion Capabilities | yes
11 | In-line fluid removal. No needle necessary | yes
12 | Patient Tube Slide Clamp | yes
13 | Collection Chamber | yes
14 | Hanger | yes
15 | Power requirments | 220v/60hz
16 | accessories | all needed accessories
17 | Other specifications | FDA approved or CE Marked
Company Stamp |  |  |  | Signature

## ITEM 67
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
67 | MIC10009 | 0 | INTUBATION DIFFICULT VIDEO ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | For documentation, teaching, monitoring and difficult direct tracheal intubation
2 | TYPE | Mobile
3 | TECHNICAL SPECIFICATIONS
4 | Imaging
5 | - Method | CMOS
6 | - Image display | LCD, high resolution (Touch screen preferred)
7 | - Size, inches (all sizes) | Specify
8 | Anti-fogging Mechanism | Yes
9 | Integrated high resolution camera | Yes
10 | Color Monitor Viewing | Yes
11 | Video Output/Recording Features | Yes
12 | Blade Angulation | 60° (approx)
13 | Field of View | Specify
14 | Viewing Direction | Direct view
15 | Illumination Method | Scope tip LED light
16 | PHYSICAL SPECIFICATIONS
17 | Dimensions | Specify
18 | Weight | Specify
19 | Accessories
20 | Adult, Blades (disposable) | Yes
21 | - Patient weight | Specify
22 | - Blade length (tip to handle) | Specify
23 | - Blade Thickness (max) | Specify
24 | - Blade width (max) | Specify
25 | - Camera | Colored
26 | - Slim blade profile | Yes
27 | Adult, Size 3 | Qty : 200 pcs / unit
28 | Adult, Size 4 | Qty : 200 pcs / unit
29 | Intlock Blade
30 | - Disposable | Yes
31 | - Suction channel | Incorporated
32 | Mobile Stand | Yes
33 | - Wheel & casters | Yes
34 | - Adjustable height stand | Yes
35 | - Cradle | Specify
36 | - Cobalt Cradle or hanger | Specify
37 | Power supply | Rechargeable Battery
38 | - Type | Specify
39 | - Battery life, hrs | ≥ 4
40 | Consumables | At least 3 months supply
41 | HIS Compatible. | Yes / Specify
42 | LATEST MODEL TO BE INSTALLED AS PER THE AVAILABILTIY FROM MANUFACTURER | Yes
43 | INSTALLATION & PRE - INSTALLATION, IF REQUIRED, SHOULD BE DONE BY THE COMPANY, Like civil, electrical, plumbing works etc….) | Yes, attach separate scope with details
44 | LINE POWER | Single Phase
45 | VAC | 220 V
46 | CURRENT | 13 A
47 | FREQUENCY | 60 Hz.
48 | PLUG TYPE | 3 Pin British
49 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
Company Stamp |  |  |  | Signature

## ITEM 68
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
68 | MIC10010 | 0 | INTUBATION DIFFICULT VIDEO PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Video Laryngoscope For difficult airway intubation
2 | Design | light weight & portable
3 | The system shall be based on the latest technological advances in difficult intubations | YES
4 | The unit shall be designed to provide a clear view of the vocal cords during intubation | YES
5 | The unit shall simplify intubation of the difficult airway | YES
6 | Trauma Airway | Yes
7 | Range in blade sizes allow coverage patient weights | Yes
8 | It shall have 60° angled blades | YES
9 | It should have pre-shaped stylet for quick intubation process | YES
10 | It should utilize minimum force for intubation | YES
11 | It should have anti-fog feature | YES
12 | The unit shall have integrated high resolution camera at good Position on blade to protect from blood and secretion | YES
13 | The unit shall mount on mobile carts and castors | YES
14 | Intubation of cervical spine immobilization | Yes
15 | Difficult Airway management and routine intubation | Yes
16 | Screen Size not less than 7 inch | LCD OR TFT
17 | Blades Size
18 | SIZE 2 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
19 | Made of medical grade plastic or stainless steel for less trauma handlingo teeth and soft tissue and easy | Specify
20 | Built in LED light for illumination | Yes
21 | Rechargable Battery on system | Yes
22 | Stylet | Qty:2
23 | Accessories | Any accessories, options and consumable items necessary to operate the offered system(s) must be clearly identified and priced separately
24 | SIZE 2 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BALDE | QTY : 120 Itemized and must be Optional Price
Company Stamp |  |  |  | Signature

## ITEM 69
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
69 | MIC10018 | 0 | MONITOR CAPNOGRAPHY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | application | monitoring CO2 measurements for neonatal, pediatric and adult patients
2 | measurements | EtCO2,FiCO2
3 | respiration rate | 0 bpm 150 bpm
4 | spo2 | yes
5 | printer | yes
6 | disply | color screen
7 | wave form trends | yes
8 | trends storage | yes 24hr.,specify
9 | accuracy | specify
10 | alarms | yes, visual and audible,specify
11 | Battery | yes
12 | Battery operating time | hr? specify
13 | power Supply | 220 VAC / 60 HZ
14 | All standard accessories required. | Yes ,Provide Itemized Price
Company Stamp |  |  |  | Signature

## ITEM 70
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
70 | MIC10026 | 0 | MONITOR PHYSIOLOGICAL BEDSIDE 17 INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Screen Size | Not less than 17Inch Medical grade TFT Type
2 | Mounting Bracket | Yes, Swivel base, wall or ceiling mounting according to Hosp. requirements.
3 | Wired remote control pad or fixed | Yes, with each monitor- specify.
4 | PARAMETERS TO BE MONITORED
5 | ECG / Resp , Heart Rate | Yes, 5 lead
6 | Spo2 | Yes , Nellcor Oximax OR Massimo set
7 | Dual Temperature | Yes
8 | NIBP | Yes
9 | Dual IBP | Yes
10 | Monitor Type | Modular
11 | Network port | Yes
12 | Interface | USP
13 | TRENDING Hrs. | Not less than 72 hrs. for all monitored parameters
14 | Haemodynamic & Drug Calculation | Yes
15 | Advance Arrhythmia & St Analysis | Yes
16 | All accessories required for the above parameters | Yes
17 | One set of the following | 1- Adult set 2- Pediatric set 3- Neonatal set
Company Stamp |  |  |  | Signature

## ITEM 71
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
71 | MIC10027 | 0 | MONITOR PHYSIOLOGICAL BEDSIDE 21 INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The system should be based on the latest technology in Monitors, patient transport that cover all applications needed in ICU. | Yes
2 | The monitor should have Intensive care software for Adult, Pediatric. | Yes
3 | The monitor should provide a high level of accuracy of monitored parameters. | Yes
4 | The monitor should be mobile stand with lockable wheels (minimum of 5), with basket, | Yes
5 | The monitor should provide a high level of accuracy (e.g. 2% approximate for ECG) of all monitored parameters. Vendors should provide parameters accuracy details. | Yes
6 | The monitor should be modular with single or multi-parameters per module. | Yes
7 | The display should be color LCD and should have an approximate size of 21 inches with a high level of contrast, brightness, resolution, and visibility from a distance and at a wide angle. Monitored parameters values should be large enough to be | Yes
8 | read from a distance.
9 | The monitor should be a touchscreen. | Yes
10 | The monitor should be capable of displaying six (8) physiological waveforms simultaneously as a minimum. | Yes
11 | The monitor should be capable of interfacing with the Hospital and Bedside Clinical Information System (HCIS). It should be able to download any recorded parameter, waveforms and numerical data, to the HCIS. The vendor should submit a list of | Yes
12 | compatible HCIS. (Please Provide it as itemized price)
13 | The vendor should provide state of the art technology for information technology used in the monitor applications. | Yes
14 | The vendor should provide an itemized price for all clinical software that is used on the monitor (vendor should support documents with evidence that prove the benefit of the clinical programs) | Yes
15 | The unit should be well constructed with durable materials to withstand typical abuse, heavy used environment. The vendor should provide support document that proves the durability of his product (International reference site should be mentioned). | Yes
16 | The monitor should be capable of providing an auxiliary ECG output signal for defibrillator/IABP synchronization, and an auxiliary BP output signal for IABP synchronization (preferable). | Yes
17 | The monitor should provide user-adjustable visual and audible alarms tailored for individual parameters. | Yes
18 | The monitor should be capable of providing graphical and numerical 24-hours trends of all monitored parameters. | Yes
19 | The monitor should be software-upgradeable to include future updates and/or additions of software application packages. | Yes
20 | The vendor should provide detailed technical information for the unique feature in the monitor. | Yes
21 | The vendor should provide detailed technical information about the following: | Yes
22 | HR Accuracy | Yes
23 | ECG HR range, bpm | Yes
24 | Interpretation | Yes
25 | Arrhythmia detect technique and number of lead analyzed | Yes
26 | ST technique and No. of leads analyzed | Yes
27 | Respiration method and Waveform displayed feature | Yes
28 | DISPLAY Size, (cm) in | Yes
29 | Display User Interface feature | Yes
30 | Remote display availability | Yes
31 | Trending Parameters | Yes
32 | Networking and Centralization. | Yes
33 | Hardwired/wireless | Yes
34 | Communication protocols | Yes
35 | Alarm Types and Feature for the machine | Yes
36 | The monitor should be capable of recording and display in waveforms and numerical values the following parameters (please provide it as itemized price): | Yes
37 | ECG, with 5 leads display, detection of pacemaker spikes, and arrhythmia. | Yes
38 | Respiration, with waveform display and alarms. | Yes
39 | Invasive Pressure (IP). | Yes
40 | Pulse Oximetry (SpO2), with Nellcor-Oxymax or Massimo RD Set cables and finger probes. | Yes
41 | Non-Invasive Blood Pressure (NIBP), and quick-connect hoses. | Yes
42 | Temperature, in two channels for two simultaneous probes, skin and rectal. | Yes
43 | End-Tidal CO2 (EtCO2), with waveform display, rapid warm-up and calibration times, and low dead-space connectors. | Yes
44 | Level of Consciousness. | Yes
45 | The vendor should provide an itemized price for different modules. | Yes
46 | The monitor should have a module that can measure at the same time: | Yes
47 | ECG, RESP. | Yes, QTY: 1
48 | Pulse Oximetry (SpO2), Nellcor-Oxymax / Massimo ACCORDING TO THE SUPPLIES CONTRACT IN SECTOR | Yes, QTY: 1
49 | Non-Invasive Blood Pressure (NIBP). | Yes, QTY: 1
50 | Invasive Blood Pressure (IBP). | Yes, QTY: 3
51 | Temperature. | Yes, QTY: 2
52 | All patient input should be defibrillator-protected. | Yes
53 | All cables, probes, and transducers (except for disposables) for all required parameters should be provided with the monitor, for adults and Pediatric. | Yes
54 | All Hardware and Accessories for each module should be designed for all type of patient use (provide itemized price). | Yes
55 | The vendor should provide an itemized price for all option, accessories on the machine. | Yes
56 | The manufacturer should guaranty that spare parts and technical support will be provided for at least ten years. | Yes
57 | Hardware and software should be industry standard with the latest specification; the vendor should provide a specification of both hardware and software used in the system. | Yes
58 | The vendor should provide an itemized price for all options in the system. | Yes
59 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and should be secured with adequate strain reliefs. | Yes
60 | The chassis should be grounded and grounding resistance should not exceed 0.15 ohm. | Yes
61 | If the unit is double insulated, it should be so labeled. | Yes
62 | Electrical leakage current from the chassis of the system should not exceed [0.5 mA per IEC 601-1 or 0.3 mA in the U.S. per NFPA 99-1993].Operator safety and system performance should not be adversely affected by fluid spills. | Yes
63 | The system performance should not be affected by EMI radiated or conducted through the power lines from another device. | Yes
64 | The power cord should be automatic retractable. | Yes
65 | The vendor should provide a full warranty of at least three (3) years. | Yes
66 | General Specifications and Requirements: | Yes
67 | Power Supply: 230 VAC 60 Hz. | Yes
68 | Hospital Grade British type Power Plugs, Line Cords, and Strain Relief. | Yes
69 | The unit must meet the applicable requirements and standards of the Underwriters Laboratories (UL), CE mark, IEC, ISO, and/or FDA Approval. | Yes
70 | The unit should be manufactured according to GMP guidelines. | Yes
71 | Labeled with all appropriate operation and safety tags and symbols. | Yes
72 | Switches, knobs, and other controls should be visible and clearly identified, and their functions should be self-evident. | Yes
73 | Switches, knobs, and other controls should be protected against accidental setting changes and should be sealed against penetration of liquids. | Yes
74 | The unit should be simple to operate and easy to clean and disinfect with Isopropanol alcohol 70%. | Yes
75 | The exterior should be well constructed with durable fluid/shock resistant materials to withstand typical abuse and cleaning. | Yes
76 | Rugged to handle in the rough transport environment. | Yes
77 | The unit should have no sharp edges. All external components should be securely mounted. | Yes
78 | All available options, consumables, or disposables deemed necessary for the intended function of the equipment and not listed in the Technical Specifications should be included and specified. | Provide Itemized Price
79 | Wall mount brackets as itemized price | Provide Itemized Price
80 | ETCO2 itemized | Provide Itemized Price
81 | Central station itemized (license of 20 beds) | Provide Itemized Price
82 | Itemized: The monitor should have a module that can measure 8 waves at the same time as follows : | Provide Itemized Price
83 | ECG, RESP. | Provide Itemized Price
84 | Pulse Oximetry (SpO2), Nellcor-Oxymax or Massimo. | Provide Itemized Price
85 | Non-Invasive Blood Pressure (NIBP). | Provide Itemized Price
86 | Invasive Blood Pressure (IBP). | Provide Itemized Price
87 | Temperature. | Provide Itemized Price
88 | EEG | Provide Itemized Price
89 | Cardiac Output. | Provide Itemized Price
Company Stamp |  |  |  | Signature

## ITEM 72
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
72 | MIC10028 | 0 | MONITOR PHYSIOLOGICAL MRI COMPATIBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Adult, Pediatric and Neonatal
2 | Type | Modular or Pre-configured
3 | MRI Compatible up to 3 Tesla | Yes
4 | PARAMETERS MEASURED
5 | ECG 3-5 Leads | Yes
6 | Heart Rate | 30 to 300 Beats per minutes
7 | NIBP | Yes
8 | SPO2 | Yes
9 | 2 x IBP | Yes, specify
10 | Respiration | Yes
11 | SCREEN
12 | Size, inch | 10?, approximate
13 | Type, Color, LCD, TFT | Specify
14 | Channel | 6 Channels OR MORE
15 | OPTIONAL PARAMETERS
16 | Temperature | Specify principle of measurement
17 | Data Recorder | Yes
18 | TRENDING (MEMORY)
19 | Parameters | Specify
20 | Graphical/Tabular | Specify
21 | Length of time, hr. | Specify
22 | DATA TRANSFER TO CONTROL ROOM | Specify
23 | Remote Control | Yes
24 | External Screen or PC | Specify
25 | ACCESSORIES
26 | ECG | Yes
27 | SPO2 | Fiber optic SPO2 sensors for adult pediatric & neonatal .
28 | NIBP | Cuff and hose for adult, pediatric & neonatal .
29 | Gases measurement accessories | Yes, for six months consumption.
30 | Pole mount bracket on 5 wheels MRI compatible. | Yes WITH STAND
31 | POWER REQUIREMENT | 220 V, 60 Hz
32 | Battery built-in operation time. | Not less than 60 min.
33 | OTHER SPECIFICATIONS | Approved by FDA, CE, ISO
Company Stamp |  |  |  | Signature

## ITEM 73
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
73 | MIC10029 | 0 | MONITOR PHYSIOLOGICAL TRANSPORT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | 12 inch Transport Monitor for Adult, Pediatric & Neonate
2 | CONFIGURATION | Configured
3 | STANDARD PARAMETERS FOR MONITOR | Multi-lead ECG, respiration, NIBP, SpO2, temperature IBP minimum 1 port & 2 is preferred
4 | ECG | Yes
5 | - Number of leads | 3, 5
6 | - HR range, bpm | 30 - 300
7 | - Max lead displayed / Simultaneouly | up to 4 / Specify
8 | - Alarm | Yes
9 | - Interpretation | Specify
10 | Arrhythmia detect | Specify
11 | - No. of leads analyzed | 3
12 | RESPIRATION | Yes
13 | - Method | Specify
14 | - Waveform display | Yes
15 | - Threshold control | Yes
16 | NIBP | Yes
17 | - Measurement technique | Osillometric
18 | - Cuff Inflation Adjustment Range for Adult/Pediatric | Automatic
19 | - Cuff Pressure Range for Adult/Pediatric | 0 to 300 mmHg
20 | Determination Time | 20 - 40 sec
21 | Pressure Reset Levels | Resolutions of 5 mmHg upto maximum limits
22 | Systolic Determination | Max:300 mmHg; Min:30 mmHg
23 | Diastolic Determination | Max:195mmHg; Min :10 mmHg
24 | Map Determination (Medium Arterial Pressure | Max:215 mmHg;Min :15 mmHg
25 | Pulse Rate Range | 30 - 200 BPM
26 | Over Pressure Cut off | 300 - 330 mmHg
27 | PULSE OXIMETRY | Yes
28 | - Accuracy | 1 Standard deviation
29 | -Adult/ Pediatric | 70 - 100 % 3.0 digits; under 70%
30 | - Saturation range | 0 - 100%
31 | - Pulse rate range | 20 - 250 bpm
32 | PROBE TYPES | Specify probe type. Disposable/reusable
33 | Patient range | Adult / Neonate/ Pediatric
34 | Cable length, m | Not less than 2.5
35 | TEMPERATURE | Yes
36 | - No.of inputs | 1
37 | - Probe type | Oral, axillary, rectal
38 | ALARMS & TRENDING
39 | ALARMS | Yes
40 | - Systolic | Preset at 30 mmHg low and 180 mmHg High
41 | - Diastolic | Preset at 15 mmHg low and 130 mmHg High
42 | Silence active alarm | Yes
43 | TRENDING | Monitor stores 24 hrs of data; up to 100 reading within time period
44 | -Parameters | All
45 | -Graphical∕tubular | Yes
46 | BATTERY | Yes, Built in with integrated charger
47 | - Operating time | Not less than 2.5 Hrs
48 | - Low battery alarm | Yes
49 | PRINTER/RECORDER | Yes, Specify the paper type
50 | PHYSICAL SPECIFICATIONS
51 | Dimensions | Specify
52 | Weight | Specify
53 | UNIT COMPLETE WITH FULL ACCESSORIES | ALL THE ACCESSORIES SHOULD BE SUPPLY AS PER PARAMETERS
54 | IMPORTANT NOTE | Please specify whether accessories are standard or optional, otherwise they will be considered standard
55 | - All sizes cuff for Adult | Yes
56 | - All sizes cuff for Pediatric & neonate | Yes
57 | - Spo2 probe/adult (reusable & disposable) | Yes
58 | - Spo2 probe/pedia and neonate (reusable & disposable) | Yes
59 | - Temperature probes | Yes
60 | - Paper roll | 1 box
61 | Reusable &Disposable ECG electrodes | Yes
62 | LATEST MODEL TO BE INSTALLED AS PER THE AVAILABILTIY FROM MANUFACTURER | Yes
63 | INSTALLATION & PRE - INSTALLATION, IF REQUIRED, SHOULD BE DONE BY THE COMPANY, Like civil, electrical, plumbing works etc….) | Yes, attach separate scope with details
64 | HIS compatible | Yes / Specify
65 | LINE POWER | Single Phase
66 | VAC | 220 V
67 | CURRENT | 13 A
68 | FREQUENCY | 60 Hz.
69 | PLUG TYPE | 3 Pin British
70 | UNIT COMPLETE WITH FULL ACCESSORIES | Yes
Company Stamp |  |  |  | Signature

## ITEM 74
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
74 | MIC10032 | 0 | NEGATIVE AIR SYSTEM WITH HEPAFILTER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | NEGATIVE AIR SYSTEM - HEPA FILTER | FOR ICU THAT HAS NO ISOLATION SYSTEM
2 | EFFICIENCY OF FILTERING OF PARTICLES | MORE THAN 99% FOR PARTICLES OF SIZE 0.3 MICRON
3 | HEPA FILTER | YES
4 | PRE-FILTER | INCLUDED
5 | HEPA FILTER CHANGE INDICATOR | INCLUDED
6 | PRESSURE MEASUREMENT & INDICATOR | YES , INCLUDED
7 | NUMBER OF MOTORS | TWO
8 | WEIGHT | SPECIFY
9 | EXHAUST OUTPUT | PROVIDE THE OFFERED TYPE INSTALLATION METHOD
10 | ELECTRICITY | 220 VOLTS , 60 HZ
11 | FDA , CDC , OSHA | MUST PROVIDE ANY OF THEM
Company Stamp |  |  |  | Signature

## ITEM 75
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
75 | MIC10033 | 0 | PACEMAKER CARDIAC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled, Single chamber , Battery backup, external invasive pacemaker with asynchronous and inhibit pacing modes, incorporating LCD display for clear visual indication of pacer parameters and warning messages
2 | TYPE | Single Chamber
3 | APPLICATION | AOO, AAI, VOO, VVI
4 | RATE :
5 | - Range, ppm | 30 - 150
6 | - Graduation Intervel | 5 ppm
7 | - Accuracy | 1 ppm
8 | Rapid atrial pacing function | up to 380/ min
9 | OUT PUT AMPLITUDE, mA | Yes
10 | - Atrium | 0.5 - 10 V, 1 - 20 mA
11 | - Ventricle | 0.5 - 10 V, 1 - 20 mA
12 | SENSITIVITY, mV
13 | - Atrium | 0.5 - 10
14 | - Ventricle | 0.5 - 10
15 | REFRACTORY PERIOD, msec
16 | - Atrium | 200 - 400
17 | - Ventricle | 200 - 400
18 | BATTERY
19 | -Type | Any common type
20 | - Operation, hr | Specify
21 | - Low-battery Indicator | Yes
Company Stamp |  |  |  | Signature

## ITEM 76
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
76 | MIC10037 | 0 | PHOTOTHERAPY DOUBLE PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | DOUBLE - PHOTOTHERAPY UNIT | FOR HYPERBILIRUBENEMIA
2 | FREESTANDING WITH LARGE CASTORS | YES
3 | TUBE HOLDER | CYLINDRICAL
4 | BULBS LIFE | 1000 hrs
5 | HEIGHT ADJUSTABLE | YES
6 | MATTRESS | INCLUDED
7 | BULB PROTECTION | YES INCASE BULBS GOT BROKEN
8 | EFFECT OF BABY TEMP | MINIMlL INCREASE TO BABY TEMP
9 | INTENSITY OF THE LIGHT | UP TO 60 w/cm2/nm
10 | POWER | 220 VOLTS
11 | OTHER SPECS | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 77
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
77 | MIC10042 | 0 | THROMBOSIS PREVENTION VIEN DEEP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled operation and safeties | YES
2 | Compression pressure shall be achieved by means of incorporated pump (no external connection to compresses air line shall be necessary) | YES
3 | Features: Specify . | YES
4 | Single Operation Button | YES
5 | One push operation selects | YES
6 | adjusts pressure manually | YES
7 | Monitor/alarm Specify. | YES
8 | One pump for calf and thigh cuffs . | YES
9 | Calf and thigh cuffs available | YES
10 | Pressure and control parameters should be specified: | YES
11 | Pressure regulation method | YES
12 | Approximate applied pressures should be as low as possible between 30 and 60 mm Hg for the calf (varies per body part) and not more than 200 mm Hg as a rule. | YES
13 | Approximate cycle time range of 30 to 60 sec. | YES
14 | The unit shall display important operating parameters and other alert messages and alarms. Specify display characteristics and list all displayed information. | YES
15 | Specify high pressure safety cutoff points (preferably more than one level, fully independent) and breaking method | YES
16 | Safety features should include: | YES
17 | Splash proof front panel design | YES
18 | Compression halt on alarm | YES
19 | Maintenance-free operation | YES
20 | 1 meter anti kink tubing can be extended for convenient and safe pump positioning. | YES
21 | secure snap lock connections to pump to prevent accidental disconnection. | YES
22 | The unit should be portable, self contained incorporating mounting bracket for IV pole and bed rails | YES
23 | The unit shall be able to connect to all types of garments available on the market (open-system) using universal tubing connectors (luer lock or snap-on) | preferred
24 | Compliance with standards & legislation: | YES
25 | The system must comply with the Electrical safety standards for electrical safety IEC-60601 | YES
26 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any. | YES
27 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements. | YES
28 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions | YES
29 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included. | YES
30 | SLEEVE THIGH SMALL | QTY : 30
31 | SLEEVE THIGH MEDIUM | QTY : 30
32 | SLEEVE THIGH LARGE | QTY : 30
33 | SLEEVE KNEE SMALL | QTY : 30
34 | SLEEVE KNEE MEDIUM | QTY : 100
35 | SLEEVE KNEE LARGE | QTY : 60
36 | PLUG | BRITISH
37 | LANGUAGE | ENGLISH
38 | CERTIFICATES | APPROVED FROM REGULATORY BODY
39 | OTHERS | PROVIDE PROVE OF USE FROM DIFFERENT HOSPITALS LOCALLY OR INTERNATIONALLY
40 | UNIT OF MEASUREMENT | SI : INTERNATIONAL STANDARD
41 | Adjustable pressure or pre-set pressure. | YES
42 | Display real-time pressure or indicators for set pressure. | YES
43 | Equipped with high pressure mechanical release valve. | yes
44 | Shall allow one limb or two limb operation. | yes
45 | Equipped with overpressure, leaking, kinked tubing, internal electronic malfunction low pressure and high pressure alarms. | yes
46 | Peripheral circulatory assist units designed for sequential compression. | YES
47 | VENDOR MUST PROVIDE ITEMIZED PRICES FOR ALL SLEEVES KIND AND SIZED TO BE SELECTED | MUST
48 | VENDOR MUST PROVIDE FOC MACHINE WITH SUGGESTED NUMBER OF SLEEVES TO BE PURCHASED | YES
Company Stamp |  |  |  | Signature

## ITEM 78
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
78 | MIC10043 | 0 | VENTILATOR BIPAP ICU |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Modes | NIV, NIV-ST , HiFLOW
2 | CPAP | 4 TO 20 cm H2O
3 | S/T | 4 to 40 cm H2O
4 | Parameters Control Ranges
5 | IPAP | 4 to 40 cm H2O
6 | EPAP | 4 to 20 cm H2O
7 | Rate | 4 to 40 BPM
8 | Timed Inspiration | 0.5 to 3.0 sec
9 | IPAP >Rise Time | 0.05 to 0.40 sec
10 | Oxygen Concentration (%O2) | 21 to 100%
11 | Display Ranges
12 | IPAP | 0 to 50 cm H2O
13 | EPAP | 0 to 50 cm H2O
14 | CPAP | 0 to 50 cm H2O
15 | Rate | 0 to 120 BPM
16 | Exhaled Tidal Volume (VT) | 20 to 2000 ml
17 | Minute Ventilation (Min Vent) | 0 to 99 L/min
18 | Total Leak (Tot Leak) | 0 to 80%
19 | Patient Leak(Pt. Leak) | 0 to 100 L/min
20 | Peak Inspiratory Pressure (PIP) | 0 to 50 cm H2O
21 | Percent of Patient Triggered Breaths (Pt. Trig) | 0 to 100%
22 | It/Tot | 0 to 100%
23 | Alarm Control Ranges
24 | High Pressure | 5 to 60 cm H2O
25 | Low Pressure | Disabled;0 to 40 cm H2O
26 | Low Pressure | Delay 0 to 60 sec
27 | Apnea | Disabled; 20 to 40 sec
28 | Low Minute Ventilation | Disabled; 1 to 99 L/min
29 | High Rate | 4 to 120 BPM
30 | Low Rate | 4 to 120 BPM
31 | POWER SUPPLY | 220 Vac - 60 Hz.
32 | Inspiratory flow on NIV, NIV-ST MODE | 2-220 L/MIN
33 | FLOW ON HIFLOW MODE | 2-70 L/MIN
34 | Hi Flow Canula, Open System | YES
35 | Screen , and Monitoring | Not less than 7' color , touch
36 | Humidifier | REQUIRED
37 | Built In Turabian | YES
38 | OTHER SPECIFICATION | FDA, CE, SPECIFY OTHERS
Company Stamp |  |  |  | Signature

## ITEM 79
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
79 | MIC10044 | 0 | VENTILATOR CPAP ADULT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Country Of origin | Specify
2 | Model | Specify
3 | Manufacture | Specify
4 | VENTILATOR | CPAP
5 | Gas supply pressure | 40,6 - 87 psi
6 | PANUMATIC OPERATION | yes
7 | Breathes per Minute / Accuracy / Range 1 second inspiretor / Range 2 second inspiretory | 10% / 0 - 30 bpm / 0 - 20 bpm
8 | Tidal Volume / Accuracy / Range 1 / Range 2 | 10 % / 200 - 600 ml / 400 - 1500 ml
9 | Adjustable tidal volume | yes
10 | Adjustable BpM | yes
11 | Chilled and Adult settings | yes
12 | Internal demand valve | yes
13 | Antibreth stochking | yes
14 | Manometer readout | yes
15 | Adjustable relife with alarm | yes
16 | Low source gas pressure alarm | yes
17 | Accessories
18 | Ventilator Circuit | yes
19 | Oxygen House | yes
20 | Oxygen Regulator | yes
21 | Internal air filter | yes
22 | CPAP ventilator circuit Disp. | yes ,50 set to be included with each unit
23 | OTHER SPECIFICATION | FDA , CE , APPROVED.
Company Stamp |  |  |  | Signature

## ITEM 80
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
80 | MIC10045 | 0 | VENTILATOR HFO ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Modes of Ventilation
2 | Controlled / Assist modes | Pressure, Volume & Volume Guaranteed
3 | Intermittent modes | SIMV-VC, SIMV-PC , SIMV-VG
4 | Two level of pressure mode | BiPAP Invasive OR Equivelent BiPAP , NIV-ST ( with Mask )
5 | Inverse Ratio mode | IRV/APRV or equivalent
6 | Spontaneous mode | CPAP/PS mode
7 | NIV | Yes
8 | Mode Settings & controls
9 | Tidal Volume | 2 - 2000 ml
10 | Respiratory Rate | 3 - 120 b/m
11 | FiO2 | 21% - 100%
12 | Inspiratory Pressure | 1 - 80 cmH2O
13 | PEEP | 0 - 50 cmH2O
14 | Pressure Support | 0 - 60 cmH2O
15 | I:E ration | 1:9 to 4:1
16 | Inspiratory Time | 0.25 - 15 sec
17 | Flow Trigger | 1 - 9 l/m
18 | Pressure Trigger | -10 to -0.25
19 | Inspiratory Pause | Yes, Specify
20 | Pressure slope/rise time | Yes, Specify
21 | Bias Flow | Yes, Specify
22 | Tube Compensation | Yes
23 | Leak Compensation | Yes
24 | Trigger Compensation | Yes
25 | Apnea backup / rate | Yes
26 | Manual Breath | Yes
27 | Inspiratory Hold | Yes
28 | Expiratory Hold | Yes
29 | AutoPEEP measurement | Yes
30 | Weaning Parameters (P0.1, NIF & RSBI) | Yes, Specify
31 | Lung Compliance & Airway Resistance | Yes
32 | Monitored Parameters
33 | PIP, Pplat, MAP & PEEP | Yes
34 | TV & MV (controlled & Spontaneous) | Yes
35 | Respiratory Rate | Yes
36 | FiO2 | Yes
37 | Clinical Alarms (High/Low)
38 | FiO2 | Yes
39 | Peak Pressure | Yes
40 | TV & MV | Yes
41 | Patient Disconnect | Yes
42 | Leak | Yes
43 | Technical
44 | O2 sensor | Specify Type
45 | Active Expiratory Valve (Autoclavable is preferred) | Yes
46 | electronic micropump nebulizer | Yes
47 | Display & Control | Colored Touch Screen not less than 12
48 | Patient circuit supported arm | Yes
49 | Backup battery | not less than 2 hrs
50 | Mobile base/ cart with locking castors | Yes
51 | Central Gas inlets (O2 & Air) | Yes
52 | Electricity 220v, 60 Hz or Auto. | Yes
53 | Air inlet Filter | Yes
54 | Technical Alarms
55 | Gas supply or power Failure, Vent inoperative & Low battery | Specify
56 | Data Displayed | Numerical, Graph/waveforms & Loops
57 | Trends | not less than 72 hrs for all patient data, Specify
58 | Patient Circuit | 1 Ea : Adult - Pediatric - Neonate
59 | WEANING MODE | MANDATORY INCLUDED : with Reusable Sensors x Qty: 3 to operate weaning mode
60 | Additional Modes: PAV, ASV, PPS, NAVA & BiLevel-VG or equivalent & Smartcare with Capno | Specify what you offer
61 | P/V tool, Low flow PV, Spirodynamic, Static PV loop, or equivalent & others | Specify
62 | Open Lung tool, Lung Recruitment, FRC with PEEP inview or equivalent & others | Specify
63 | EtCO2, EtO2, VCO2, VO2, EE & RQ | Specify Optional : Itemized Price
64 | Humedifier with Cahmber & Heater Wire : Approved type in M.O.H | To be supplied with the unit
Company Stamp |  |  |  | Signature

## ITEM 81
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
81 | MIC10046 | 0 | VENTILATOR ICU ADULT & PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SYSTEM SELF CHECK | YES
2 | Controlled / Assist modes | Pressure, Volume & Volume Guaranteed
3 | Intermittent modes | SIMV-VC, SIMV-PC , SIMV-VG
4 | Two level of pressure mode | BiPAP Invasive OR Equivelent BiPAP , NIV-ST ( with Mask )
5 | Spontaneous mode | CPAP/PS mode
6 | NIV , NIV-ST | REQUIRED
7 | Mode Settings & controls
8 | Tidal Volume | 50 - 2000 ml
9 | Respiratory Rate | 3 - 80 b/m
10 | FiO2 | 21% - 100%
11 | Inspiratory Pressure | 1 - 60 cmH2O
12 | PEEP | 0 - 40 cmH2O
13 | Pressure Support | 0 - 60 cmH2O
14 | I:E ration | 1:9 to 4:1
15 | Inspiratory Time | 0.25 - 15 sec
16 | Flow Trigger | 1 - 9 l/m
17 | Pressure Trigger | -10 to -0.25
18 | Inspiratory Pause | Yes, Specify
19 | Pressure slope/rise time | Yes, Specify
20 | Bias Flow | Yes, Specify
21 | Tube Compensation | Yes
22 | Leak Compensation | Yes
23 | Trigger Compensation | Yes
24 | Apnea backup / rate | Yes
25 | Manual Breath | Yes
26 | Inspiratory Hold | Yes
27 | Expiratory Hold | Yes
28 | AutoPEEP measurement | Yes
29 | Weaning Parameters (P0.1, NIF & RSBI) | Yes, Specify
30 | Lung Compliance & Airway Resistance | Yes
31 | Monitored Parameters
32 | PIP, Pplat, MAP & PEEP | Yes
33 | TV & MV (controlled & Spontaneous) | Yes
34 | Respiratory Rate | Yes
35 | FiO2 | Yes
36 | Clinical Alarms (High/Low)
37 | Peak Pressure | Yes
38 | TV & MV | Yes
39 | Patient Disconnect | Yes
40 | Leak | Yes
41 | Technical
42 | O2 sensor | BUILT IN
43 | Active Expiratory Valve (Autoclavable is preferred) | Yes
44 | electronic micro pump nebulizer | OPTIONAL ITEMIZED PRICE
45 | Display & Control | Colored Touch Screen not less than 12 INCH
46 | Patient circuit supported arm | ADJUSTABLE
47 | Backup battery | not less than 2 hrs
48 | Mobile base/ cart with locking castors | Yes
49 | Central Gas inlets (O2 & Air) | BRITISH STANDARD BOC
50 | Electricity 220v, 60 Hz or Auto. | Yes
51 | Air inlet Filter | Yes
52 | Technical Alarms
53 | Gas supply or power Failure, Vent inoperative & Low battery | Specify
54 | Data Displayed | Numerical, Graph/waveforms & Loops
55 | Trends | not less than 72 hrs for all patient data, Specify
56 | WEANING MODE | yes
57 | Humedifier with Cahmber & Heater Wire : Approved standard by sector | INCLUDED WITH UNIT
58 | MOBILE CART | INCLUDED WITH BASKET
59 | PLUG STANDARD | BRITISH
60 | LANGAUGE | ENGLISH LANGUAGE
61 | UNIT OF MEASUREMENT | INTERNATIONAL STANDARD SI
62 | APRV or equivalent | included
63 | Additional Modes: PAV, ASV, PPS, NAVA & BiLevel-VG or equivalent & Smartcare with Capno | OPTIONAL
64 | P/V tool, Low flow PV, Spirodynamic, Static PV loop, or equivalent & others | OPTIONAL
65 | EtCO2, EtO2, VCO2, VO2, EE & RQ | OPTIONAL
66 | Open Lung tool, Lung Recruitment, FRC with PEEP inview or equivalent & others | OPTIONAL
67 | BREATHING CIRCUIT | 100 WITH EACH UNIT , OPEN STANDARD
68 | DEVICE CERTIFICATION | APPROVED FROM REGULATORY BODY
69 | AIR COMPRESSOR INTEGRATED WITH THE DEVICE | ITEMIZED PRICE
70 | PROVIDE WHERE THE DEVICE & TECHNOLOGY WERE USED | PROVIDE REF HOSPITALS
71 | Modes of Ventilation
72 | BI-LEVEL MODE ( APRV ) | OPTIONAL ITEMIZED
73 | Inverse Ratio mode : IRV | optional
Company Stamp |  |  |  | Signature

## ITEM 82
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
82 | MIC10047 | 0 | VENTILATOR ICU NEONATAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Modes of Ventilation
2 | Controlled / Assist modes | Pressure, Volume & Volume Guaranteed
3 | Intermittent modes | SIMV-VC, SIMV-PC , SIMV-VG
4 | Two level of pressure mode | BiPAP Invasive OR Equivelent BiPAP , NIV-ST ( with Mask )
5 | Inverse Ratio mode | IRV/APRV or equivalent
6 | Spontaneous mode | CPAP/PS mode
7 | NIV | Yes
8 | Mode Settings & controls
9 | Tidal Volume | 2 - 2000 ml
10 | Respiratory Rate | 3 - 120 b/m
11 | FiO2 | 21% - 100%
12 | Inspiratory Pressure | 1 - 80 cmH2O
13 | PEEP | 0 - 50 cmH2O
14 | Pressure Support | 0 - 60 cmH2O
15 | I:E ration | 1:9 to 4:1
16 | Inspiratory Time | 0.25 - 15 sec
17 | Flow Trigger | 1 - 9 l/m
18 | Pressure Trigger | -10 to -0.25
19 | Inspiratory Pause | Yes, Specify
20 | Pressure slope/rise time | Yes, Specify
21 | Bias Flow | Yes, Specify
22 | Tube Compensation | Yes
23 | Leak Compensation | Yes
24 | Trigger Compensation | Yes
25 | Apnea backup / rate | Yes
26 | Manual Breath | Yes
27 | Inspiratory Hold | Yes
28 | Expiratory Hold | Yes
29 | AutoPEEP measurement | Yes
30 | Weaning Parameters (P0.1, NIF & RSBI) | Yes, Specify
31 | Lung Compliance & Airway Resistance | Yes
32 | Monitored Parameters
33 | PIP, Pplat, MAP & PEEP | Yes
34 | TV & MV (controlled & Spontaneous) | Yes
35 | Respiratory Rate | Yes
36 | FiO2 | Yes
37 | Clinical Alarms (High/Low)
38 | Peak Pressure | Yes
39 | TV & MV | Yes
40 | Patient Disconnect | Yes
41 | Leak | Yes
42 | Technical
43 | O2 sensor | Specify Type
44 | Active Expiratory Valve (Autoclavable is preferred) | Yes
45 | electronic micropump nebulizer | Yes
46 | Display & Control | Colored Touch Screen not less than 12
47 | Patient circuit supported arm | Yes
48 | Backup battery | not less than 2 hrs
49 | Mobile base/ cart with locking castors | Yes
50 | Central Gas inlets (O2 & Air) | Yes
51 | Electricity 220v, 60 Hz or Auto. | Yes
52 | Air inlet Filter | Yes
53 | Technical Alarms
54 | Gas supply or power Failure, Vent inoperative & Low battery | Specify
55 | Data Displayed | Numerical, Graph/waveforms & Loops
56 | Trends | not less than 72 hrs for all patient data, Specify
57 | Patient Circuit | 50 Ea : 30 Adult - 10 Pediatric - 10 NEONATE
58 | WEANING MODE | MANDATORY INCLUDED :with Reusable Sensors x Qty: 3 to operate weaning mode
59 | Additional Modes: PAV, ASV, PPS, NAVA & BiLevel-VG or equivalent & Smartcare with Capno | Specify what you offer
60 | P/V tool, Low flow PV, Spirodynamic, Static PV loop, or equivalent & others | Specify
61 | Open Lung tool, Lung Recruitment, FRC with PEEP inview or equivalent & others | Specify
62 | EtCO2, EtO2, VCO2, VO2, EE & RQ | SpecifyOptional : Itemized Price
63 | Humedifier with Cahmber & Heater Wire : Approved standard by sectors | To be supplied with the unit
Company Stamp |  |  |  | Signature

## ITEM 83
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
83 | MIC10048 | 0 | VENTILATOR TRANSPORT MRI COMPATABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Adult, Pediatric & Neonatal
2 | Type | Transport
3 | MRI compatible Up to 3 Tesla | Yes 10 CM FROM THE CORE
4 | VENTILATION MODE | CMV, PEEP, CPAP, Other Specify
5 | Tidal Volume | 0 - 2000 ml or wider
6 | Ventilation Rate | 0 - 60 per minute or wider
7 | Inspiratory time range | Specify
8 | Expiratory time range | Specify
9 | Expiratory flow range | Specify
10 | Pressure relief range | Specify
11 | Air way pressure indication | Either, mechanically gauge or digital display
12 | Single circuit patient Disp. | Yes, other specify , 50 set to be included with each unit
13 | AIR OXYGEN BLENDER, MRI COMPATIBLE | built-in
14 | FiO2 Percent | 21 - 100 %
15 | PNEUMATIC POWERED | Yes, O2 & Medical air
16 | UNIT WEIGHT , KG, LBS. | Specify
17 | ACCESSORIES
18 | Pole mount with cylinder bracket MRI Compatible | Yes
19 | E cylinder regulator oxygen, MRI compatible. | Yes, Qty. 1
20 | E cylinder regulator medical Air MRI compatible. | Yes, Qty. 1
21 | PMG Hose with female connection compatible with Hospital gases for oxygen and medical. | Yes
22 | Aluminium E cylinder oxygen | Yes, Qty. 3
23 | Aluminium E cylinder air | Yes, Qty. 3
24 | OTHER SPECIFICATIONS | Approved by FDA, CE, ISO
Company Stamp |  |  |  | Signature

## ITEM 84
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
84 | MIC10049 | 0 | VENTILATOR TRANSPORT WITH MONITOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Monitor included | Yes : For patient vital Parameters
2 | Parameters | ECG , Resp , NIBP , 2x IBP , TEMP , SPO2
3 | SPO2 | Masimo or Nellcor Oxi-max ( must comply with Sector Supplies Standard )
4 | Temp | Yes
5 | Waveform | 3 - 5 Wavforms
6 | Screen Size | >=12"
7 | Monitor Type | Ventilation parameters and Patient Vital Signs displayed on Same Monitor
8 | Ventilator
9 | Configuration | Adult, Pediatric , Neonatal
10 | Tidal volume, L | 2-2000 mL
11 | Breath rate, br/min | 0 - 80
12 | Inspiratory time, sec | 0.2 - 3 Sec
13 | Inspiratory flow, L/min | 3 - 100 L/Min, specify
14 | Pressure level, cm H2O | Specify
15 | Pressure ramp | Specify
16 | Pressure support | 0 - 50
17 | Trigger Mechanism | Flow or pressure or both, Specify.
18 | Sensitivity, cm H2O | Flow or pressure trigger, specify
19 | FiO2, Percentage | 21-100 %
20 | I:E ratio | 1:9 TO 4:1
21 | Adjustable PEEP | 0 - 40 cm H2O/ higher
22 | OPERATING MODES
23 | Assist / Control Mode | Yes
24 | Volume Mode | Yes
25 | Pressure Mode | Yes
26 | Spontaneous / CPAP Mode | Yes
27 | SIMV Mode | Yes
28 | MONITOR PARAMETERS
29 | Pressure, PA , MAP, PEEP. | Yes
30 | Volume, Tidal, Minute | Yes
31 | Inspiratory / Expiratory Time | Yes
32 | I:E Ratio | Yes
33 | Respiratory Rate | Yes
34 | PATIENT ALARMS
35 | FiO2, Low/High minute volume, Low inspiratory pressure, High pressure, Loss of PEEP, Apnea, Inverse IE, High continuous pressure occlusion, High respiratory rate, Others. | All
36 | EQUIPMENT ALARMS
37 | Gas-supply failure, Power failure, Vent inoperative, Low battery, Self Diagnostic, Others. | All
38 | Silence | 2 minutes
39 | CONTROL TYPES | Dial, soft keys, toggle switch
40 | Breathing circuit Adult & Pediatric Reusable | 3 set ( Option )
41 | Air Comp /Turbine as Air Backup Source | YES, Specify
42 | BATTERY WITH CHARGER | Built-in, operational hours on Battery >= 6hrs
43 | Weight of ( Ventilator + Monitor ) | <=12KG
44 | OTHER SPECIFICATION | FDA, CE, ISO Approved.
45 | plug type | british
46 | mobile cart | required , 4-5 castors with basket
47 | certifications | approved regulatory body
48 | circuit | 100 each included with the unit
Company Stamp |  |  |  | Signature

## ITEM 85
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
85 | MIC10054 | 0 | ULTRASOUND POC ADVANCE PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Regulation Standard :
2 | FDA 510 (K) Clearance / CE Mark & SFDA Registration | Yes
3 | Validated Up to Date Certificates Must be Submitted with the Offer | Yes, Failure of Submission and Comply Will Disqualify the Offer
4 | Product Description :
5 | Range | Laptop Ultrasound
6 | Purpose | Used for point of care General Clinical Routine Applications.
7 | Console Design :
8 | Laptop / Hand Carried on height adjusted mobility Cart | Cart Should Include the Following :Integrated Power SupplyProbe Portswheels with locking braking system
9 | Total Weight - (incl. battery) | ≤ ( 5 ) kg .
10 | Esay Acess mode keys | Yes, (Specify)
11 | Screen Monitor | ≥ (12) inch. LCD with High Resolution
12 | Image To Screen Ratio @ highest Resolution | >=75%
13 | Security Lock | Yes
14 | Built-In Battery | Yes
15 | Battery scanning time | 1 hour
16 | Power Requirement | 220-240 V / 60 Hz
17 | Computing Capabilities :
18 | Operation System | Windows or equivalent
19 | Integrated Hard Desk Type | Flash memory≥15 GB for better shock Absorbance
20 | Fast System boot up | Yes , Less than 30 Sec
21 | DVD /CD burner/USB port | Yes
22 | Clinical Applications :
23 | Radiology, abdominal, Small parts | Yes
24 | OB / Gyn, Breast | Yes
25 | Musculoskeletal | Yes
26 | Vascular | Yes
27 | Urology | Yes
28 | pediatrics Radiology | Yes
29 | Other Available Applications | Yes
30 | Multi Frequency Probes : | One Each , With dedicated Operational Software
31 | Linear array | 3 -12 MHz (±2)
32 | Convex array | 2 - 5 MHz (±2)
33 | Phased array (TCD) | 1 - 5 MHz (±2)
34 | Active Probe Ports | 2-3 Active Ports
35 | Operating Modes :
36 | Brightness Mode (B-Mode) (2D) | Yes
37 | Motion Mode (M-Mode) | Yes
38 | Color Doppler mode | yes
39 | Pulsed wave Doppler (PW) with HPRF | Yes
40 | High PRF Doppler Mode | Yes
41 | Color Flow Doppler Mode (CFM) | Yes
42 | Power Doppler Mode | Yes
43 | Continuous Wave Doppler | Yes
44 | Doplex Mode | Yes
45 | Triplex Mode | Yes
46 | Scanning Parameters :
47 | Imaging Depth, cm | Not less than 30 cm
48 | Dynamic Range | Up to 165 dB
49 | TGC Control | Yes
50 | Frame Rate 2D | Specify Range
51 | Frame rate Color Doppler | Specify Range
52 | GrayScale Levels | 256 ( 8 bits )
53 | Scanning Automatic Optimization | Capability for Auto optimization in 2D & Doppler
54 | Digital Calipers | Yes
55 | Preprocessing | Yes
56 | Postprocessing | Yes, (Measurment, adjust the Daynamic Range, TGC, color map & Gain) on stored images
57 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
58 | Doppler :
59 | Type | PW in all probes
60 | Frequency Display | Yes
61 | Doppler baseline shift after freeze | Yes
62 | Velocity Display | Yes
63 | User – defined annotation | Yes
64 | GrayScale Levels | Yes
65 | Transmit focal zones | Select between 1 to 8
66 | Digital reconstructed zoom | 16 levels
67 | Gain Control after freeze | Yes
68 | Adaptive Gain Compensation | Yes
69 | Directional color power Doppler | Yes
70 | Cine Features :
71 | Cine Storage | Yes, (Still images / Clips / Volume)
72 | Length / Capacity | 140 MB (up to 3000 images)
73 | Scanning Features :
74 | Speckle Reduction Imaging | Yes, Organ Specific
75 | Trapeziodal Imaging / Virtual Convex | Yes
76 | B-Mode Steering | Yes
77 | Compound Imaging | Yes
78 | (Panoramic Imaging) - Free Style Extended Imaging | Yes, Min 30 cm
79 | Tissue Harmonic Imaging | Yes
80 | TGC (Time Gain Compensation) | Yes
81 | Color flow mapping | Standard
82 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
83 | Automatic Doppler Calculations | Yes
84 | Adjustable Transmit Focus | Yes
85 | Dynamic Receive Focus | Yes
86 | Selectable Dynamic Range | Yes
87 | Advance Needle Visualization | Yes
88 | B - Color For Color Coding of Grey Levels | Yes
89 | Automatic Frequency and focal zone optimization with depth adjustment | Yes
90 | Zoom / Pan :
91 | Real- time image | Yes
92 | Frozen image | Yes
93 | Connectivity :
94 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicaple
95 | Wireless Connection | Yes
96 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
97 | Storage and Playback of M-mode, PW Spectral Doppler. | yes
98 | Recalculation for stored images | Yes
99 | Security / Privacy :
100 | Accessories :
101 | Secondary Battery | Yes
102 | Carrying Case | Yes
103 | Requirement :
104 | Supplier is an authorized representative of the manufacturer by SFDA | Yes (Attach official documents)
105 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 86
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
86 | MIC10056 | 0 | CENTRAL STATION PHYSIOLOGICAL 4 BEDS LICENCE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | Adult , Pediatric , NICU
2 | CENTRAL STATION | Yes MEDICAL GRADE
3 | No. of Patient | Liscence for 4 Beds
4 | Display Monitor | High resolution color not less than 19 Inch , Flat screen TFT. Type
5 | Number Of Monitors | TO FIT ALL LISENCES REQUESTED
6 | Trending Full Disclosure | Not Less Than 72 Hrs.
7 | Operation | Fixed key and menu driven/touch screen, or other - Specify
8 | Communication | Network, central station, remote bedside monitor
9 | Bed To Bed Communication | Yes
10 | Configurable Reporting | Yes
11 | ECG interpretation | Yes
12 | Computer medical grade | YES
13 | keyboard and mouse. | Wireless or wired Specify
14 | RECORDER | 4 Channel, Digital
15 | Recording | Yes, One for the system
16 | Laser Printer | Yes, One for the system
17 | Data | All, Patients
18 | Waveforms | Yes
19 | Trends | Yes
20 | Strip Chart | Yes
21 | Remote Operation | From bedside monitor
22 | Alarm - Audio and Visible | Yes
23 | Alarm Lamp and flashing | Yes
24 | Central Station Must Be Compatible with the selected Bed Side Monitor | MUST BE INTERFACED WITH BED SIDE MONITOR FROM SAME MANUFACTURER
Company Stamp |  |  |  | Signature

## ITEM 87
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
87 | MIC10057 | 0 | CENTRAL STATION PHYSIOLOGICAL 8 BEDS LICENCE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | Adult , Pediatric , NICU
2 | CENTRAL STATION | Yes MEDICAL GRADE
3 | No. of Patient | Liscence for 8 Beds
4 | Display Monitor | High resolution color not less than 19 Inch , Flat screen TFT. Type
5 | Number Of Monitors | TO FIT ALL LISENCES REQUESTED
6 | Trending Full Disclosure | Not Less Than 72 Hrs.
7 | Operation | Fixed key and menu driven/touch screen, or other - Specify
8 | Communication | Network, central station, remote bedside monitor
9 | Bed To Bed Communication | Yes
10 | Configurable Reporting | Yes
11 | ECG interpretation | Yes
12 | Computer medical grade | YES
13 | keyboard and mouse. | Wireless or wired Specify
14 | RECORDER | 4 Channel, Digital
15 | Recording | Yes, One for the system
16 | Laser Printer | Yes, One for the system
17 | Data | All, Patients
18 | Waveforms | Yes
19 | Trends | Yes
20 | Strip Chart | Yes
21 | Remote Operation | From bedside monitor
22 | Alarm - Audio and Visible | Yes
23 | Alarm Lamp and flashing | Yes
24 | Central Station Must Be Compatible with the selected Bed Side Monitor | MUST BE INTERFACED WITH BED SIDE MONITOR FROM SAME MANUFACTURER
Company Stamp |  |  |  | Signature

## ITEM 88
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
88 | MIC10058 | 0 | CENTRAL STATION PHYSIOLOGICAL 12 BEDS LICENCE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | Adult , Pediatric , NICU
2 | CENTRAL STATION | Yes MEDICAL GRADE
3 | No. of Patient | Liscence for 12 Beds
4 | Display Monitor | High resolution color not less than 19 Inch , Flat screen TFT. Type
5 | Number Of Monitors | TO FIT ALL LISENCES REQUESTED
6 | Trending Full Disclosure | Not Less Than 72 Hrs.
7 | Operation | Fixed key and menu driven/touch screen, or other - Specify
8 | Communication | Network, central station, remote bedside monitor
9 | Bed To Bed Communication | Yes
10 | Configurable Reporting | Yes
11 | ECG interpretation | Yes
12 | Computer medical grade | YES
13 | keyboard and mouse. | Wireless or wired Specify
14 | RECORDER | 4 Channel, Digital
15 | Recording | Yes, One for the system
16 | Laser Printer | Yes, One for the system
17 | Data | All, Patients
18 | Waveforms | Yes
19 | Trends | Yes
20 | Strip Chart | Yes
21 | Remote Operation | From bedside monitor
22 | Alarm - Audio and Visible | Yes
23 | Alarm Lamp and flashing | Yes
24 | Central Station Must Be Compatible with the selected Bed Side Monitor | MUST BE INTERFACED WITH BED SIDE MONITOR FROM SAME MANUFACTURER
Company Stamp |  |  |  | Signature

## ITEM 89
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
89 | MIC10060 | 0 | CENTRAL STATION PHYSIOLOGICAL 20 BEDS LICENCE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | Adult , Pediatric , NICU
2 | CENTRAL STATION | Yes MEDICAL GRADE
3 | No. of Patient | Liscence for 20 Beds
4 | Display Monitor | High resolution color not less than 19 Inch , Flat screen TFT. Type
5 | Number Of Monitors | TO FIT ALL LISENCES REQUESTED
6 | Trending Full Disclosure | Not Less Than 72 Hrs.
7 | Operation | Fixed key and menu driven/touch screen, or other - Specify
8 | Communication | Network, central station, remote bedside monitor
9 | Bed To Bed Communication | Yes
10 | Configurable Reporting | Yes
11 | ECG interpretation | Yes
12 | Computer medical grade | YES
13 | keyboard and mouse. | Wireless or wired Specify
14 | RECORDER | 4 Channel, Digital
15 | Recording | Yes, One for the system
16 | Laser Printer | Yes, One for the system
17 | Data | All, Patients
18 | Waveforms | Yes
19 | Trends | Yes
20 | Strip Chart | Yes
21 | Remote Operation | From bedside monitor
22 | Alarm - Audio and Visible | Yes
23 | Alarm Lamp and flashing | Yes
24 | Central Station Must Be Compatible with the selected Bed Side Monitor | MUST BE INTERFACED WITH BED SIDE MONITOR FROM SAME MANUFACTURER
Company Stamp |  |  |  | Signature

## ITEM 90
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
90 | MIC10069 | 0 | BEDSIDE PHYSIOLOGICAL MONITOR 14 BED STATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION AREA | Adult , Pediatric , NICU
2 | CENTRAL STATION | Yes
3 | No. of Patient | Liscence for 16 Beds
4 | Capability | upgradable to 24
5 | Display Monitor | High resolution color not less than 19 Inch Medical grade, Flat screen TFT. Type
6 | Number Of Monitors | Two screen with configurable Number of patient , Specify
7 | Trending Full Disclosure | Not Less Than 72 Hrs.
8 | Operation | Fixed key and menu driven/touch screen, or other - Specify
9 | Communication | Network, central station, remote bedside monitor
10 | Bed To Bed Communication | Yes
11 | Configurable Reporting | Yes
12 | ECG interpretation | Yes
13 | Computer medical grade | YES
14 | keyboard and mouse. | Wireless or wired Specify
15 | RECORDER | 4 Channel, Digital
16 | Recording | Yes, One for the system
17 | Laser Printer | Yes, One for the system
18 | Data | All, Patients
19 | Waveforms | Yes
20 | Trends | Yes
21 | Strip Chart | Yes
22 | Remote Operation | From bedside monitor
23 | Alarm - Audio and Visible | Yes
24 | Alarm Lamp and flashing | Yes
25 | Central Station Must Be Compatible with the selected Bed Side Monitor | MUST BE INTERFACED WITH BED SIDE MONITOR FROM SAME MANUFACTURER
26 | MONITORS , Bedside . | 14 Monitors
27 | Screen Size | Not less than 19 Inch Medical grade TFT Type
28 | Mounting Bracket | Yes, Swivel base, wall or ceiling mounting according to Hosp. requirements.
29 | Wired remote control pad or fixed | Yes, with each monitor- specify.
30 | PARAMETERS TO BE MONITORED
31 | ECG / Resp , Heart Rate | Yes, 5 lead
32 | Spo2 | Yes , Nellcor Oximax OR Massimo set
33 | Dual Temperature | Yes
34 | NIBP | Yes
35 | Dual IBP | Yes
36 | OTHER SPECIFICATION | FDA, CE , HL7 ICU information system
37 | Monitor Type | Modular with ability to provide continuous monitoring during transport in the patient
38 | Network port | Yes
39 | Interface | USB
40 | TRENDING Hrs. | Not less than 72 hrs. for all monitored parameters
41 | Haemodynamic & Drug Calculation | Yes
42 | Advance Arrhythmia & St Analysis | Yes
43 | All accessories required for the above parameters | Adult : specify what will be provided in details
44 | EEG : with accessories for 16 patient | Included - Qty : 1
45 | Etco2 : with accessories for 16 patient | Included - Qty : 1
46 | BIS : with accessories for 16 patient | Included - Qty : 1
47 | Non Invasive Cardiac Output | Itemized - Qty : 1
48 | Cardiac Output / PICO | Itemized - Qty : 1
Company Stamp |  |  |  | Signature

## ITEM 91
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
91 | MIC20001 | 0 | ANALYZER NITRIC OXIDE NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LIGHT WEIGHT, PORTABLE DESIGN
2 | CONTINUOUS DELIVERY OF INHALED NITRIC OXIDE, DOSING-MONITORING SYNCHRONIZATION , CLOSED LOOP TYPE
3 | LONG BATTERY LIFE: LEAD ACID TYPE WITH AUTO RECHARGE 5 HOURS FULLY CHARGED
4 | PUSH BUTTON ZEROING FOR CALIBRATION
5 | SHALL BE COMPATIBLE WITH HIGH FREQUENCY VENTILATION
6 | BIDDER SHALL SPECIFY THE SIZE AND TYPE OF DISPLAY
7 | VISUAL INDICATORS ON DISPLAY AT LEAST:
8 | RANGE OF NO AND NO2 IN PPM
9 | O2 CONCENTRATION
10 | NO DELIVERY FLOW RATE IN LPM
11 | BATTERY INDICATOR
12 | SAMPLE LINE OCCLUSION
13 | DIMENSIONS : 25.4 CM H X 30.5 CM W X 12.7 CM D
14 | WEIGHT <=6.5KG
15 | COMPATIBLE WITH STANDARD AND PORTABLE NITRIC OXIDE CYLINDERS
16 | GASES DETECTED: NITRIC OXIDE (NO), NITROGEN DIOXIDE (NO2), OXYGEN (O2)
17 | CONCENTRATION RANGE:
18 | NITRIC OXIDE (NO) :P 0 TO 100 PPM, DOSE SET BY TOUCH SCREEN IN 0.1PPM RESOLUSION AT ALL TREATMENT RANGE
19 | NITROGEN DIOXIDE (NO2): 0 TO 12 PPM
20 | OXYGEN (O2): 5 TO 100%
21 | NO MASS FLOWMETER : 0 TO 2 LPM,10 CC/MIN RESOLUTION
22 | BIDDER SHALL LIST ALL THE AVAILABLE ACCESSORIES
23 | COMPLIANCE WITH STANDARDS & LEGISLATION:
24 | SHOULD HAVE A FDA APPROVAL AND/OR CE MARK & SFDA REGISTRATION, WHERE APPLICABLE.LIST ANY OTHER INTERNATIONAL STANDARDS (CE, UL, TUV, CSA), IF ANY.
25 | ALL OTHER BASIC ACCESSORIES DEEMED NECESSARY THAT ARE NOT MENTIONED IN THIS SPECIFICATION BUT ARE REQUIRED FOR FULL FUNCTION AND HIGHEST CLINICAL OUTCOMES AND OUTPUT OF THE EQUIPMENT MUST BE INCLUDED
Company Stamp |  |  |  | Signature

## ITEM 92
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
92 | MIC20008 | 0 | INCUBATOR PREMIUM NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | INCUBATOR CONVERTED TO WARMER | Yes
2 | TYPE | Mobile-Microprocessor controlled
3 | TEMPERATURE CONTROL | Yes
4 | RANGE, C degree | Specify
5 | Air (increment) | 20-42 approx.
6 | Skin (increment) | 36-38 approx.
7 | TEMPERATURE DISPLAY
8 | Type | Digital
9 | Temperature range 0c for Air | 20 - 42
10 | Temperature range for Skin | 22 - 40
11 | ALARMS
12 | High air temp | Yes
13 | Low air temp | Yes
14 | High skin temp | Yes
15 | Low skin temp | Yes
16 | Fan failure | Yes
17 | Sensor failure | Yes
18 | Power failure | Yes
19 | BACKUP THERMOSTAT | Yes
20 | HEATER POWER INDICATOR | Yes
21 | HAND PORTS | Yes
22 | TUBING PORTS | Yes
23 | O2 INLET PORTS CONTROLLERS | Yes
24 | SUPPLEMENTAL HUMIDITY | Yes , Up to 95%
25 | INTERNAL NOISE, dB | Less Than 40 dB
26 | DOUBLE WALL CANOPY | Yes
27 | Handling Flaps
28 | Lars. Front Flap | Yes
29 | WEIGHING DIGITAL SCALE | Yes
30 | BEND ANGLE | Variable
31 | Trendeln Burs | Yes
32 | Fowler mechanism | Yes
33 | SEPARATE X-RAY FIL;M | Cassette compartments, Yes
34 | OVERHEAD SHELF FOR MONITORS | Yes
35 | EQUIPMENT RAIL SYSTEM | Yes
36 | I.V. ROD TELESCOPIC ASSEMBLY | Yes
37 | BASE COMPARTMENT WITH SHELVES | Yes
38 | ANTI-STATIC CASTORS | 4 (All with locking device)
39 | POWER SUPPLY | 220 Volt - 60 Hz.
40 | WATER TANK | Yes
41 | Screen Display , Type & Size | LCD , Other Specify
42 | Display Temp ( different modes ) humidity , O2 Concentration , babys weight | Yes, Other Specify
43 | ACCESSORIES
44 | Asperators Assembly with clamping Device | Yes
45 | Oxygen /Air Blender Assembly | Yes
46 | Oxygen Flow Meter Assebly | Yes
47 | 0 to 5 liter /minute | Yes
48 | PMG Oxygen & Medical connecting oses. | Yes
49 | Oxygen Monitor AND analyser | Yes
50 | Infant skin temp probe, | Yes
51 | Oxygen Hood for Neonate | Yes
52 | Humidification Pads | Yes pack / specify
53 | Air Intake filler pack | Yes / specify
54 | Mattress foam - staple - check covered. | Yes
55 | Mattress covers reusable | Yes
56 | I.V. Rod | Yes
57 | Temprature Probe | Yes
58 | SAFETY STANDARDS | IEC-601Part 1&2, EN 46000 or others.
59 | REGULATORY COMPLIAANCES | UL, CE, ISO FDA or others.
60 | THE INCUBATOR ELECTRONIC & WIRE SHOLD | Yes
61 | BE WATER BROVE | Yes
62 | BATTERY BUILLT IN 8HR RECHARGABLE | Yes
63 | EASE TO SAPARATED IN PART | Yes
64 | Can be Operated From Both Sides | Yes
65 | OTHER SPECIFICATIONS | All access needed
Company Stamp |  |  |  | Signature

## ITEM 93
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
93 | MIC20009 | 0 | INCUBATOR STANDARD NEONATAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Mobile-Microprocessor controlled
2 | TEMPERATURE CONTROL
3 | RANGE, C degree
4 | Air (increment) | 20- 42 approx.
5 | Skin (increment) | 36 - 38 approx.
6 | TEMPERATURE DISPLAY
7 | Type | Digital
8 | Temperature range 0c for Air | 20 - 42
9 | Temperature range for Skin | 22 - 40
10 | ALARMS
11 | High air temp | Yes
12 | Low air temp | Yes
13 | High skin temp | Yes
14 | Low skin temp | Yes
15 | Fan failure | Yes
16 | Sensor failure | Yes
17 | Power failure | Yes
18 | BACKUP THERMOSTAT | Yes
19 | HEATER POWER INDICATOR | Yes
20 | HAND PORTS | 4
21 | TUBING PORTS | Yes
22 | SUPPLEMENTAL HUMIDITY | Yes , Up to 95%
23 | INTERNAL NOISE, dB | Less Than 40 dB : Specify
24 | DOUBLE WALL CANOPY | Yes
25 | HNADLING FLIP
26 | WEIGHING DIGITAL SCALE | Yes
27 | BEND ANGLE | Variable
28 | Trendelnburg | Yes
29 | FOWLER MECHANISM | Yes
30 | SEPARATE X-RAY FILM | Cassette compartments, Yes
31 | EQUIPMENT RAIL SYSTEM | Yes
32 | I.V ROD TELESCOPIC ASSEMBLY | Yes
33 | BASE COMPARTMENT WITH DRAWER | Yes
34 | ANTI -STATIC CASTORS | 4 (All with locking device)
35 | WATER TANK | Yes
36 | Screen Display , Type & Size | LCD , Other Specify
37 | Display Temp ( different modes ) humidity , babys weight | Yes , Other Specify
38 | ACCESSORIES
39 | Asperators Assembly with clamping Device | Yes
40 | Mattress covers reusable | Yes
41 | Temprature Probe | Yes
42 | Infant skin temp probe, | Yes
43 | Humidification Pads | Yes pack / specify
44 | Air Intake filler pack | Yes / specify
45 | SAFETY STANDARDS | IEC - 601Part 1&2
46 | REGULATORY COMPLIANCES | CE , FDA
Company Stamp |  |  |  | Signature

## ITEM 94
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
94 | MIC20010 | 0 | INCUBATORS TRANSPORT INFANT WITH VENTILATOR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MOBILE OR TRANSPORT | Transport
2 | TEMPERATURE CONTROL
3 | Range, C
4 | Air | 22 - 38 approx.
5 | Skin | 34 - 42
6 | TEMPERATURE DISPLAY
7 | Air, display type | Digital
8 | Range, C | 30 - 45 approx.
9 | Skin, display type | Digital
10 | Range, ?C | 17 - 45
11 | ALARMS
12 | High air temp | Yes
13 | Low air temp | Yes
14 | High skin temp | Yes
15 | Low skin temp | Yes
16 | Fan failure | Yes
17 | Sensor failure | Yes
18 | Power failure | Yes
19 | BACKUP THERMOSTAT | Yes
20 | HEATER POWER INDICATOR | Yes
21 | HAND PORTS | Yes , at least 3 ports
22 | TUBING PORTS | Yes,1 respiratory, 2 accessory
23 | O2 INLET PORTS/ CONTROLLERS | Standard/Optional, Specify
24 | SUPPLEMENTAL HUMIDITY, ADJUSTABLE | Yes
25 | INTERNAL NOISE, dB
26 | LINE POWER, VAC | 220-240 V 60Hz
27 | RECHARGABE BATTERY, | Yes
28 | Type | 12 V sealed lead acid/ acid gel
29 | DOUBLE WALL VERTICAL HOOD TO MATTRESS DISTANCE, cm (in) | Not less than 21
30 | TROLLEY | Yes
31 | CYLINDER BRAKET | Yes
32 | IV POLE | Yes
33 | DISPOSBLE MATRESS COVER | Yes Qty. 50 Pcs
34 | TYPE | Transport
35 | CONTROLS
36 | Inspiratory time, sec | 0.2-3
37 | Expiratory time, sec | 0.4-30
38 | Ventilation | Time cycle, pressure-limeted, continuous flow
39 | Respiratory frequency, breaths/min | 0 150 bpm or more
40 | Inspiratory flow, L/min | 0-180
41 | Peak inspiratory pressure, cm H2O | 60 cm H2O
42 | CPAP/PEEP, cm H2O | 0-20 or more
43 | FiO2 | 0.21-1 (21 100%)
44 | Trigger mechanism | Flow/Flow pressure/ pressure
45 | OPERATING MODES
46 | Assist/control | Yes
47 | Manual inspiration | Yes
48 | IMV/SIMV | Yes
49 | SIMV pressure support | Yes
50 | Sigh | Yes
51 | Demand/flow | Yes
52 | MONITORED PARAMETERS
53 | PIP | Yes
54 | Expiratory volume | Yes /tidal minute
55 | FiO2 | Yes
56 | SpO2 | Yes
57 | MAP | Yes
58 | CPAP/PEEP | Yes
59 | Ventilation frequency | Yes
60 | Inspiratory time | Yes
61 | Expiratory time | Yes
62 | I:E ratio | Yes
63 | Reverse I:E ratio | Yes
64 | Inspiratory volume | Yes (tidal)
65 | ALARMS | Yes
66 | Low PIP | Yes
67 | Low battery | Yes
68 | Vent inoperative | Yes
69 | High PIP | Yes
70 | Low CPAP/PEEP | Yes
71 | Volume | Yes
72 | FiO2 | Yes
73 | Fail to cycle | Yes
74 | Gas supply loss | Yes
75 | Power failure | Yes
76 | AUXILIARY POWER | Yes
77 | Battery | 12/6 V recharable with built in charger
78 | Operating time | 12 hrs approx.
79 | LINE POWER, VAC | 220-240 V 60 Hz
80 | Power Cord | 3Pin BT, 1-2 Meter Length
81 | REGULATORY COMPLIANCES | ISO, FDA, CE, or Others. (Specify in offer).
82 | SAFETY STANDARD | UL, IEC, CSA or other international Standard, (Specify in offer).
83 | ACCESSORIES | All standard Accessories, [Specify in offer]
84 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Signature

## ITEM 95
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
95 | MIC20011 | 0 | INTUBATION DIFFICULT VIDEO NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Video Laryngoscope For difficult airway intubation
2 | Design | light weight & portable
3 | The system shall be based on the latest technological advances in difficult intubations | YES
4 | The unit shall be designed to provide a clear view of the vocal cords during intubation | YES
5 | The unit shall simplify intubation of the difficult airway | YES
6 | Trauma Airway | Yes
7 | Range in blade sizes allow coverage patient weights | Yes
8 | It shall have 60° angled blades | YES
9 | It should have pre-shaped stylet for quick intubation process | YES
10 | It should utilize minimum force for intubation | YES
11 | It should have anti-fog feature | YES
12 | The unit shall have integrated high resolution camera at good Position on blade to protect from blood and secretion | YES
13 | The unit shall mount on mobile carts and castors | YES
14 | Intubation of cervical spine immobilization | Yes
15 | Difficult Airway management and routine intubation | Yes
16 | Screen Size not less than 7 inch | LCD OR TFT
17 | Blades Size
18 | SIZE 0 | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
19 | SIZE 1 | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
20 | Made of medical grade plastic or stainless steel for less trauma handlingo teeth and soft tissue and easy | Specify
21 | Built in LED light for illumination | Yes
22 | Rechargable Battery on system | Yes
23 | Stylet | Qty:2
24 | Accessories | Any accessories, options and consumable items necessary to operate the offered system(s) must be clearly identified and priced separately
25 | SIZE 0 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 120 Itemized and must be Optional Price
26 | SIZE 1 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 120 Itemized and must be Optional Price
Company Stamp |  |  |  | Signature

## ITEM 96
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
96 | MIC20017 | 0 | RESUSCITATOR & WARMER INFANT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | INFANT RESUSCITATOR - RADIANT WARMER | WITH THE FOLLOWING INCLUDED :
2 | MOBILE WITH LARGE LOCKABLE CASTORS | yes
3 | FULLY MICROPROCESSOR CONTROLLED | yes
4 | LARGE LED / LCD to disply spo2,heart rate,temp.baby weight option abgare time ,trends. | yes
5 | UNIT TRAY MADE FROM RADIO-LUCENT MATERIAL | yes
6 | TRAY WITH TENDLENBURG FOWLER SYSTEM | yes
7 | TRAY SHOULD HAVE REMOVABLE TRANSPARENT WALL + X-ray tray | yes
8 | SYSTEM TO HAVE DIGITAL DISPLAY OF TEMPRATURE | yes
9 | UNIT TO BE WITH INTEGRATED IV POLEMONITOR TRAYOBSERVATION LAMP | yes
10 | INSTRUMENT SHELF | yes
11 | RESUSCITATOR WITH POSITIVE PRESSURE TO THE MASK OR ENDO TRACHEAL TUBE WITH MANUAL CONTROL AND FREE EXPIRATION | yes
12 | MINIMUM & MAX. PRESSURE VALUE SHOULD BE PRESET AND CLEARLY INDICATED ON THE MONOMETER | yes
13 | RESUSCITATOR SHALL OPERATE WHEN CONNECTED TO AIR / O2 MIXER OR ANY FLOW METER | yes
14 | UNIT MUST HAVE INTEGRATED APGAR TIMER WITH AUDIOVISUAL NOTIFICATION TO REMIND FOR INFANT CHECK AFTER CERTAIN PERIOD OF TIME AS PER APGAR SCHEME | yes
15 | SYSTEM SHALL HAVE HEATER built in | yes
16 | ELECTRICAL ELEVATING BED | yes
17 | SLOW SUCTION UNIT WITH CONTINOUS VACUUM ADJUSTMENT FROM 0 - 400 mmHg | yes
18 | TEMP. PROBE SHALL HAVE RANGE OF 32 - 42 C | yes
19 | SYSTEM MUST HAVE STARTUP CHECK-UP FOR :BABY PROBE FAILUREHIGH & LOW TEMP.POWER FAILUREHEATER FAILURESUCTION LIMITS | yes
20 | HEATER POWER MUST BE ADJUSTED AUTOMATICALLY BY SKIN TEMP PROBE | yes
21 | pulse oximeter builtt in (spo2) | yes
22 | neonatal mask | yes (3 sizes)
23 | blender | yes
24 | oxygen ,air cylender | yes size E
25 | UNIT MuST HAVE PROVESION FOR O2 CYLINDER
26 | POWER | 220 VLOTS
27 | FDA , CE | YES
28 | ITEMIZED PRICE FOR | TEMP. PROBES
Company Stamp |  |  |  | Signature

## ITEM 97
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
97 | MIC20020 | 0 | THERAPY HIGH FLOW NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Built in blender for mixing O2 and normal Air. | Yes
2 | Built in humidifier. | Yes
3 | Able to deliver 100 % relative humidity at all flow rates. | Yes
4 | Adjustable Temperature 33c up to 43c. | Yes
5 | Adjustable FiO2 21% up to 100%. | Yes
6 | Alarms for: blocked tube, low water , low gas and temperature variation. | Yes
7 | Triple lumen design for patient delivery tube. | Yes
8 | back up battery for 15 min. | Yes
9 | Display for temperature, FiO2 and flow rate. | Yes
10 | Has filter for Air and O2 sources. | Yes
11 | all accessories needed must be provided for complete practice (itemized price) | Yes
12 | Other specifications | FDA or CE marked
Company Stamp |  |  |  | Signature

## ITEM 98
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
98 | MIC20023 | 0 | VENTILATOR CPAP NASAL NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | VENTILATOR | nCPAP for neonate
2 | humidifire | yes
3 | disply | LCD or TFT 11 screen or more
4 | Breathes per Minute | 0 - 60 bpm or better
5 | Tidal Volume | 5 - 200 ml or better
6 | Adjustable tidal volume | yes
7 | Adjustable BpM | yes
8 | neonate settings | yes
9 | FiO2 | yes 21%-100% or other specify
10 | Low source gas pressure alarm | yes
11 | system failure alarm | yes
12 | Accessories | yes all needed accessories must be provided
13 | Power requirements | 220v /60hz
14 | OTHER SPECIFICATION | FDA , CE , APPROVED.
Company Stamp |  |  |  | Signature

## ITEM 99
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
99 | MIC20024 | 0 | VENTILATOR HFO NEONATE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | NEONATE -pediatric high frequency
2 | CONTROLS
3 | Inspiratory time, sec | 0.1- 2 or better
4 | tidle volume start from 2ml | 2ml
5 | Expiratory time, sec | 0.2 or better
6 | Ventilation | Time cycle, pressure- limited, continuous flow
7 | Respiratory Frequency, breaths/min | 0 - 270 bpm or more
8 | Inspiratory Flow, L/min | 0 - 30 or more
9 | CPAP/PEEP, cm H2O | 0-20 or more
10 | FiO2 | 0.21-1 ( 21 ? 100 % )
11 | Trigger mechanism | Flow / pressure
12 | MANUAL INSPIRATION | Yes
13 | OPERATING MODES
14 | Assist/control | Yes
15 | Spontaneous/CPAP | Yes
16 | IMV/SIMV | Yes
17 | pressure Support ventilation | Yes
18 | Others | Specify
19 | ADVANCE OPERATIONS MODE
20 | High Frequency/Oscillator Operating mode | Yes, Specify
21 | Technology, either Diaphragm or Electrical valve oscillation based. | Yes, Specify
22 | Respiration rate BPM or Hz. | Yes, Specify
23 | MONITORED PARAMETERS
24 | Screen Display Size & Type | TFT, LCD, LED, Specify
25 | PIP | Yes
26 | MAP | Yes
27 | PEEP | Yes
28 | Respiration Freq., breath/min | Yes [specify]
29 | Inspiratory time | Yes [specify]
30 | Expiratory time | Yes [specify]
31 | I:E ratio | Yes [specify]
32 | Inspiratory tidal volume, ml | Yes [specify]
33 | Expiratory minute volume, ml | Yes [specify]
34 | FiO2 | 21-100% either Built in or separate
35 | PATIENT ALARMS | Yes
36 | Hi/Lo minute volume | Yes
37 | Low inspiration pressure | Yes
38 | High PIP | Yes
39 | High PEEP | [specify]
40 | Loss of PEEP | [specify]
41 | FiO2 | either Built in or separate
42 | Apnea | Yes
43 | High continuous pressure /occlusion | Yes
44 | Inverse I :E | Yes
45 | High respiration rate | Yes
46 | EQUIPMENT ALARMS
47 | Fail to cycle | Yes
48 | Gas supply loss | Yes
49 | Power failure | Yes
50 | Low battery | Yes
51 | Vent inoperative | Specify
52 | Self-diagnostic | Yes
53 | COMPUTER INTERPHASE | Specify
54 | TRANSPORTABLE/Mobile | Mobile on stand
55 | BATTERY OR UPS | Yes
56 | Operating time./hr | Specify
57 | Rechargeable | Specify
58 | POWER SUPPLY
59 | VOLTAGE/ FREQUENCY | 220 VOLT / 60 Hz
60 | PATIENT CIRCUIT disp. | 100 set for each unit
61 | HEATED HUMIDIFIER WITH TEMPERATURE SENOR MODEL MR 850 FISHER AND PAYKEL OR EQUIVALENT . | Yes
62 | Reusable Chamber for Neonatal and Pre-mature . | Yes
63 | Connections and water traps for each patient set | Yes
64 | STANDARD ACCESSORIES | All standard accessories required for commissioning equipment
65 | OPTIONAL ACCESSORIES | Quote separate price for optional accessories
66 | Either built in or separate Air compressor | Yes
67 | Screen 10 Inch colored or more for waves and curve display, | Yes
68 | SAFETY STANDARD | International Standard, IEC 60601-1, CSA 22.2 , UL. Specify.
69 | OTHER SPECIFICATIONS | FDA, CE, ISO Approved
Company Stamp |  |  |  | Signature

## ITEM 100
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
100 | MIC20027 | 0 | WARMER INFANT WITH BASSINET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | INFANT RADIANT WARMING UNIT
2 | SAFE WARMING RATE ON BABY SKIN | YES
3 | SKIN TEMP SENSOR | YES
4 | SELF TEST FOR THE UNIT | YES
5 | HEATER STATUS ( ON / OFF ) | YES VISUAL
6 | UNIT WORKING MODE | AUTOMATI C TEMP CONTROLMANUAL TEMP CONTROL
7 | IRRADIANCE NEAR I.R | YES
8 | HEATING ELEMENT | HALOGEN BULB ORCERAMIC OR OTHER SPECIFY
9 | SET TEMP FEATURE | ELECTRONIC TYPE
10 | SET TEMP RANGE | 34 C - 38 C
11 | SAFETY OF BABY | SHOULD NOT RECIEVE TEMP MORE THAN 39 Degree FROM AIR SORROUNDING OR IF CONTACT TO UNIT METAL
12 | ALARM
13 | SET TEMP ALARM | YES
14 | TEMP OUT OF RANGE | YES
15 | POWER | 220 VOLTS
16 | OTHER SPECS | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 101
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
101 | MIF40003 | 0 | OZONE AND HYDROGEN PEROXIDE DISINFECTION SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | MOBILE
2 | MICROPROCESSOR | YES
3 | AUTOMATED | FULLY AUTOMATED
4 | WI-FI CONNECTION | YES
5 | TYPE OF SOLUTIONS | FLUID VAPORISING SYSTEM AND AN EMBEDDED OZONE
6 | GENERATION | OZONE GENERATION AND DETECTION
7 | REMOTE CONTROL | YES
8 | MOTION SENSOR | YES
9 | DISINFECTANT | USE H2O2 CONCENTRATION 7.5% HYDROGEN PEROXIDE
10 | SPACE DISINFECTION | 200 M3 OR MORE
11 | DISINFECTION METHOD | DRY MIST DIFUSSION
12 | HUMIDITY SENSOR | YES
13 | POWER SUPPLY | 220 V , 60 HZ
14 | BIO COMPATIBILITY ELECTRONIC DEVICES | YES
15 | REGULATORY COMPLIANCE | ISO , CE , SPECIFY
16 | SAFTY STANDARD | MET APPROVAL
17 | COMPLIANCE | NFT 72-281 (2009)
Company Stamp |  |  |  | Signature

## ITEM 102
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
102 | MIF40004 | 0 | SURFACE DECONTAMINATION SYSTEM H2O2 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SURFACE DECONTAMINATION SYSTEM - H2O2 , Hydrogen Peroxide Sterilant space up to 250 m3
2 | Features :
3 | Microprocessor Controlled
4 | Hydrogen peroxide vapor
5 | Specifications :
6 | System shall have User Configuration Programs: Specify
7 | Blower speed : 10-40 m3/h
8 | Pump Injection rates :1-16 grams/min.
9 | System shall have Audio visual Alarm.
10 | System shall include all needed Hoses and connections .
11 | Compliance wh standards & legislation:
12 | The system must comply wh the Electrical safety standards for electrical safety IEC-60601
13 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
14 | All electrical connections and plugs should be hospal grade and follow international, local and hospal requirements.
15 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and condions
16 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 103
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
103 | MIF50013 | 0 | CHAMBER ISOLATION PATIENT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mobile Isolation Unit | Yes
2 | Patient Controls | Yes
3 | Nurse Control | Yes
4 | Frame Mattress Fuctions | Head High/Low, Knee up& Down
5 | Dimensions (Outer dimensions) | L : 2285 mm W : 640 mm H : 695 mm approx
6 | Patient weight | Max 180 kg / 395 lb
7 | Patient Length | Max 198 cm
8 | Operating temperature | 0 to +40 °C (32 to +104 °F)
9 | Blower unit | CleanAir Chemical 2F (w/ display)
10 | Power source | Rechargeable Li-ion battery 14,4V 2,6 Ah
11 | Air exchange | More than 15 air exchanges per hour
12 | Operating negative pressure | Min 15 Pa relative to the environment
13 | Operating positive pressure | Min 15 Pa relative to the environment
14 | Operator ports | 8 ports for gloves, waste bag or sluice bag configurations
15 | Wire port | Membrane inlet for IV lines, monitoring cables and similar equipment
16 | Ventilator port | Generic port for all types of mechanical ventilator circuits
17 | Transparent Waterproof Protective Hard Shell | yes
18 | Crash Tested | yes
19 | Easy Transport Option | Ems System Compatible (Air/Sea/Land)
20 | Patient Retention/Safety Belt | Quick Release 6-point torso/hip restraint, chest strap, and upper / lower leg strap
21 | Strecher Adapter | yes
22 | Sun Shade | yes
23 | Storage Cover | yes
24 | Filter (A2B3 type) 4pcs. | yes
25 | FDA / CE approved | yes
Company Stamp |  |  |  | Signature

## ITEM 104
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
104 | MLB10003 | 0 | APHERESIS UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Fully automated apheresis unit for separating, collecting and re-infusion of blood components for donating or therapeutic purposes
2 | The unit should be microprocessor controlled with pre-programmed protocols. Specify.
3 | Automatic parameter optimization based on donor profile and customer needs
4 | Interface to the Laboratory Information System
5 | Recognize ISBT128 barcodes and the barcodes generated by the LIS
6 | Digital alphanumeric display and keyboard for:
7 | Entry of patient demographic information
8 | Specifying type of procedure to be performed
9 | Provide users with set-up information and apheresis set installation/removal guidance
10 | Monitoring of session progress and related parameters list.
11 | Troubleshooting and system performance monitoring and verification
12 | Display of alarming conditions with user instructions
13 | Variable anticoagulant control (~ 1:8 – 1:16)
14 | Microprocessor controlled pumps for inlet, collect, anticoagulant and plasma lines with fail safe mechanisms. Provide details.
15 | Single and dual needle operation. Specify collection parameters for both modes (flow type, inlet rate, extra-corporeal volume, etc.)
16 | Component separation by centrifugation. Specify centrifugal speed range.
17 | The unit should be designed for all apheresis applications:
18 | Collection
19 | Exchange
20 | Therapeutic removal
21 | Shall list all the safety precaution and features
22 | Compliance with standards & legislation:
23 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
24 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
25 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
26 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
27 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 105
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
105 | MLB10027 | 0 | CABINET FLAMMABLE STORAGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Welded construction from 1 mm thick steel or close similar
2 | The unit shall have a tough, chemical resistant finish
3 | Supplied with adjustable spillage trays
4 | Welded liquid tight lift-out base spillage sump
5 | Door with latching mechanism
6 | Internal fire barrier around door opening
7 | Strong chrome lever locking handle
8 | High visibility tough powder coated yellow finish (preferably)
9 | Appropriate warning label and cabinet color to be as per international standards and codes
10 | Dimensions: 600 x 600 x 1200 mm approx.
11 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
Company Stamp |  |  |  | Signature

## ITEM 106
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
106 | MLB10035 | 0 | CART UTILITY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CART UTILITY , General Purpose cart.
2 | Features and Specifications :
3 | Supported by four uprights with tops welded closed
4 | Four-sided guards rails
5 | Tubular Strong Frame
6 | Worktop and handles :18 gauge S.S.
7 | Capacity : 80-100KG.
8 | Top and shelves made of 18 gauge Stainless steel with edges bent and hemmed.
9 | 16 gauge 25mm (1") O.D
10 | Stainless steel legs
11 | Dimension : approx.. 700x 400x 850 mm
12 | Three shelves
13 | Integrated push handles
14 | swivel castors
15 | 76mm (3") ball bearing .
16 | Antistatic
17 | locking Capability.
18 | Directional
19 | A sample of items, including its parts shall be submitted for approval before an order is placed ,Approved samples shall be used as standards of finish and workmanship.
Company Stamp |  |  |  | Signature

## ITEM 107
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
107 | MLB10045 | 0 | CHAIR BLOOD DRAWING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Wide with padded arm support that rotates 360 degrees.
2 | Suitable for right or left hand use.
3 | Arm rest adjusts to any size/height patient
4 | Stain resistant upholstery
5 | The armrest should lock into place at the desired location.
6 | Capability to carry up to 200Kg
7 | Storage drawer
8 | Non slip rubber feet
9 | Color to be coordinated with interior design / Engineer.
10 | The offered equipment shall have an approved international certificate( CE,FDA,TUV, etc..)
Company Stamp |  |  |  | Signature

## ITEM 108
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
108 | MLB20002 | 0 | ANALYZER AUTOMATED BLOOD GAS PH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | To incorporate a LCD screen and alphanumeric keyboard for user interfacing. Provide detailed specs such as touch control, size, color or mono, resolution, etc..
2 | To incorporate a built-in thermal printer. Provide detailed specs such as, paper size and type (roll, Z-fold, etc.), resolution, printing format and information (data, tables, patient demographics, normal values, abnormal flags, etc..
3 | Measured blood gas parameters to include. Specify measurement range, units, accuracy as well as any other relevant information:
4 | pH
5 | pCO2
6 | pO2
7 | CO-Oximetry panel shall be integrated.
8 | Calculated blood gas parameters to include. Specify measurement range, units, accuracy as well as any other relevant information:
9 | SO2 (Oxygen saturation of hemoglobin in blood
10 | HCO-3 (Concentration of hydrogen carbonate in plasma
11 | tO2 (total oxygen concentration in blood
12 | AaDpO2 (Alveolar-arterial oxygen tension difference for arterial blood)
13 | SBE and ABE (standard and actual base excess)
14 | SBC (standard bicarbonate)
15 | Plasma tCO2 (Concentration of total CO2 in plasma)
16 | PT pH (pH of plasma at patient temperature)
17 | PT pCO2 (CO2 tension in blood at patient temperature)
18 | PT H+ (Concentration of hydrogen ions in blood at patient temperature
19 | PT pO2 (O2 tension in blood at patient temperature)
20 | Measured electrolytes. Specify measurement range, units, accuracy as well as any other relevant information:
21 | Sodium
22 | Potassium
23 | Calcium
24 | Chloride
25 | User entered data shall include (provide details):
26 | Patient demographics (name, ID, age, sex, height, weight)
27 | Technician ID
28 | Physician name
29 | Location (outpatient, in-hospital department or bed, etc.)
30 | Sample type and time taken
31 | Patient temperature
32 | Hb
33 | FiO2
34 | Automatic washing capability. Specify parameters.
35 | Hands-free sample administration (syringe or capillary tube)
36 | Fully automated aspiration, transport, analysis and processing of sample
37 | State sample volume in all applicable modes (normal, micro, etc.)
38 | State type, stability and accuracy of temperature control
39 | Specify test throughput (analysis time) from “READY” state as well as start up time from off to ready
40 | Automatic calibration. State levels, schedule (or time interval) and duration. Calibration interruption for STAT samples
41 | On-board QC management
42 | Barcode reading capability for sample and reagent handling. Automatic reagent tracking (level detection, expiry date, etc.)
43 | Built-in automatic maintenance program, preferably zero user maintenance
44 | State type of required gas mixture and method of delivery
45 | Internal troubleshooting (self diagnostic) software capability is an asset. Specify details.
46 | Computer connectivity for Bi-directional data, results and report transmission from and to the Hospital Information System including archiving capabilities
47 | The following features of the offered system are preferred. State any that apply to your system. If optional, Price separately:
48 | Sealed waste container
49 | On-board help function
50 | Remote diagnostics
51 | Data management capabilities
52 | Shall list all the safety precaution and features
53 | Compliance with standards & legislation:
54 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
55 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
56 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
57 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
58 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 109
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
109 | MMI110012 | 0 | X-RAY DIGITAL MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | GENERAL STANDARD
2 | (FDA 510 (K) Clearance OR CE Mark) & SFDA Registration | Yes
3 | MOBILE X-RAY UNIT
4 | Microprocessor Controlled | Yes
5 | Forward & Reverse Motor Driven | Yes, Self Propelled Handle Drive Control
6 | Max incline, degree | ( 5-7 ) Degree is Preferable
7 | Speed | Adjustable from ( 0.5 - 1.3 ) m/s
8 | AC line | Yes
9 | Line-voltage Compensator | Yes / Automatic
10 | Type of Power Sources | Both AC & Battery, The Generator Should Work on AC in Case of Battery Failure (Preferable)
11 | Power Requirements | 220 Volt / 60 Hz.
12 | Single or Dual Battery System | Two Independent Systems Provide Power for Driving Motors & Imaging Exposure (Preferable)
13 | Battery Type | Lead acid sealed lead / NiCad
14 | Battery Charging Time - ( From Empty to Full Charged ) | Not more than ( 10 ) hrs
15 | Low-Battery Indicator | Yes
16 | Battery Power Storage | Specify number of maximum oprating hours
17 | Power Cord Length, m | Not less than ( 2.5-5 ) m or More
18 | Wired Handheld Switch & Cordless Remote control Switch | Yes, Exposure By Using Remote Control for Optimal Radiation Protection
19 | Manual Movement in Case of Battery or Motor Failure | Yes
20 | Remote Control Operating range | Not Less than ( 7 ) m
21 | Anti-Collision Protection System or Equivalent | Yes
22 | System Weight | Specify
23 | Mobile X-ray can be Moved Forward & Backward by using the Bedside Drive Controls Located on the Collimator or on the Articulated Arm | Preferred
24 | X-RAY TUBE
25 | Rotating Anode | Yes
26 | Anode Heat Storage Capacity | Not Less Than ( 120 ) kHU
27 | Maximum output Voltage, kVp | Approx. ( 140 ) KV , ( ±10 )
28 | Nominal Focal Spot Value, mm | ( 0.8 / IEC ) or Dual Focal Spots
29 | Nominal Anode Input power | Not Less Than ( 30 ) KW
30 | Tube Movement & Angulation | Yes
31 | Clear view while moving (telescopic movement of the column or dropdown) | Yes
32 | Shortest Exposure Time | 2 ms or less
33 | Working Column Movement Range | Rotating ( 180 ) Degree
34 | X-RAY GENERATOR
35 | High Frequency X-Ray Generator | Yes
36 | Nominal power Rating | Not Less Than ( 30 ) KW
37 | Kv Range | Approx. ( 40 - 140 ) Kv ( ±10 ) step of 5 (as minimum)
38 | mA range | Not Less Than ( 300 ) mA
39 | mAs range | Min.: ≤ 0.4 , Max.: ≥ 320
40 | Increments | Step Up / Step Down / Continuous
41 | Digital Indicator meters | Yes, (If applicable) For KVp, mAs, mA, Dose
42 | X-RAY COLLIMATOR
43 | Type | Manual Collimator
44 | Rotation | Approx. ( ±90 ) Degree
45 | Aluminum filter, mm | ≤ (3.5) mm
46 | SID Range, cm | Min.: ≤ 68 , Max.: ≥ 200
47 | Centering indicator (light /laser) | Yes, for Quick & Easy Indication of SID
48 | LARGE DETECTOR ASSEMBLY
49 | Waterproof | Yes
50 | Shock Sensor | Yes
51 | Portable Digital Detector | Wireless
52 | Semiconductor material | Amorphous silicon ( A-Si )
53 | Scintillator | Cesium iodide ( Csl )
54 | Detector Active Area Size | ≥ ( 35 x 43 ) cm, Approx.
55 | Image Depth / A/D Conversion / Acquisition Depth | ( 14 ) Bit
56 | Pixel size | 150 Micron or less is Preferred
57 | Matrix Size / Resolution | Not Less Than ( 2022 X 2022 )
58 | Detective Quantum Efficiency ( DQE ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
59 | Modulation Transfer Function ( MTF ) | Not Less Than ( 60% ) at ( 1 lp/mm ) or Higher
60 | Detector Weight Including Battery | Less Than or equal ( 4 ) kg
61 | Max. Load Capacity On Lying Position | More Than ( 100 ) kg
62 | Rechargeable Battery (Built in or Removable) | Lithium-ion , in Case of Removable Second Battery is Required or equivalent
63 | Charging Time | Not More Than ( 4 ) hours for 100% Charge
64 | Battery Operation Time | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
65 | Charging Station | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
66 | Rigidity of Detector Housing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
67 | Detector Handle | Not Less Than ( 3 ) hours at Normal Operation ,Typical at ( 90 ) images per hour
68 | Detector Sharing | Not Less Than ( 3 ) hours at Normal Operation ,Typical at (90) images per hour
69 | PROCESING & VIEWING WORKSTATION (Operator Console)
70 | Acquisition Workstation | Built in
71 | LCD Touch Screen Monitor | Not Less Than ( 15 ) inch
72 | Radiation Dose Monitoring ( DAP ) | indicate
73 | Had Disk Storage | Not Less Than ( 3000 ) images
74 | Automatic Programming ( APR ) | Yes
75 | IMAGE POST PROCESSING FUNCTIONS
76 | Window Width & Leveling | Yes
77 | Gray Scale Invert / Annotation | Yes
78 | Image Rotate / Free Rotation | Yes
79 | Electronic L/R Marker | Yes
80 | Patient Edit / Emergency Exam | Yes
81 | NETWORKING & SECURITY
82 | DICOM Compatible | Yes
83 | DICOM Store and DICOM structured dose report | Yes
84 | Image auto transfer | Yes
85 | Query, retrieve, Send, Receive | Yes
86 | Modality worklist | Yes
87 | Storage commitment | Yes
88 | DICOM print | Yes
89 | Patient edit | Yes
90 | DICOM MPPS | Yes
91 | DICOM viewer on CD "Burn exam on CD with DICOM viewer" | Yes
92 | HIPAA Patient Data Security | Yes
93 | DATA PROTECTION AGAINST DELETION & export "Different privilege levels " | Yes
94 | Security Package | Yes
95 | Virus Protection | Yes
96 | The ability to interact with the processing workstation from remote PC to view and process studies within the same hospital | Yes
97 | ACCESSORIES
98 | Two light weight Aprons (0.5) Equivalent of ( Non Leaded Material ) | ( Two ) - (medium & large)
99 | One Moveable Detector holder | ( One ) with Vertical, Horizontal & Swivel Movement
100 | Add-on specifications on top of base specs | Vendor input
101 | Extra detector (small)
102 | Extra detector (large)
103 | Above specs with telescopic column movement for X-Ray tube (if available)
104 | Extra detector (small) with exchange price
105 | Extra detector (large) with exchange price
106 | One Mobile Apron Hanger, and must hold 5 aprons
107 | Please specify Add-on specifications on top of base specs
Company Stamp |  |  |  | Signature

## ITEM 110
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
110 | MMI120001 | 0 | PROTECTION X-RAY APRON |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | X-ray protective apron made of flexible lightweight material. Specify weight and length for each size
2 | Wrap around model to provide 0.50 mm protective lead equivalency in the front and 0.3 mm in the back
3 | Wide shoulder cut to provide wide weight distribution and reduce inflicted fatigue
4 | Fastening shall be easy and reliable (Velcro type fasteners for example). Specify.
5 | Set of three different sizes
6 | Specify all available colors.
7 | Features ;
8 | High Quality alloy.
9 | Easily cleaned front
10 | Non-slip backing
11 | Stain-resistant binding
12 | Permit hanging
13 | Protective multi-layered
14 | Provides protection secondary (scatter) radiation.
15 | radiological lab inspected and certified.
16 | Souple, drapes comfortably
17 | Lightweight, easy handling
18 | Life cycle and store time :7-10 years specify
Company Stamp |  |  |  | Signature

## ITEM 111
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
111 | MMI140003 | 0 | PUMP INFUSION & SYRINGE MRI COMPATIBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Single channel IV volumetric infusion pump MRI compatible for use during MRI examination to all patient populations (adult, pediatric and neonates)
2 | Made of materials compatible with MRI up to 3 Tesla, bidder to specify
3 | Microprocessor controlled, easy to use, with large, digital LCD alphanumeric display of parameters and alarms
4 | Variable rate ranging from
5 | 0.1 to 99.9 ml/hr. (or better) in 0.05 mL/hr. increments in micro mode
6 | 0.1 to 999.9 ml/hr. (or better) in 1 mL/hr. increments in normal mode
7 | 5 % accuracy or better. State whether a dedicated infusion set is required to maintain the specified accuracy range. If open system, state so.
8 | Dose rate calculation
9 | Variable volume-to-be-infused from 1 to 10,000 mL
10 | Variable infusion time from 1 minute to 60 hours or better
11 | Keep Vein Open (KVO) rate 1 to 5 ml/hr. (specify)
12 | Piggyback capability (primary / secondary)
13 | Enteral feeding program
14 | Digitally displayed parameters to include:
15 | Infusion rate
16 | Volume to be infused
17 | Total infused volume
18 | Remaining / elapsed infusion time
19 | Battery / AC operation
20 | Running indicator
21 | Time to dose
22 | Dose remaining
23 | Reservoir volume
24 | Alarming condition when active, with indication of alarm type or code
25 | Back pressure monitor / indicator. If variable occlusion pressure, specify pressure range and default value
26 | Programming modes shall include but not be limited to:
27 | Normal mode
28 | Micro mode
29 | Ramp up
30 | Ramp down
31 | Primary, secondary and sequential
32 | Bolus
33 | Combinations of above listed modes (list)
34 | Audiovisual alarms shall include but not be limited to the following:
35 | Infusion set installation and integrity
36 | Door open
37 | Air in line
38 | Line disconnection or free flow (sudden drop in back pressure)
39 | Occlusion pressure pre-alarm (upstream and downstream)
40 | Occlusion pressure (upstream and downstream)
41 | Near end of infusion
42 | End of infusion
43 | Empty fluid container
44 | Unlocked container
45 | Low battery pre-alarm
46 | Discharged battery
47 | Internal malfunction
48 | Other, specify
49 | Data log capability and data port for data transmission, display and printing. Any required software for such function shall be included.
50 | Logged data to include:
51 | Settings
52 | Alarms
53 | Errors
54 | Other (state)
55 | Safety features shall include but not be limited to:
56 | Self test at start-up
57 | Nurse call interfacing capability
58 | Splash proof design
59 | Auto priming
60 | Adjustable alarm volume. No permanent silencing shall be possible.
61 | Keypad lock
62 | Impossibility to improperly install infusion set
63 | Free flow prevention system
64 | Last parameter setting retention
65 | Additional features (if available) shall be listed with their corresponding specs. Features such as:
66 | Preset drug labels
67 | User parameter setting storage memory
68 | Programmable profiles for different patients / areas
69 | Air trapping capability with air accumulation quantity measurement
70 | Other (specify)
71 | IV stand mounting clamp shall be included
72 | Line and rechargeable battery operation
73 | Battery autonomy of 3 hrs. or more when fully charged. Specify:
74 | Battery type and characteristics (voltage and current capacity)
75 | Autonomy at 10 mL/hr.
76 | Recharging time from depleted to 90%
77 | Compliance with standards & legislation:
78 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
79 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
80 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
81 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
82 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 112
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
112 | MMI90003 | 0 | ULTRASOUND 4D GENERAL STATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | High end, high performance color imaging system with 4D imaging capabilities for abdominal, vascular, obstetrics, gynecology, neonatal, and small parts applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing , Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R/W Drive and USB
22 | The system shall be suitable for the following applications:
23 | Abdominal
24 | Obstetrical
25 | Gynecological
26 | Musculoskeletal
27 | Vascular
28 | Urological
29 | Small Parts and Superficial
30 | Pediatric and Neonatal
31 | Transcranial Doppler
32 | The system shall possess capabilities for the following scanning methods:
33 | Electronic Sector
34 | Electronic Convex
35 | Electronic Linear
36 | The system shall be capable of accommodating the following transducer types:
37 | Phased array sector
38 | Convex
39 | Micro-convex
40 | Linear
41 | Operating modes shall include but not be limited to:
42 | 4D
43 | B-Mode
44 | M-Mode
45 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
46 | Color Flow Mode
47 | Power Doppler Imaging
48 | PW Doppler
49 | M-Color Flow Mode
50 | Steerable CW Doppler
51 | Trapezoidal Imaging
52 | Optional pan view
53 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
54 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
55 | Tilt/rotate adjustment
56 | Digital Brightness/Contrast Adjustment
57 | Integrated Task Light
58 | Slave monitor :
59 | A second slave monitor Wi-Fi connected with the scanner shall also be included in the offer, for patient viewing. All necessary mounting accessories shall be supplied.
60 | System must be offered with an above 19-inch high resolution,
61 | flat panel, medical grade
62 | monitor with wide viewing angles & good colour resolution.
63 | Resolution : 1024x768pixels or better
64 | The user interface panel shall possess ergonomic and friendly functional design, including:
65 | Alphanumeric keyboard for data entry with ergonomic key operations
66 | Interactive backlighting
67 | Indicator lights to identify activated keys
68 | The system cart shall possess features of high stability and maneuverability, including:
69 | On-board Storage for color-printer
70 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
71 | Multiple probe holders
72 | Gel dispenser holder
73 | Rear handle for transport maneuverability
74 | Video connection shall be integrated for film printing by laser camera.
75 | A color small format (A4), high resolution printer shall be included with the offer
76 | The system shall possess extensive display annotations including:
77 | Institution name
78 | Date / Time
79 | Patient Demographics
80 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
81 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
82 | Focal zone markers
83 | TGC Curve (LGC or Lateral Gain Compensation)
84 | Body patterns
85 | System messages display
86 | Trackball functionality status
87 | Biopsy guide line and zone
88 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
89 | h as IM
90 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
91 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
92 | Endo-vaginal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint. (Biopsy guide shall be quoted)
93 | Linear Probe for vascular, small parts, breast, neonatal and pediatrics: specify the frequency range, the number of elements, the convex radius and the FOV
94 | 4D obstetrical probe with highest performance specifications at time of delivery
95 | 4D Endo-vaginal probe with highest performance specifications at time of delivery
96 | Compliance with standards & legislation:
97 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
98 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
99 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
100 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
101 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 113
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
113 | MMI90004 | 0 | ULTRASOUND 4D GENERAL PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Advanced, high performance color imaging system for abdominal, vascular, obstetrics, gynecology, neonatal, and small parts applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing, Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R/W Drive and USB
22 | The system shall be suitable for the following applications:
23 | Abdominal
24 | Obstetrical
25 | Gynecological
26 | Musculoskeletal
27 | Vascular
28 | Urological
29 | Small Parts and Superficial
30 | Pediatric and Neonatal
31 | Transcranial Doppler
32 | The system shall possess capabilities for the following scanning methods:
33 | Electronic Sector
34 | Electronic Convex
35 | Electronic Linear
36 | The system shall be capable of accommodating the following transducer types:
37 | Phased array sector
38 | Convex
39 | Micro-convex
40 | Linear
41 | Operating modes shall include but not be limited to:
42 | B-Mode
43 | M-Mode
44 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
45 | Color Flow Mode
46 | Power Doppler Imaging
47 | PW Doppler
48 | M-Color Flow Mode
49 | Steerable CW Doppler
50 | Trapezoidal Imaging
51 | Optional pan view
52 | shear-wave elastography, and contrast-enhanced imaging , Specify.
53 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
54 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
55 | Tilt/rotate adjustment
56 | Digital Brightness/Contrast Adjustment
57 | Integrated Task Light
58 | Slave monitor :
59 | A second slave monitor Wi-Fi connected with the scanner shall also be included in the offer, for patient viewing. All necessary mounting accessories shall be supplied.
60 | System must be offered with an above 19-inch high resolution,
61 | flat panel, medical grade
62 | monitor with wide viewing angles & good colour resolution.
63 | Resolution : 1024x768pixels or better
64 | The user interface panel shall possess ergonomic and friendly functional design, including:
65 | Alphanumeric keyboard for data entry with ergonomic key operations
66 | Interactive backlighting
67 | Indicator lights to identify activated keys
68 | The system cart shall possess features of high stability and maneuverability, including:
69 | On-board Storage for color-printer
70 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
71 | Multiple probe holders
72 | Gel dispenser holder
73 | Rear handle for transport maneuverability
74 | Video connection shall be integrated for film printing by laser camera.
75 | A color small format (A4), high resolution printer shall be included with the offer
76 | The system shall possess extensive display annotations including:
77 | Institution name
78 | Date / Time
79 | Patient Demographics
80 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
81 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
82 | Focal zone markers
83 | TGC Curve (LGC or Lateral Gain Compensation)
84 | Body patterns
85 | System messages display
86 | Trackball functionality status
87 | Biopsy guide line and zone
88 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
89 | h as IM
90 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
91 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
92 | Endo-vaginal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint. (Biopsy guide shall be quoted)
93 | Linear Probe for vascular, small parts, breast, neonatal and pediatrics: specify the frequency range, the number of elements, the convex radius and the FOV
94 | Compliance with standards & legislation:
95 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
96 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
97 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
98 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
99 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 114
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
114 | MMI90006 | 0 | ULTRASOUND 4D OBGYN PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | High performance color imaging system for obstetrics and gynecology applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing, Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R/W Drive and USB
22 | The system shall be suitable for the following applications:
23 | Abdominal
24 | Obstetrical
25 | Gynecological
26 | Musculoskeletal
27 | Vascular
28 | Urological
29 | Small Parts and Superficial
30 | The system shall possess capabilities for the following scanning methods:
31 | Electronic Sector
32 | Electronic Convex
33 | Electronic Linear
34 | The system shall be capable of accommodating the following transducer types:
35 | Phased array sector
36 | Convex
37 | Micro-convex
38 | Linear
39 | Operating modes shall include but not be limited to:
40 | B-Mode
41 | M-Mode
42 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
43 | Color Flow Mode
44 | Power Doppler Imaging
45 | PW Doppler
46 | M-Color Flow Mode
47 | Steerable CW Doppler
48 | Trapezoidal Imaging
49 | Optional pan view
50 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
51 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
52 | Tilt/rotate adjustment
53 | Digital Brightness/Contrast Adjustment
54 | Integrated Task Light
55 | The user interface panel shall possess ergonomic and friendly functional design, including:
56 | Alphanumeric keyboard for data entry with ergonomic key operations
57 | Interactive backlighting
58 | Indicator lights to identify activated keys
59 | The system cart shall possess features of high stability and maneuverability, including:
60 | On-board Storage for color-printer
61 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
62 | Multiple probe holders
63 | Gel dispenser holder
64 | Rear handle for transport maneuverability
65 | Video connection shall be integrated for film printing by laser camera.
66 | A color small format (A4), high resolution printer shall be included with the offer
67 | The system shall possess extensive display annotations including:
68 | Institution name
69 | Date / Time
70 | Patient Demographics
71 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
72 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
73 | Focal zone markers
74 | TGC Curve (LGC or Lateral Gain Compensation)
75 | Body patterns
76 | System messages display
77 | Trackball functionality status
78 | Biopsy guide line and zone
79 | Basic Calculations , gynecology / obstetrical calculation package to be included , Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated
80 | parameter
81 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
82 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
83 | convex probe for Gyne/OB applications : specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
84 | Endo-vaginal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint. (Biopsy guide shall be quoted)
85 | Linear Probe for vascular, small parts, breast: specify the frequency range, the number of elements, the convex radius and the FOV
86 | Compliance with standards & legislation:
87 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
88 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
89 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
90 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
91 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 115
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
115 | MMI90008 | 0 | ULTRASOUND CARDIAC ADVANCE STATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The system shall be an advanced Echocardiography Ultrasound System suitable for Echo, Vascular, Superficial Parts, Transcranial and Contrast Imaging.
2 | Architecture
3 | Two Interactive graphical touch screens, 19” Flat monitor, adjustable height.
4 | System architecture should have the capability of multiple data streams simultaneously built for 2D, Panoramic and MPR.
5 | Digital broadband acoustic beamforming technology or multi-frequency technology (please specify).
6 | At least 50,000 Digital channels.
7 | At least 180 dB Full time input dynamic Range.
8 | An advanced algorithm for Adaptive Image Processing for noise and artifact reduction to improve tissue conspicuity.
9 | Automation
10 | Single button optimization key for adaptive gain compensation.
11 | Single button optimization of focal range position.
12 | Single button optimization key that automatically adapts system performance for different patient size and different flow states.
13 | Tissue Harmonic Imaging.
14 | Automatic Doppler trace analysis.
15 | Real-time Compound Imaging
16 | Automated Doppler (maintains scanning angle)
17 | Automatic Spectral Doppler drawing contours and Analysis
18 | Advanced Imaging Features
19 | TDI (Tissue Doppler Imaging), Tissue Harmonic Imaging, ECG capability and Contrast imaging.
20 | Cardiac 3D Quantification software to provide easy access to Live 3D, 3D Zoom, Full Volume and 3D Color data sets, viewing, cropping, slicing and quantification including distance measurements, area, Bi-plane LV Volume, Ejection Fraction (EF) and
21 | LV Mass
22 | Measurements of LV endocardial Volumes, Stroke Volume (SV) and true 3D ejection fraction (EF) using a semi-automated border detection in 3D space.
23 | To compute global and regional LV volumes.
24 | To display global LV volume waveform and provide selective display of regional volume waveforms.
25 | Imaging modes
26 | 2D grayscale imaging with advanced pulse coding, pulse shaping and frequency compounding technologies
27 | M-mode
28 | M-mode color Doppler
29 | M-mode Tissue Doppler
30 | Live 3D Echo (instantaneous volume rendering of cardiac anatomy)
31 | Live xPlane imaging (simultaneous display of two live imaging planes)
32 | Tissue Harmonic Imaging (THI)
33 | Color Doppler
34 | High-PRF pulsed wave (PW) Doppler
35 | Duplex and simultaneous 2D/PW Doppler
36 | Duplex continuous wave (CW) Doppler
37 | Duplex color flow and CW Doppler
38 | Duplex 2D, color flow, PW Doppler
39 | Duplex 2D, CPA, PW Doppler
40 | Tissue Doppler Imaging (TDI)
41 | 3D Applications
42 | Applications
43 | Adult Echocardiography
44 | Stress Echocardiography
45 | Vascular (peripheral, cerebrovascular, temporal and orbital TCD, and abdominal vascular)
46 | Transesophageal echocardiography (pediatric)
47 | Transesophageal echocardiography (adult)
48 | Contrast echocardiography (LVO, low MI and high MI detection)
49 | Perioperative
50 | Epicardial echocardiography
51 | Strain imaging capability and automated analysis
52 | Software and features
53 | Adult and pediatrric Cardiology Software and Calculation Package
54 | Vascular Imaging Software and Calculation Package.
55 | Pre-defined TGC curves optimized for consistently excellent imaging with minimal adjustment.
56 | Ability to remove virtually all clutter and artifact; available on linear and curved array transducers for B-mode.
57 | Harmonic imaging: available on all transducers
58 | 16-level digital reconstructed zoom with pan capability
59 | High Definition zoom concentrates all image processing power into a user-defined area of interest; possible to combine High Definition (HD) Zoom with Pan Zoom
60 | Cineloop image review
61 | Multiple technologies for one-button approach to automatically and instantly adjust system performance for different patient sizes, flow states and clinical requirements
62 | Live 3D Echo: 3D LVO setting
63 | Rotation in either absolute or relative trackball motion, Easy use, 3D quantification.
64 | Protocols with smart exam tools: Exam guide with on-screen display, and Automatic Changing.
65 | Stress Echo
66 | Data storage formats include DICOM
67 | 3D DICOM
68 | Offline Cardiac analysis Software
69 | Up to three on-board peripheral devices
70 | DVD recorder.
71 | Small format digital color printer (USB)
72 | Small format digital B/W printer (USB)
73 | USB port
74 | ScanHeads and Probes:
75 | Non-imaging 2MHz PW/CW Doppler transducer for cardiac applications on all systems
76 | Phased array probes: frequency range: 1 to 5 MHz or better: Extended frequency range fully sampled 2D phased array probe with both elevation and lateral dimensional control for real-time 3D volume image acquisition and highest number of scanning
77 | elements
78 | Phased array probes: frequency range: 3 to 7 MHz or close similar with highest number of scanning elements with all systems
79 | Linear array probes: frequency range: 3 to 12 MHz or close similar with highest number of scanning elements with all systems
80 | Pediatric transesophageal phased array probe: specify frequency range and number of elements, one to be supplied with all accessories for full operation .
81 | Adult transesophageal phased array probe: specify frequency range and number of elements, one to be supplied with all accessories for full operation .
82 | 4D capable adult TEE probe : specify frequency range and number of elements, one to be supplied with all accessories for full operation .
83 | Neonatal phased area probe: frequency 12 MHz shall provide high-resolution imaging for pediatric echo and neonatal cephalic applications.
84 | Reporting
85 | User configurable report generation
86 | Measurements and calculations should be automatically inserted to report.
87 | Selected images automatically inserted in the report.
88 | Connectivity
89 | DICOM Store to send images to a Workstation or to a DICOM node.
90 | DICOM Print to send images to a network printer.
91 | DICOM MEDIA enabling recording on DVD of DICOM images.
92 | DICOM Worklist
93 | Compliance with standards & legislation:
94 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
95 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
96 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
97 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
98 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 116
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
116 | MMI90009 | 0 | ULTRASOUND CARDIAC PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The system shall be an advanced Echocardiography Ultrasound System suitable for Echo, Vascular, Superficial Parts, Transcranial and Contrast Imaging.
2 | Architecture
3 | Two Interactive graphical touch screens, 19” Flat monitor, adjustable height.
4 | System architecture should have the capability of multiple data streams simultaneously built for 2D, Panoramic and MPR.
5 | Digital broadband acoustic beamforming technology or multi-frequency technology (please specify).
6 | At least 50,000 Digital channels.
7 | At least 180 dB Full time input dynamic Range.
8 | An advanced algorithm for Adaptive Image Processing for noise and artifact reduction to improve tissue conspicuity.
9 | Automation
10 | Single button optimization key for adaptive gain compensation.
11 | Single button optimization of focal range position.
12 | Single button optimization key that automatically adapts system performance for different patient size and different flow states.
13 | Tissue Harmonic Imaging.
14 | Automatic Doppler trace analysis.
15 | Real-time Compound Imaging
16 | Automated Doppler (maintains scanning angle)
17 | Automatic Spectral Doppler drawing contours and Analysis
18 | Imaging modes
19 | 2D grayscale imaging with advanced pulse coding, pulse shaping and frequency compounding technologies
20 | M-mode
21 | M-mode color Doppler
22 | M-mode Tissue Doppler
23 | Live xPlane imaging (simultaneous display of two live imaging planes)
24 | Tissue Harmonic Imaging (THI)
25 | Color Doppler
26 | Color Power Angio imaging (CPA)
27 | High-PRF pulsed wave (PW) Doppler
28 | Duplex and simultaneous 2D/PW Doppler
29 | Duplex continuous wave (CW) Doppler
30 | Duplex color flow and CW Doppler
31 | Duplex 2D, color flow, PW Doppler
32 | Duplex 2D, CPA, PW Doppler
33 | Tissue Doppler Imaging (TDI)
34 | Strain imaging capability and automated analysis
35 | Applications
36 | Adult Echocardiography
37 | Vascular (peripheral, cerebrovascular, temporal and orbital TCD, and abdominal vascular)
38 | Contrast echocardiography (LVO, low MI and high MI detection)
39 | Perioperative
40 | Epicardial echocardiography
41 | Software and features
42 | Adult and pediatrric Cardiology Software and Calculation Package
43 | Vascular Imaging Software and Calculation Package.
44 | Pre-defined TGC curves optimized for consistently excellent imaging with minimal adjustment.
45 | Ability to remove virtually all clutter and artifact; available on linear and curved array transducers for B-mode.
46 | Harmonic imaging: available on all transducers
47 | 16-level digital reconstructed zoom with pan capability
48 | High Definition zoom concentrates all image processing power into a user-defined area of interest; possible to combine High Definition (HD) Zoom with Pan Zoom
49 | Cineloop image review
50 | Multiple technologies for one-button approach to automatically and instantly adjust system performance for different patient sizes, flow states and clinical requirements
51 | Protocols with smart exam tools: Exam guide with on-screen display, and Automatic Changing.
52 | Data storage formats include DICOM
53 | Offline Cardiac analysis Software
54 | Up to three on-board peripheral devices
55 | DVD recorder.
56 | Small format digital color printer (USB)
57 | Small format digital B/W printer (USB)
58 | USB port
59 | ScanHeads and Probes:
60 | Non-imaging 2MHz PW/CW Doppler transducer for cardiac applications on all systems
61 | Phased array probes: frequency range: 1 to 5 MHz or better
62 | Phased array probes: frequency range: 3 to 7 MHz or close similar with highest number of scanning elements with all systems
63 | Linear array probes: frequency range: 3 to 12 MHz or close similar with highest number of scanning elements with all systems
64 | Pediatric transesophageal phased array probe: specify frequency range and number of elements, one to be supplied with all accessories for full operation .
65 | Adult transesophageal phased array probe: specify frequency range and number of elements, one to be supplied with all accessories for full operation (f
66 | Reporting
67 | User configurable report generation
68 | Measurements and calculations should be automatically inserted to report.
69 | Selected images automatically inserted in the report.
70 | Connectivity
71 | DICOM Store to send images to a Workstation or to a DICOM node.
72 | DICOM Print to send images to a network printer.
73 | DICOM MEDIA enabling recording on DVD of DICOM images.
74 | DICOM Worklist
75 | Biopsy guide line and zone
76 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
77 | h as IM
78 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
79 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
80 | Compliance with standards & legislation:
81 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
82 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
83 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
84 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
85 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 117
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
117 | MMI90010 | 0 | ULTRASOUND GENERAL PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Advanced, high performance color imaging system for abdominal, vascular, obstetrics, gynecology, neonatal, and small parts applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing, Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R/W Drive and USB
22 | The system shall be suitable for the following applications:
23 | Abdominal
24 | Obstetrical
25 | Gynecological
26 | Musculoskeletal
27 | Vascular
28 | Urological
29 | Small Parts and Superficial
30 | Pediatric and Neonatal
31 | Transcranial Doppler
32 | The system shall possess capabilities for the following scanning methods:
33 | Electronic Sector
34 | Electronic Convex
35 | Electronic Linear
36 | The system shall be capable of accommodating the following transducer types:
37 | Phased array sector
38 | Convex
39 | Micro-convex
40 | Linear
41 | Operating modes shall include but not be limited to:
42 | B-Mode
43 | M-Mode
44 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
45 | Color Flow Mode
46 | Power Doppler Imaging
47 | PW Doppler
48 | M-Color Flow Mode
49 | Steerable CW Doppler
50 | Trapezoidal Imaging
51 | Optional pan view
52 | shear-wave elastography, and contrast-enhanced imaging , Specify.
53 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
54 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
55 | Tilt/rotate adjustment
56 | Digital Brightness/Contrast Adjustment
57 | Integrated Task Light
58 | Slave monitor :
59 | A second slave monitor Wi-Fi connected with the scanner shall also be included in the offer, for patient viewing. All necessary mounting accessories shall be supplied.
60 | System must be offered with an above 19-inch high resolution,
61 | flat panel, medical grade
62 | monitor with wide viewing angles & good colour resolution.
63 | Resolution : 1024x768pixels or better
64 | The user interface panel shall possess ergonomic and friendly functional design, including:
65 | Alphanumeric keyboard for data entry with ergonomic key operations
66 | Interactive backlighting
67 | Indicator lights to identify activated keys
68 | The system cart shall possess features of high stability and maneuverability, including:
69 | On-board Storage for color-printer
70 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
71 | Multiple probe holders
72 | Gel dispenser holder
73 | Rear handle for transport maneuverability
74 | Video connection shall be integrated for film printing by laser camera.
75 | A color small format (A4), high resolution printer shall be included with the offer
76 | The system shall possess extensive display annotations including:
77 | Institution name
78 | Date / Time
79 | Patient Demographics
80 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
81 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
82 | Focal zone markers
83 | TGC Curve (LGC or Lateral Gain Compensation)
84 | Body patterns
85 | System messages display
86 | Trackball functionality status
87 | Biopsy guide line and zone
88 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
89 | h as IM
90 | The following probes shall be included (all probes shall be multi-frequency, wide band technology):
91 | Abdominal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
92 | Endo-vaginal convex probe: specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint. (Biopsy guide shall be quoted)
93 | Linear Probe for vascular, small parts, breast, neonatal and pediatrics: specify the frequency range, the number of elements, the convex radius and the FOV
94 | Compliance with standards & legislation:
95 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
96 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
97 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
98 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
99 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 118
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
118 | MMI90014 | 0 | ULTRASOUND VASCULAR PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Advanced, high performance color imaging system for vascular applications.
2 | The system shall possess extensive computational power, image-manipulation capability and workflow flexibility.
3 | The system shall employ digital beam-forming technology and multi-frequency technology.
4 | The system shall be compatible with one-button image optimization and adaptive image processing for noise and artifact reduction to improve tissue conspicuity.
5 | The system shall be supplied with a gel warmer, to be mounted on each unit.
6 | The following minimal system performance characteristics shall be incorporated:
7 | Multi-Frequency/Digital Broadband Beam former Technology
8 | Displayed Imaging Depth shall be provided
9 | Minimum Depth of Field shall be provided
10 | Maximum Depth of Field shall be provided
11 | dynamic auto focusing ,Focal Zone, Transmission Focus Points
12 | User adjustable B Mode gain, TGC, and dynamic range controls
13 | 256 Shades of Gray
14 | Up to 232 dB Dynamic Range in 3dB steps
15 | Adjustable Field of View (FOV)
16 | Image Reversal
17 | Image Rotation
18 | Three Active Probe Ports
19 | Integrated HDD (≥ 500 GB. Specify capacity)
20 | CINE memory capacity ≥ 256 MB for up to 1500 frames or close similar (variable speed and sequence selection. Provide details)
21 | Integrated DVD-R Drive and USB
22 | The system shall be suitable for the various vascular applications, cerebrovascular and peripheral vascular. Add abdominal vascular if the scanner will be used for evaluation of renal arteries, portal venous system, etc.
23 | The system shall possess capabilities for the following scanning methods:
24 | Electronic Sector
25 | Electronic Convex
26 | Electronic Linear
27 | The system shall be capable of accommodating the following transducer types:
28 | Phased array sector
29 | Convex
30 | Micro-convex
31 | Linear
32 | Operating modes shall include but not be limited to:
33 | B-Mode
34 | M-Mode
35 | Color M-Mode and Optional M-mode perpendicular to the anatomy independent of transducer orientation)
36 | Color Flow Mode
37 | Power Doppler Imaging
38 | PW Doppler
39 | M-Color Flow Mode
40 | Steerable CW Doppler
41 | Trapezoidal Imaging
42 | Optional pan view
43 | Display modes and capabilities shall include simultaneous display of duplex or triplex as well as alternating and multi-image split screen modes. Provide list of modes and possible combinations.
44 | The system shall incorporate a 17” non-interlaced, high-resolution flat display color monitor with:
45 | Tilt/rotate adjustment
46 | Digital Brightness/Contrast Adjustment
47 | Integrated Task Light
48 | Slave monitor :
49 | A second slave monitor Wi-Fi connected with the scanner shall also be included in the offer, for patient viewing. All necessary mounting accessories shall be supplied.
50 | System must be offered with an above 19-inch high resolution,
51 | flat panel, medical grade
52 | monitor with wide viewing angles & good colour resolution.
53 | Resolution : 1024x768pixels or better
54 | The user interface panel shall possess ergonomic and friendly functional design, including:
55 | Alphanumeric keyboard for data entry with ergonomic key operations
56 | Interactive backlighting
57 | Indicator lights to identify activated keys
58 | The system cart shall possess features of high stability and maneuverability, including:
59 | On-board Storage for color-printer
60 | Swiveling castors (≥ 5” diameter), heavy duty, capable of withstanding shocks from elevator and similar uneven floor transportation within the hospital.
61 | Multiple probe holders
62 | Gel dispenser holder
63 | Rear handle for transport maneuverability
64 | Video connection shall be integrated for film printing by laser camera.
65 | A color small format (A4), high resolution printer shall be included with the offer
66 | The system shall possess extensive display annotations including:
67 | Institution name
68 | Date / Time
69 | Patient Demographics
70 | System parameters (frequency, power, gray/color Bar, cine gauge, probe type, application name, etc.)
71 | Imaging parameters by mode (gain, edge enhancement, frame averaging, image depth, dynamic range, etc.)
72 | Focal zone markers
73 | TGC Curve (LGC or Lateral Gain Compensation)
74 | Body patterns
75 | System messages display
76 | Trackball functionality status
77 | Biopsy guide line and zone
78 | Extensive digital image pre-processing and post-processing (calculation packages) capabilities shall be incorporated. Provide a list for all modes with associated parameters and ranges. ( offline optional analysis software for advanced analysis suc
79 | h as IM
80 | A linear probe for vascular applications shall be included in the offer. (Specify the frequency range, the number of elements, the convex radius and the FOV )
81 | Abdominal convex probe for deeper vascular applications (renal arteries) and obese patient : specify the frequency range, the number of elements, the convex radius, the FOV and the physical footprint.
82 | Compliance with standards & legislation:
83 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
84 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
85 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
86 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
87 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 119
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
119 | MOP20002 | 0 | EEG UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | EEG MONITORING SYSTEM | 32 Channel, transport
2 | EEG channels input | 25 Monopolar 14 Bipolar ,8DC
3 | CMRR , dB | 110 dB @ 60 Hz
4 | Sensitivity control range , V / mm | Software controlled Infinte steps
5 | Noise , V | 1.5 Micro VP - P
6 | Sampling rate , Hz | 200 - 10000 per channels
7 | Electrode - Impedance check | Yes
8 | SYSTEM CONTENT
9 | Junction box | Yes
10 | Stand with Support arm assembly
11 | Isolated Power supply for system | Yes , Optional back up battery, Itemies prices
12 | Reusable EEG cup electrode | Yes
13 | Availability to use head cap | Yes
14 | Carrying Case, either soft or hard | Yes
15 | Carrying Case, either soft or hard | Yes, specify
16 | PHOTIC ( Flash ) stimulator With Variable Frequency | Optional, Itemized Price
17 | COMPUTER SYSTEM | Laptop / notebook ( medical grade )
18 | INTEL CORE 2 DUO 2.2 GHZ | Yes or latest
19 | 2 GB RAM OR MOER | Yes or latest
20 | VGA Card Resolution 256 - 512 MB | Yes or latest
21 | 320 GB HARD DISK | Yes or latest
22 | 10/ 100 MBit TP Network card (Ethernet) | Yes or latest
23 | CD-DVD-RW drive 8/4/32 x | Yes or latest
24 | Flat screen with high resolution LCD or TFT, Size 14?- 17? HD | Yes, Specify
25 | MS Windows 7 PROFETIONAL | Yes or latest
26 | MS office 2007 | Yes or latest
27 | STANDARD SOFTWARE
28 | Unlimited number of montages | Yes or latest
29 | Availability to divide the display for screen A&B for comparing signals | Yes
30 | Software for automatic color report generator | Yes
31 | Software for measurement, zooming, review & analysis | Yes
32 | Color amplitude map, brain map 2D | Yes
33 | Color power spectrum map Alpha, Beta, | Yes
34 | Theta, Delta | Yes
35 | ACCESSORIES TO INCLUDE
36 | Electrode Ag/Agl or Gold | Standard set
37 | Patient cable assembly | Yes
38 | Adhesive electrode paste | Yes
39 | Black / white printer compatible with above system | Yes, laser or ink jet printer
40 | POWER & FREQUENCY | 220 Volt - 60 Hz.
41 | OPTIONAL ITEMIZED PRICES | Yes
42 | Patient chair with height adjustable, electrical | Yes
43 | OTHER SPECIFICATION | Up gradbale to SLEEP LAB ,Optinal itemized price.
Company Stamp |  |  |  | Signature

## ITEM 120
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
120 | MOR10007 | 0 | CABINET WARMER BLANKET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | WARMER | To be used for Blanket warming
2 | Microprocessor Controlled | Yes
3 | Number of Doors | Two
4 | Door Configuration | Glass
5 | Number of Shelves | 2 Each Compartment
6 | Temp. Range | 30 to 40 Degree Celsius
7 | Outer Material | Stainless Steel
8 | Castors | Four two of them lockable
9 | Capacity | at least 300 Litres each Compartment
Company Stamp |  |  |  | Signature

## ITEM 121
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
121 | MOR10029 | 0 | FIXATOR SKULL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Stainless Steel | Yes
2 | Steam Auto cleavable with storage | Yes
3 | Complete skull fixator | Yes
4 | Skull pins For Adult | Yes
5 | Skull pins for Children | Yes
6 | Cross Bar Attachment | Yes
7 | Compatible with operating table | Specify the table
8 | Adapter for Head Position | Yes
9 | FDA or CE Approved | Yes
Company Stamp |  |  |  | Signature

## ITEM 122
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
122 | MOR10087 | 0 | SUCTION UNIT HIGH VACUUM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Stable mobile stand made of lightweight, heavy-duty, antirust metal with 4 antistatic swiveling castors and wide base to prevent accidental tripping.
2 | The pump shall possess (or exceed) the following technical specifications:
3 | Double piston pump continuous rating type.
4 | Pressure guage up to 800 mm Hg
5 | Pressure regulator .
6 | All Castors with brakes .
7 | Over flow protection.
8 | Silicon Tubings.
9 | Suction Bottle Capacity - 2 x 2500 ml minimum (with safety valve)
10 | Bottles graduation scale 100ml - material polysulphone . sterelisable at 134°C .
11 | Vacuum control : ≥ 50 LPM specify
12 | Variable (regulating valve) vacuum pressure from 0 to approximately – 750 mm Hg, with negative pressure gauge indicating actual pressure.
13 | Low noise operation (≤ 45 dB @ 1 meter)
14 | Incorporated filter (specify type)
15 | Foot switch operation .
16 | Hand control .
17 | The offer shall include all the accessories and parts necessary for the full and efficient operation of the system. A detailed list of such standard accessories (with part numbers and quantities) shall be included.
Company Stamp |  |  |  | Signature

## ITEM 123
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
123 | MOR10093 | 0 | TABLE MAYO SMALL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MAKE | Stainless Steel AISI 304
2 | HEIGHT . | Adjustable with foot operated hydraulic pump from 70-110 cm.
3 | APPROX. TOP DIMENSIONS | 60 x 40cm
4 | TOP | Rotable
5 | LOAD CAPACITY | 8 Kg
6 | CASTORS | Yes , with 2 lockable .
7 | CASTOR DIAMETER . | More than 50mm .
8 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 124
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
124 | MOR10095 | 0 | TABLE OPERATING ENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Electric or Electro-hydraulic, specify
2 | NO.OF SECTIONS | 5 or more
3 | PADS | integrated
4 | HEIGHT ADJUSTMENT | 600 - 1200mm, approx.
5 | TRENDELENBURG | +30 degree approx
6 | REVERSE TRENDELENBURG | -30 degree approx
7 | LATERAL TILT | 20 degree approx
8 | LONGITUDNAL DISPLACEMENT | 25 cms
9 | HEAD SECTION UP & DOWN | 45 degree approx
10 | BACK SECTION UP & DOWN | +50 to -30 degree approx.
11 | FOOT / LEG SECTION UP & DOWN | +45 to -90 degree approx.
12 | LOAD LIMIT | 250 kgs approx
13 | HAND SWITCH CONTROL | Yes, soft touch
14 | RADIOLUCENT TABLETOP | Yes, all sections
15 | C-ARM ACCESSIBLE | Yes, Full length
16 | X-RAY CASSETTE TUNNEL FROM EITHER END | Yes
17 | KIDNEY ELEVATOR | Yes
18 | BASE TYPE | Mobile
19 | COLUMN HOUSING | Stainless steel
20 | NON-CONDUCTIVE CASTORS | Yes
21 | LOCKABLE CASTORS | Yes
22 | BACK UP BATTERY | Yes
23 | BATTERY CHARGER | Yes, Built in
24 | POWER SUPPLY | 220V, 60 Hz
25 | ACCESSORIES | a)Arm support b)Lateral support c)Buttock support d)Pelvis or chest support e)Wristlet f)Anaesthesia screen g)Circular socket clamp h)Pair shoulder rest i)Knee support
26 | Accessories :- | As Follow
27 | Individual modular setup | Yes
28 | Pads and carrier plates | Removed without tools
29 | Viscoelastic foam for decubitus prevention and reduction of shearing forces | Yes
30 | hygienity | High
31 | Prevention of X-ray absorption | Yes
32 | Fixation | Without fleece and Velcro straps
33 | Longitudinal displacement | Motor powered
34 | Back plate drive | Motor powered
35 | Leg plate drive | Motor powered
36 | Zero positioning function | Automatic
37 | Collision protection | Yes
38 | Transfer of operating table top | Two-side
39 | Recognition of Orientation direction | Automatic
40 | Automatic locking | Yes
41 | Connection for external operating table top control | Yes
42 | Technical specification | Specify
43 | Table width without accessory | Specify
44 | Radiotranslucent window between the bars Min /max height (top edge/stationary column) | Specify
45 | Trendelenburg/ low head | Specify
46 | Trendelenburg/ low feet | Specify
47 | Lateral tilt left | Specify
48 | Lateral tiltright | Specify
49 | Leg sections down | Specify
50 | Leg sections up | Specify
51 | Longitudinal displacement | Specify
52 | Self-supporting, radiotranslucent length in direction of the head | Specify
53 | Self-supporting, radiotranslucent length in direction of the feet | Specify
54 | Joint module | Motor operated
55 | Motor-powered adjustments | Specify
56 | Dimensions(L x W x H) in mm | Specify
57 | Four-part leg section | Specify
58 | Leg section suited for kneeled positioning of the patient | folded for 95 degree / raised for 90 degree
59 | With PUR structural foam pads and electrical discharging capacity intended length | approx. 940
60 | Dual-joint head section | Yes
61 | Tilting angle with the head section pad not being tilt up: | Specify
62 | Tilting-up angle of head section pad: | Specify
63 | Dimensions: L x W x H (mm) | Specify
64 | The operating table column | Yes
65 | transferral of the operating table tops from both sides | Yes
66 | IR hand control | Yes
67 | Corded hand control | Yes
68 | Foot switch with three functions | Yes
69 | Operating control panel being integrated into the operating table column | Yes
70 | Connection values for mobile transformer unit | 220 V AC / 60 Hz
71 | Height including universal operating table top | Specify
72 | Transporter
73 | Trolley and frame | Made of corrosion-resistant material.
74 | Length of trolley | Specify
75 | Overall width including rollers | Specify
76 | Height | Specify
77 | Trendelenburg tilt head/foot low | Specify
78 | Electromotive adjustment | interchangeable storage battery
79 | Lifting of height adjustment | 170 mm
80 | Batteries | high performance Lithium-Ion battery
81 | Long back section for general surgery | Yes
82 | Shortened back section for general surgery Pair of leg sections, with four-fold sectioning, articulated and hinged | Yes
83 | All access. Needed for ENT surgery | Yes, itemiz price
Company Stamp |  |  |  | Signature

## ITEM 125
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
125 | MOR10096 | 0 | TABLE OPERATING GENERAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Electric or Electro-hydraulic, specify
2 | NO.OF SECTIONS | 5 or more
3 | PADS | integrated
4 | HEIGHT ADJUSTMENT | 600 - 1200mm, approx.
5 | TRENDELENBURG | +30 degree approx
6 | REVERSE TRENDELENBURG | -30 degree approx
7 | LATERAL TILT | 20 degree approx
8 | LONGITUDNAL DISPLACEMENT | 25 cms
9 | HEAD SECTION UP & DOWN | 45degree approx
10 | BACK SECTION UP & DOWN | +50 to -30 degree approx.
11 | FOOT / LEG SECTION UP & DOWN | +45 to -90 degree approx.
12 | LOAD LIMIT | 250 kgs approx
13 | HAND SWITCH CONTROL . | Yes, soft touch
14 | RADIOLUCENT TABLETOP . | Yes, all sections
15 | C-ARM X-RAY ACCESSIBLE . | Yes, Full length
16 | X-RAY CASSETTE TUNNEL FROM EITHER END . | Yes
17 | KIDNEY ELEVATOR . | Yes
18 | BASE TYPE | Mobile
19 | COLUMN HOUSING | Stainless steel
20 | NON-CONDUCTIVE CASTORS | Yes
21 | LOCKABLE CASTORS | Yes
22 | BACK UP BATTERY . | Yes
23 | BATTERY CHARGER . | Yes, Built in
24 | POWER SUPPLY . | 220V, 60 Hz
25 | ACCESSORIES . | a)Arm support b)Lateral support c)Buttock support d)Pelvis or chest support e)Wristlet f)Anaesthesia screen holder . g)Circular socket clamp h)Pair shoulder rest i)Knee support for ACL & BCL cases . J) Scope Holder .
26 | OTHER SPECIFICATION .
Company Stamp |  |  |  | Signature

## ITEM 126
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
126 | MOR10114 | 0 | TELESCOPE 90 DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scopes Excellent images, significant depth of focus | Yes
2 | Aspherical Lens | Yes
3 | Integrated fiber optic light illumination | Yes
4 | 4X magnification focusing device | Specify
5 | Excellent Depth of Field
6 | Diameter: 10 mm approx. | Yes
7 | Length: 200 mm approx. | Yes
8 | Direction of view: 90° | Yes
9 | Completely made from stainless steel | Yes
10 | Case for cleaning and sterilization | Yes
Company Stamp |  |  |  | Signature

## ITEM 127
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
127 | MOR10122 | 0 | ULTRASOUND VASCULAR DOPPLER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Regulation Standard :
2 | FDA 510 (K) Clearance + CE Mark & SFDA Registration | Yes
3 | Validated Up to Date Certificates Must be Submitted with the Offer | Yes, Failure of Submission and Comply Will Disqualify the Offer
4 | Product Description : | Laptop Ultrasound
5 | Console Design :
6 | Laptop / Hand Carried on height adjusted mobility Cart | Cart Should Include the Following :Integrated Power SupplyProbe Portswheels with locking braking system
7 | Total Weight - (incl. battery) | ≤ ( 7 ) kg .
8 | Esay Acess mode keys | Yes, (Specify)
9 | User Friendly | Yes
10 | Easy to Clean-Sterilize | Yes
11 | Screen Monitor | ≥ (15) inch. LCD with High Resolution
12 | Security Lock | Yes
13 | Built-In Battery | Yes
14 | Battery scanning time | 1 hour
15 | Power Requirement | 220-240 V / 60 Hz
16 | Computing Capabilities :
17 | Operation System | Windows or equivalent
18 | Integrated Hard Desk | At least 80 GB
19 | Fast System boot up | Yes
20 | DVD /CD burner | Yes
21 | Clinical Applications :
22 | Vascular & Cardiac | Yes, Analysis Packages Standard
23 | Abdominal, Small parts | Yes
24 | Vascular Access & Superficial | Yes
25 | Musculoskeletal | Yes
26 | Other Available Applications | Yes
27 | Multi Frequency Probes : | One Each , With dedicated Operational Software
28 | Linear array | 4 - 10 MHz (±2)
29 | Convex array | 1 - 5 MHz (±2)
30 | Phased array | 3 - 8 MHz (±1)
31 | Active Probe Ports | 1 in the main console3 ports with the cart
32 | Operating Modes :
33 | Brightness Mode (B-Mode) (2D) | Yes
34 | Motion Mode (M-Mode) | Yes
35 | Color Doppler mode | yes
36 | Pulsed wave Doppler (PW) with HPRF | Yes
37 | Color Flow Doppler Mode (CFM) | Yes
38 | Power Doppler Mode | Yes
39 | Continuous Wave Doppler | Yes
40 | Dopex Mode | Yes
41 | Triplex Mode | Yes
42 | Scanning Parameters :
43 | Imaging Depth, cm | Not less than 30 cm
44 | Dynamic Range | Up to 150 dB
45 | TGC Control | Yes
46 | Frame Rate 2D | Up to 350 fps
47 | Frame rate Color Doppler | Up to 100 fps
48 | GrayScale Levels | 256 ( 8 bits )
49 | Scanning Automatic Optimization | Capability for Auto optimization in 2D, Doppler & Color Dopler
50 | Digital Calipers | Yes
51 | Preprocessing | Yes
52 | Postprocessing | Yes, (Measurment, adjust the Daynamic Range, TGC, color map & Gain) on stored images
53 | Zoom , freeze , gain , save , print , cine review , text / picto , depth | Yes
54 | Doppler :
55 | Type | CW, PW, in all probes
56 | Frequency Display | Yes
57 | Doppler baseline shift after freeze | Yes
58 | Velocity Display | Yes
59 | User – defined annotation | Yes
60 | GrayScale Levels | Yes, Up to 20 Sets
61 | Transmit focal zones | Select between 1 to 8
62 | Digital reconstructed zoom | 16 levels
63 | Gain Control after freeze | Yes
64 | Adaptive Gain Compensation | Yes
65 | Directional color power Doppler | Yes
66 | Cine Features :
67 | Cine Storage | Yes, (Still images / Clips / Volume)
68 | Length / Capacity | 140 MB (up to 3000 images)
69 | Scanning Features :
70 | Intima Wall Thickness Measurments | Yes
71 | Speckle Reduction Imaging | Yes, Organ Specific
72 | Trapeziodal Imaging | Yes
73 | Compound Imaging | Yes
74 | Tissue Harmonic Imaging | Yes, Standard
75 | TGC (Time Gain Compensation) | Yes
76 | Color flow mapping | Standard
77 | Automatic Optimization | Yes, for B-Mode, Color Doppler and PW Doppler
78 | Stearable Color Doppler with linear probe | Yes
79 | Color Doppler Frequency Change indepent from 2D | Yes
80 | Color Compare | Yes
81 | Removal of Color map from the tissue during digital reply | Yes
82 | Adjustable Transmit Focus | Yes
83 | Dynamic Receive Focus | Yes
84 | Selectable Dynamic Range | Yes
85 | Automatic Frequency and focal zone optimization with depth adjustment | Yes
86 | Zoom / Pan :
87 | Real- time image | Yes
88 | Frozen image | Yes
89 | Split Screen | Yes
90 | Connectivity :
91 | Integration and Network | Yes, LAN Speed not less than 1G bit and Wifi if applicaple
92 | Wireless Connection | Yes
93 | Full DICOM Connectivity | (Verify, Print, Store, Modality work list, multi frame, storage commitment, Modality Performed Procedure Step, MPPS, Media Exchange). Query / Retrieve, Structured Reporting
94 | Performing Measurements and Calculations in Current as well as already Stored Images and Clips. | yes
95 | Storage and Playback of M-mode, CW and PW Spectral Doppler. | yes
96 | Full connection to PACS system | Yes
97 | Full connection to RIS | Yes
98 | Full Connectivity to USB Printers | Yes
99 | Printing of Patient Reports Via a USB Connection. | Yes
100 | Recalculation for stored images | Yes
101 | Possibility to Export on different Formats | Raw data, PC formats, DICOM Viewer
102 | 3D clips in DICOM Format | Yes
103 | Security / Privacy :
104 | HIPAA Patient Data Security | Yes, Vendor should Submit Official Letter of Compliance - (Failure of Submission and Comply Will Disqualify the Offer)
105 | Data Protection & Against Deletion & Export | Yes , Each User Should have User Name Management and Password for Data Access Protection, and Vendor should Submit Official Letter of Compliance - (Failure of Submission and Comply Will Disqualify the Offer)
106 | Ultrsound System security Virus protection Or equivilant | Yes , Vendor should Submit Official Letter of Compliance - (Failure of Submission and Comply Will Disqualify the Offer)
107 | Accessories :
108 | Secondary Battery | Yes
109 | Carrying Case | Yes
110 | MOH Requirement :
111 | Demo Unit for at least 1 Weeks (in the presence of qualified application specialist) | Yes, Demo Unit has to be ready in (KSA at bider sotre) at the day of submission of the offer to be delivered for evaluation by the committee from radiology Department & General Supplies Department within (15) days and Sites will be defined by MOH
112 | Clinical & Application Training Course for Doctors | Yes, on site initial application training for (1) Week within (3) Months in total followed by another (1) Week within (6) Months for another advanced applications (Please Provide Detailed Training Plan) .
113 | Warranty | 5 years full warranty Unlimited including Battery & Probes
114 | A-All the vendors should supply all available options of their machine not stated in the above specifications as an optional items with separate price. B- All complince letters must be submitted on separate file as part of the offer (please | Yes
115 | indicate loaction on your answer )
116 | sample images in DICOM format acquired from similar machine for different clinical cases with attached radiation dose for each study should be supplied on CD or DVD as part of the tender document | Yes
117 | Maintenance and Service Contract Value with Spare Parts for the First Year after Warranty | Specify Value
118 | Maintenance and Service Contract Value with Spare Parts for the Second Year after Warrant | Specify Value
119 | Maintenance and Service Contract Value with Spare Parts for the Third Year after Warranty | Specify Value
120 | Maintenance and Service Contract Value with Spare Parts for the Fourth Year after Warranty | Specify Value
121 | Maintenance and Service Contract Value with Spare Parts for the Fifth Year after Warranty | Specify Value
122 | Vendors Should be the Provider of Both Sales & Service | Yes, Submit a file contain the name and CVs of sale, application specialist and biomedical engineers
Company Stamp |  |  |  | Signature

## ITEM 128
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
128 | MOR20006 | 0 | ANESTHESIA MACHINE MRI COMPATIBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | ANAESTHESIA UNIT MRI
2 | APPLICATION | Adult and Pediatric , NEONATAL
3 | MRI Compatible up to 3 Tesla | Yes close to core 10cm or less
4 | PIPELINE GAS INLETS | 3 (O2, N20, Air)
5 | Pressure Indicator for all above gasses. | Gauge Indicator
6 | FLOWMETER FOR O2 , N2O & AIR | Yes, Mechanical control, tube flow display with cascade and low flow.
7 | Oxygen range | lit/min, specify
8 | Nitrous oxide range | lit/min, specify
9 | Compressed Air range | Lit/min, specify
10 | EMERGENCY O2 | Yes, closed
11 | PRESSURE RELIEVE VALVE | Yes
12 | ANAESTHETIC GAS SCAVENGING SYSTEM | Should be either passive or active according to the hospital gas system
13 | CO2 ABSORBER COMPLETE, INCLUSIVE VENTILATOR & ABSORBER BYPASS | Yes,
14 | DRAWER FOR UNIT WIDE & DEEP | Yes
15 | SUCTION SYSTEM | Yes
16 | O2 FAIL-SAFE | Yes
17 | HYPOXIC MIXTURE FAIL-SAFE | Yes
18 | VAPORIZERS | Yes, Not Less Than Three
19 | Agents | Sevoflurane, Isoflurane ,.key filler or quick filler
20 | Type | Variable by pass/direct injection
21 | Number | Not less than 3
22 | Interlock | Yes
23 | AUTOMATIC VENTILATOR | Yes
24 | Operated by compressed air and in case of emergency, capable to operate by compressed oxygen. | Yes
25 | Screen Display and Size | Yes, LCD, TFT & 17 Screen FOR SLAVE MONITOR
26 | Primary controls / Ventilation modes | Volume control specify mode of ventilation / Pressure Control specify mode of ventilation.
27 | Rate alarm | Yes
28 | Apnea alarm | Yes
29 | High/low min vol | Yes
30 | High/low flow | Specify
31 | O2 concentration | Specify , ( 21 % - 100 % )
32 | Response time, Vol % | Yes
33 | POWER REQUIRED, VAC | 220 Volt , 60 Hz
34 | Auxillary outlets | Yes, specify
35 | BACK-UP BATTERY | Yes
36 | Type & Time | Specify
37 | ACCESSORIES | Yes,
38 | Patient Circuit system | Yes,. Patient circuit for adult 3 sets, & pediatric circuit 3 sets ,
39 | Anesthesia mask | Yes, 3 sets for each size adult & pediatrics
40 | Magill tube | Yes
41 | Jackson Rees or equivalent for open patient circuit. | Yes
42 | OPTIONAL ACCESSORIES | Itemized price
43 | Aluminium E cylinder oxygen | Yes, Qty. 4
44 | Aluminium E cylinder medical air | Yes, Qty. 4
45 | Aluminium E cylinder Nitrous oxide (N20) | Yes, Qty. 4
46 | OTHER SPECIFICATION | .Approved by FDA, CE, ISO
47 | ANEASTHESIA PATIENT MONITOR , | Yes , Preconfigured 2 MRI Compatible with bis or equivalent ONE SLAVE MONITOR
48 | Application | Adult, Pediatric ,neonat
49 | Type | SPECIFY
50 | MRI Compatible up to 3 Tesla | Yes AT 10cm FROM THE CORE
51 | Parameter Monitor
52 | ECG 3-5 leads | Yes
53 | Heart Rate | 30 to 300 beats per minute
54 | NIBP | Yes
55 | SPO2 | Yes
56 | 2 X IBP | Yes, specify
57 | N2O | Yes
58 | FiO2 | Yes
59 | Respiration | Yes
60 | Multi gases anesthetic 5 agents, Isoflurane, Halothane, Sevoflurane, | Yes
61 | Screen | 2 monitor
62 | Size, inch | SPECIFY
63 | Type, Color, LCD, TFT | Specify
64 | Channel | 8 Channels
65 | Optional Parameters
66 | Temperature | Specify principle of measurement
67 | Data Recorder | Yes
68 | Trending (Memory)
69 | Parameters | Specify
70 | Graphical/Tabular | Specify
71 | Length of time, hr. | Specify
72 | Data Transfer To Control Room
73 | Remote Control | Yes
74 | External Screen or PC | Specify
75 | Accessories
76 | ECG | Yes
77 | SPO2 | Fiber optic SPO2 sensors for adult pediatric & neonatal .
78 | NIBP | Cuff and hose for adult, pediatric & neonatal .
79 | IBP | Yes, complete
80 | Gases measurement accessories | Yes, for six months consumption.
81 | Power Requirements | 220 V, 60 Hz.
82 | Battery built-in operation time | Not less than 60 min.
83 | OTHER SPECIFICTIONS | Approved by FDA, CE, ISO - DICOM.3-All access. Needed- slave screen out of the room
Company Stamp |  |  |  | Signature

## ITEM 129
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
129 | MOR20009 | 0 | BLANKET WARMING SYSTEM AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Inflation time | 90 sec
2 | Warming up time | 4 minutes
3 | Length of Air Hoses | 1500 mm
4 | Temperature | 20°C - 42°C, step of 1°C
5 | Pressure | 0 – 60 mb. Step of 5 mb
6 | Alarm | Audio and visual
7 | Hand Controller | Specify
8 | Warm-Air Mattress, Wide | Length 200 cm, Width 60 cm.
9 | Warm-Air Mattress, Small | Length 100 cm, Width 50 cm
10 | Strap | Yes
11 | Hose seal | 2 pcs
12 | Hose set complete | 2 pcs
13 | Itemized price for all accessories | Yes
Company Stamp |  |  |  | Signature

## ITEM 130
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
130 | MOR20010 | 0 | CART ANESTHESIA SUPPLY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | ANAESTHESIA, SUPPLY CART
2 | Mobile unit, impact resistant plastic top with raised edges. | Drawers unit with large and small drawers
3 | Swing out side storage unit with tray divider. | Sharps container holder, and tilt out bin.
4 | Multi-purpose overbridge unit . | With universal clamp, blade /bag holder, small. large metal bin, label/tape dispenser, & bins w/ cover.
5 | COMPLETE WITH THE FOLLOWING:
6 | Forceps, Magil, Adult. | set qty 2 set
7 | Forceps, Magil, Child. | set qty 2 set
8 | All needed accessories. | Specify
9 | 4 non-conductive castors | at least 5 inch
10 | other specifications | FDA, CE OR OTHER SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 131
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
131 | MOR20011 | 0 | COMPRESSION SYSTEM SEQUENTIAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Microprocessor controlled operation and safeties
2 | Compression pressure shall be achieved by means of incorporated pump (no external connection to compresses air line shall be necessary)
3 | Treats one or two limbs at a time.specify.
4 | Sequential Cycle: specify
5 | Upper extremities:Providing Different sizes cuffs specify (axilla, Bicep, Elbow, Forearm, Wrist and axilla to finger tip length ..etc)
6 | Lower extremities:Providing Different sizes cuffs specify (Thigh below crotch, Mid thigh, Mid knee, Mid Calf, Ankle ..etc)
7 | Pressure and control parameters should be specified:
8 | Pressure regulation method
9 | Fully variable range from 20 to 80 mmhg Specify.
10 | Inflation and deflation cycle times specify
11 | The unit shall digitally display important operating parameters and other alert messages and alarms. Specify display characteristics and list all displayed information.
12 | Safety features should include: Specify.
13 | Specify high pressure safety cutoff points (preferably more than one level, fully independent) and breaking method
14 | Splash proof .
15 | Internal self test
16 | Compression halt on alarm
17 | Maintenance-free operation
18 | The unit should be portable, self contained incorporating mounting bracket for IV pole and bed rails
19 | The unit shall be able to connect to all types of garments available on the market (open-system) using universal tubing connectors (luer lock or snap-on)
20 | Compliance with standards & legislation:
21 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
22 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
23 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
24 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
25 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 132
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
132 | MOR20015 | 0 | INTUBATION SET COMPLETE CASE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FORCEPS, MAGILL, 24cm, adult (Qty: 1)
2 | FORCEPS, MAGILL, 19cm, child small (Qty: 1)
3 | FORCEPS, MAGILL, 16cm, child extra small (Qty: 1)
4 | (laryngoscope fiber) BLADE, Miller n°0, new born, reusable. (Qty: 1)
5 | (laryngoscope fiber) BLADE, Miller n°1, baby, reusable. (Qty: 1)
6 | (tube, endotracheal) GUIDE, size 3, L 50cm, tubes from 6.0 (Qty: 10)
7 | (tube, endotracheal) GUIDE, size 2, L 45cm, tubes 4.5-6.0 (Qty: 10)
8 | (tube, endotracheal) GUIDE, size 1, L 45cm, tubes 3.5-4.5 (Qty: 10)
9 | ESCHMANN AIRWAY EXCHANGE GUIDE, BUGIE, 15FR, 70cm, single use (Qty: 10)
Company Stamp |  |  |  | Signature

## ITEM 133
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
133 | MOR20017 | 0 | INTUBATION DIFFICULT VIDEO |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Video Laryngoscope For difficult airway intubation
2 | Design | light weight & portable
3 | The system shall be based on the latest technological advances in difficult intubations | YES
4 | The unit shall be designed to provide a clear view of the vocal cords during intubation | YES
5 | The unit shall simplify intubation of the difficult airway | YES
6 | Trauma Airway | Yes
7 | Range in blade sizes allow coverage patient weights | Yes
8 | It shall have 60° angled blades | YES
9 | It should have pre-shaped stylet for quick intubation process | YES
10 | It should utilize minimum force for intubation | YES
11 | It should have anti-fog feature | YES
12 | The unit shall have integrated high resolution camera at good Position on blade to protect from blood and secretion | YES
13 | The unit shall mount on mobile carts and castors | YES
14 | Intubation of cervical spine immobilization | Yes
15 | Difficult Airway management and routine intubation | Yes
16 | Screen Size not less than 7 inch | LCD OR TFT
17 | Blades Size
18 | SIZE 2 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
19 | SIZE 2 SINGLE USE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
20 | SIZE 3 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
21 | SIZE 3 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
22 | SIZE 4 REUSABLE | Qty:1 ITEMIZED ( PRICE NOT TO BE INCLUDED IN OFFERED UNIT PRICE )
23 | SIZE 4 SINGLE USE INTEGRATED CAMERA & LIGHT SOURCE WITHIN THE BLADE | QTY : 80 MUST BE AN OPTIONAL ITEMIZED PRICE
24 | Made of medical grade plastic or stainless steel for less trauma handlingo teeth and soft tissue and easy | Specify
25 | Built in LED light for illumination | Yes
26 | Rechargable Battery on system | Yes
27 | Stylet | Qty:3
28 | Accessories | Any accessories, options and consumable items necessary to operate the offered system(s) must be clearly identified and priced separately
Company Stamp |  |  |  | Signature

## ITEM 134
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
134 | MOR20019 | 0 | LARYNGOSCOPE SET COMPLETE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE, FIBER OPTIC | Yes
2 | Macintosh blades | Yes
3 | Adult | 3, 4
4 | pediatric | 2
5 | Miller blades for sizes 00, 0, 1 for neonatal and infant | Included
6 | Handle | 2
7 | ILLUMINATION | LED
8 | Spare lamps | 5
9 | Materials of handle | Stainless steel
10 | Cleaning and sterilization for handle and blades | specify
11 | Case | Included
12 | Power supply AA battery | Yes
13 | Charger with rechargeable battery | Option
14 | Itemized Price for charger with battery | Yes
15 | Itemized Price for Lamp | Yes
16 | Itemized Price for handle | Yes
17 | Itemized Price for all blades | Yes
18 | All Accessories (itemized price) | Yes
Company Stamp |  |  |  | Signature

## ITEM 135
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
135 | MOR20022 | 0 | PUMP PCA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Easy to use , setup | yes
2 | Small Size , Rugged , lightweight unit. | yes
3 | Prime/Purge Control capability | Must
4 | Bolus Delivery | One touch - without interrupting current rate set
5 | Programmable to store drug library | yes
6 | Programmable unit for different protocol /profile | yes
7 | Delivery Modes | Continuous , Volume over Time , Bolus , Continous with Bolus
8 | Programming Units | ml, mg, mcg
9 | Specify drug reservoir type - for the PCA Pump | Cassette - Syringe
10 | Drug Concentration | 0.1 -100 mg/ml, 1-500 mcg/ml
11 | Continuous Rate | 0 - 50 ml/hr, 0 - 5,000 mg/hr
12 | Flow Rate | 0.1 - 9.9 ml , 0 - 990 mg ,Bolus , 0-2,000 mg , 0-10,000 mcg
13 | Demand Dose Lockout | 1 Min - 24 Hrs.
14 | Demand Dose per Hour | Specify Qty.
15 | Drive Mechanism | Specify Provided Mechanism
16 | Drug library | Specify How many drugs provided .
17 | LCD display | yes with high level of contrast , brightness , visibility
18 | LCD Size | Specify Size
19 | Protocol | Stores up to 500 protocols : Specify Provided
20 | Patient Information | Specify the Data related to patient that can be Regsitered
21 | Display | Run Indicator Light, Bolus Infusing, Pause Indicator, Infusion Complete,Volume Infused : specify
22 | Patient Activation Set | Yes : Specify Mechanism
23 | Event Memory | Yes.
24 | Alarms
25 | Occlusion | Yes
26 | Infusion Near End | Yes
27 | End Of Infusion | Yes
28 | Low Battery | Yes
29 | Alarm Volume | High, medium, low
30 | Graphs | PCA dose graph : Specify your`s
31 | Reports | Provide full information about generated report
32 | Security of Drugs | Specify Mechanism in your system
33 | Accuracy | ±3%
34 | Occlusion Detection | Specify Type & Mechanism
35 | Battery Operations | Specify the Max. time unit will work on battery.
36 | RS-232 , USB , other Interface | Specify Provided
Company Stamp |  |  |  | Signature

## ITEM 136
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
136 | MOR20026 | 0 | TOURNIQUET SYSTEM ANESTHESIA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LED or LCD Digital display | Yes.
2 | monitoring shows actual cuff pressure | Yes.
3 | Fast cuff inflation and deflation | Yes.
4 | Limits pressure to max. 600mmHg approx. | Yes.
5 | Auto & Manual deflate setting facility | Specify.
6 | Flush-Tester | Yes.
7 | Visual and audible alarm to detect leaks, pressure deviations and low battery voltage | Yes.
8 | Dual channel | Yes.
9 | Extension Hoses | Yes.
10 | Color coded Coil Extension Hoses | Yes.
11 | Easy attachment and detachment for safe and leak-free inflation | Yes.
12 | Automatic Timer | Yes.
13 | automatic time recording after cuff inflation | Yes.
14 | provides elapsed cuff inflation time | Yes.
15 | Stand with basket | Yes.
16 | very stable | Yes.
17 | Rides on 4 antistatic dual wheels, 2 equipped locking brakes | Yes.
18 | Automatic switchover from off line to on line(vise versa). | Specify.
19 | Battery back-up system
20 | Double cuffs for I.V.-Regional Anesthesia Qty: 2 for each child and adult | Yes.
21 | Double cuffs for I.V.-Regional Anesthesia Qty: 2 for Arm / 46 cm | Yes.
22 | Double cuffs for I.V.-Regional Anesthesia Qty: 2 for Leg / 61 cm | Yes.
Company Stamp |  |  |  | Signature

## ITEM 137
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
137 | MOR20029 | 0 | WARMER BLOOD |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | HEAT EXCHANGER
2 | Technology. Dry heat, sealed water cirduit, water bath, warm forced air, direct radial conduction, and magnetic induction,. | Specify
3 | MAX. TEMP. SETTING degree C | 38 - 42 degree C, Specify
4 | FLOW RANGE, mL/min, With 10OC input to at least 35OC | > OR = 70
5 | APPLICATIONS
6 | IV , Irrigation | Specify
7 | DISPLAY | LCD, LED, TFT, Specify
8 | Temp range, C. Degree | From 35 TO 41 C Degree
9 | Increments, OC | 0.5 C degree
10 | HEATER POWER, W | Specify
11 | PRIMARY TEMPERATURE
12 | CONTROL | Electronic thermostatic
13 | HIGH-TEMPERATURE CUTOFF, OC | 40 - 43, Others specify
14 | HIGH-TEMPERATURE ALARM Audible, Visual, Alarm test | Specify
15 | ON/OFF SWITCH | Yes
16 | MOUNTING
17 | Freestanding, IV pole attachment | Specify
18 | ACCESSORIES
19 | 500 disposable set, | Yes
20 | life for disposable sets | 6 M0NTH
21 | Power Supply | 220 Vac , 60 Hz
22 | Other Specifications | FDA , CE , Iso Approved
Company Stamp |  |  |  | Signature

## ITEM 138
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
138 | MOR20030 | 0 | WARMER BLOOD RAPID INFUSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | HEAT EXCHANGER
2 | Technology. Dry heat, sealed water cirduit, water bath, warm forced air, direct radial conduction, and magnetic induction,. | Specify
3 | MAX. TEMP. SETTING degree C | 38 - 42 degree C, Specify
4 | FLOW RANGE, mL/min, With 10OC input to at least 35OC | > OR = 1000
5 | APPLICATIONS
6 | IV , Irrigation | Specify
7 | DISPLAY | LCD, LED, TFT, Specify
8 | Temp range, C. Degree | From 35 TO 41 C Degree
9 | Increments, OC | 0.5 C degree
10 | HEATER POWER, W | Specify
11 | PRIMARY TEMPERATURE
12 | CONTROL | Electronic thermostatic
13 | HIGH-TEMPERATURE CUTOFF, OC | 40 - 43, Others specify
14 | HIGH-TEMPERATURE ALARM Audible, Visual, Alarm test | Specify
15 | ON/OFF SWITCH | Yes
16 | MOUNTING
17 | Freestanding, IV pole attachment | Specify
18 | ACCESSORIES
19 | 500 disposable set, | Yes
20 | life for disposable sets | 6 M0NTH
21 | Power Supply | 220 Vac , 60 Hz
22 | Other Specifications | FDA , CE , Iso Approved
Company Stamp |  |  |  | Signature

## ITEM 139
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
139 | MOR30001 | 0 | ASPIRATION SYSTEM THROMBUS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | System components:
2 | Aspiration pump
3 | Aspirator tubing
4 | Liner set
5 | Catheter
6 | Specification of the aspiration pump:
7 | Max. vacuum kPa/mmHg: -95/-713
8 | Flow [l/min]: 40, 50, 60
9 | Drive unit: Long-lasting piston/cylinder technology
10 | Rpms [at 40l/min, 50l/min, 60l/min]: 63, 79, 102
11 | Noise level [dB(A)][at 40l/min, 50l/min, 60l/min]: 37.8, 38.4, 39.8
12 | Vacuum regulator: Membrane
13 | pump shall be mounted on a trolley
14 | Catheter specifications:
15 | Inner Diameter mm (in): 1.52 (0.060)
16 | Outer Diameter F (mm) [in]: 6.0F (2.01) [0.079] / Dist: 5.4F (1.81) [0.071]
17 | Effective length (cm): 132
18 | Overall length (cm): 137
19 | Autoclavable
20 | Bidder shall list the specifications of the liner set
21 | Bidder shall list the specifications of the aspirator tubing
22 | Bidder shall list all the other accessories
23 | Compliance with standards & legislation:
24 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
25 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
26 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
27 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
28 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 140
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
140 | MOR30013 | 0 | MONITOR CARDIAC OUTPUT THROUGH ESOPHAGUS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Noninvasive Hemodynamic Monitoring system
2 | The system must be able to measure the following parameters
3 | SV/SI (Stroke volume / index)
4 | CO/CI (Cardiac output / index)
5 | HR (Heart Rate)
6 | PEP (Pre-ejection period) – An indicator of Iso-Volumetric Contraction Time.
7 | FTC (Flow Time Corrected) – a Preload indicator for spontaneously Breathing Patients
8 | TFC (Thoracic Fluid Contents) – Indicator of Total Chest Fluid
9 | ICON (Index of Contractility) – Left Ventricular global contractility Index.
10 | STR (Systolic time ratio) – A qualitative indicator of Ejection Fraction.
11 | SVV (Stroke Volume Variation) – Fluid Responsiveness Indicator
12 | Calculate SVR (Systemic vascular resistance / index)
13 | Calculate DO2 (Delivered Oxygen)
14 | HRV/C (Heart Rate Variability/Complexity).
15 | The system should be usable for all patient size Adults, Pediatric & Neonates
16 | The technology used should be totally non-invasive to facilitate the use in emergency situations
17 | Each parameter should be displayed with it's normal range
18 | The monitor should be handheld, & Easy setup.
19 | The monitor should be connectable to a software provides parameter trending, Passive Leg Raising Test, Printing, ...etc
20 | The monitor should provide an online signal quality indication
21 | A transportation trolley, Screen and Printer should be provided with the monitor.
22 | The monitor should have a built-in battery and large storage memory.
23 | Bidder shall list all the available accessories
24 | Compliance with standards & legislation:
25 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
26 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
27 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
28 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcomes and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 141
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
141 | MOR30015 | 0 | PUMP BALLON |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TIMING LOGIC
2 | ECG TRIGGER
3 | AUTO TIMING/TRIGGER
4 | INFLATION VOL ADJUST INCREMENTS,CC | YES, SPECIFY
5 | TRIGGERING MODES
6 | ECG, PRESGGERING, ASYNCHRONOUS TRIGGERING | YES
7 | TIMING SHOULD BE ADJUSTABLE | YES
8 | MAX PUMP RATE, BPM | 200
9 | FREQUENCY WEANING | YES
10 | SIGNAL INPUTS
11 | ECG | YES
12 | PRESSURE | YES
13 | AUTO CONDENSATION REMOVAL | Yes
14 | DISPLAY
15 | TYPE / SIZE | LCD,TFT,COLORS
16 | PARAMETER | inFlation / deflation intervals on the pressure waveform to allow timing adjustments
17 | HEART RATE ALARM | Optional
18 | QRS INDICATOR | Yes
19 | PRESSURE ALARM | Yes
20 | LEAD SELECT | 3-lead
21 | AL ARMS
22 | TRIGGER LOSS/CHANGE | Yes
23 | VAC/PRESSURE LOSS | Yes
24 | BALLON DISCONNECT | Yes
25 | BALLON LEAK | Yes
26 | BALLON OVERINFLATE | Yes
27 | SYSTOLIC INFLATE | Yes
28 | LOADING FAILURE | Yes
29 | HIGH PRESSURE | Yes
30 | POWER LOSS
31 | BALLOON SIZES, CC | SPECIFY
32 | COMPATIBILITY | SPECIFY
33 | HELIUM CYLINDER SIZE
34 | VOLUME, L | SPECIFY
35 | PRESSURE | SPECIFY
36 | POWER SUPPLY
37 | LINE POWER, VAC | 220, 60Hz
38 | CHARGING TIME, HR | SPECIFY
39 | OPERATING TIME, HR | One
40 | ACCESSORIES | ALL REQUIRED
Company Stamp |  |  |  | Signature

## ITEM 142
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
142 | MPH40001 | 0 | REFRIGERATOR MEDICATION 700L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Floor standing
2 | CONFIGURATION | Storage of pharmaceuticals under safe, controlled conditions
3 | CAPACITY IN LITERS | 700 L Approximatlly
4 | SHELVES | 4 or 5 heavy gauge zinc plated /powder coated
5 | DOORS
6 | DOOR TYPE | Glass door
7 | DOOR SEALING SYSTEM | Yes, para-magnetic
8 | DOOR LOCK | Yes
9 | TEMPERATURE CONTROL | 2 degree Centigrate TO 10 degree Centigrate
10 | INTERIOR | Stainless steel
11 | EXTERIOR | Powder coated enemal paint
12 | TEMP-ALARM . And Door Open Alarm | Yes, High & Low / and Yes for Door alarm
13 | TEMPERATURE INDICATOR | Yes, Digital display
14 | TEMP. CHART RECORDER | Yes
15 | POWER FAILURE ALARM, AUDIO & VISUAL | Yes
16 | RAPID TEMP. RECOVERY AFTER DOOR OPENINGS | Yes
17 | INSULATION | CFC free
18 | INSIDE LIGHT | Yes
19 | CASTERS | Yes 4, (two lockable)
20 | POWER SUPPLY | 220 V, 60 Hz
21 | OTHER SPECIFICATION | FDA OR CE
Company Stamp |  |  |  | Signature

## ITEM 143
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
143 | MPH40005 | 0 | REFRIGERATOR MEDICATION 160L |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Floor standing
2 | CONFIGURATION | Storage of pharmaceuticals under safe, controlled conditions
3 | CAPACITY IN LITERS | 160 L Approximatlly
4 | SHELVES | Specify
5 | DOORS
6 | DOOR TYPE | Glass door
7 | DOOR SEALING SYSTEM | Yes, para-magnetic
8 | DOOR LOCK | Yes
9 | TEMPERATURE CONTROL | 2 degree Centigrate TO 10 degree Centigrate
10 | INTERIOR | Stainless steel
11 | EXTERIOR | Powder coated enemal paint
12 | TEMP-ALARM . And Door Open Alarm | Yes, High & Low / and Yes for Door alarm
13 | TEMPERATURE INDICATOR | Yes, Digital display
14 | TEMP. CHART RECORDER | Yes
15 | POWER FAILURE ALARM, AUDIO & VISUAL | Yes
16 | RAPID TEMP. RECOVERY AFTER DOOR OPENINGS | Yes
17 | INSULATION | CFC free
18 | INSIDE LIGHT | Yes
19 | CASTERS | Yes 4, (two lockable)
20 | POWER SUPPLY | 220 V, 60 Hz
21 | OTHER SPECIFICATION | FDA OR CE
Company Stamp |  |  |  | Signature

## ITEM 144
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
144 | MPH50001 | 0 | CABINET MEDICATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | يستخدم لحفظ الادوية. | نعم
2 | يحتوي على 4 أدراج بحد أدنى. | نعم
3 | مدهون بدهان حراري غير قابل للصدأ | نعم
4 | يحتوي على قفل. | نعم
5 | شكل عصري وعملي | نعم
6 | وجود فهرسة | نعم
7 | تقديم ضمان شامل لا يقل عن 5 سنوات | نعم
8 | الابواب زجاجية | نعم
9 | الابعاد | طول 100 سم, عرض 50 سم ,ارتفاع 180 سم
Company Stamp |  |  |  | Signature

## ITEM 145
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
145 | MPH50002 | 0 | CABINET NARCOTIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Storage of Narcotics
2 | FINISH | Enamelled
3 | NUMBER OF DOORS
4 | door type | piano hinge
5 | locks | Yes separate for each door
6 | DIMENSION | 350mm x 250mm x 200mm H. approximately
7 | OTHER SPECIFICATION
8 | ALARMS
9 | visual | Yes
10 | audible | Yes
11 | POWER
12 | vac
13 | hz
Company Stamp |  |  |  | Signature

## ITEM 146
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
146 | MPH50005 | 0 | CART MEDICATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mobile unit, impact resistant plastic top with raised edges. | Drawers unit with large and small drawers
2 | Hard Top Surface | Yes.
3 | Swing out side storage unit with tray divider. | Sharps container holder, and tilt out bin.
4 | Multi-purpose overbridge unit . | With universal clamp, small. large metal bin, label/tape dispenser, & bins w/ cover.
5 | Dual-sided patient drawer | Specify
6 | Waste Container | Yes
7 | Adjustable Dividers | Yes
8 | Writing Surface | Yes
9 | Removable Trays | Yes
10 | Electronic Locking | Yes
11 | Cart Bumper | Yes.
12 | Four castor ,two of them with lock | Yes.
13 | other specifications | FDA, CE OR OTHER SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 147
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
147 | MPH50006 | 0 | CART DISPENSING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | PHARMACY DISPENSING CART
2 | MOBILE | YES
3 | CASTORS | MEDIUM SIZE , ANTI-STATIC
4 | NUMBER OF DRAWERS
5 | SECTIONS PER DRAWERS | 6 - 8 SECTIONS ( PARTITIONS )
6 | PARTITIONS | PROVIDE DIFFERNET SIZES FOR EACH DRAWER TO ACCOMODATE DIFFERNET PHARMACY RELATED ITEMS
7 | LARGE BOTTOM DRAWER
8 | TOP SURFACE | FLUSHED TYPE FOR INFECTION CONTROL
9 | BASKET | SIDE BASKET
10 | BOTTOM BUMBER | PREFFERED
11 | MATERIAL MAIN CART | ALUMINUM
12 | DRAWERS MATERIAL | ALUMINUM - SLOTTED FOR PARTIOTIONS
13 | PARTIRIONS | PLASTIC ADJUSTABLE
Company Stamp |  |  |  | Signature

## ITEM 148
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
148 | MPH50007 | 0 | SAFE NARCOTIC LARGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Applications | Narcotic Safe
2 | Dimensions | Approx. 35" H x 20" W x 19" D
3 | Weight | Approx. 300 Kg
4 | Double Door Narcotic Storage | Yes
5 | 1/2" solid steel door, sledgehammer & pry bar resistant. | Yes
6 | Auto door detent | Yes
7 | Thick 1" in diameter, chromed live locking bolts | Yes
8 | Spring loaded re-locker | Yes
9 | Scratch resistant | Yes
10 | shelves | approx. 8
11 | 1/2" diameter anchor holes | Yes
12 | Electronic Locking System | Prefered
13 | LCD screen for easy programming | Prefered
14 | Multiple registered users | Prefered
15 | Programmable time delay and time window | Prefered
16 | Penalty lockout | Prefered
17 | Single or dual control modes | Prefered
18 | Standard footprint for easy retrofit | Prefered
Company Stamp |  |  |  | Signature

## ITEM 149
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
149 | MPH50008 | 0 | SAFE NARCOTIC SMALL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Narcotic Safe
2 | Height | 15 Inches Approx.
3 | Width | 10 Inches Approx.
4 | Depth | 5 Inches Approx.
5 | Double door | Each door has an individual keys and locks
6 | High quality welded construction | Yes
7 | Steel shelves are removable | Yes
8 | Material | Steel
9 | OTHER SPECIFICATION | SPECIFY
Company Stamp |  |  |  | Signature

## ITEM 150
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
150 | MRH10006 | 0 | BALANCE APPARATUS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Country Of origin | Specify
2 | Model | Specify
3 | Manufacture | Specify
4 | FOR DETERMINING THE INDIVIDUAL WEIGHT BEARING LINE FOR | YES
5 | TRANSFEMORAL PROSTHESES | YES
6 | KNEE DISARTICULATION PROSTHESES | YES
7 | APPROX. DIMENSIONS WxDxH | 1200x500x900mm
8 | FULL ACCESSORIES | YES
9 | OTHER SPECS
Company Stamp |  |  |  | Signature

## ITEM 151
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
151 | MSE10043 | 0 | GAUGE ROOM PRESSURE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Portable, hand-held
2 | LCD | Yes
3 | PRESSURE MEASURING RANGES | from 0 to 800 mmHg
4 | VARIOUS PRESSURE UNITS | cmH2O, mmHg, mBar, Pascal etc.
5 | ABSOLUTE PRESSURE RANGE | (0 to 180) kPa
6 | PRESSURE ACCURACY | within ±2% of full scale
7 | BATTERY POWERED | yes
8 | POWER SUPPLY | yes
9 | VOLTAGE | 220-240 Volts
10 | FREQUENCY | 60 Hz
11 | REGULATORY COMPLIANCES | ISO, FDA, CE, or other. (Specifyin offer).
12 | SAFETY STANDARD | EMC,UL, IEC, CSA or other Intl. Standard, (Specify in offer).
13 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Signature

## ITEM 152
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
152 | MWD10002 | 0 | BED BURN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Dedicated Burns Bed
2 | INDICATIONS
3 | Moderate to Severe Burns | Yes
4 | Prevention of wounds | Yes
5 | Treatment of wounds up to stage IV | Yes, Specify stage
6 | Flaps | Yes
7 | Skin Graft | Yes
8 | Pain management | Yes
9 | Oedema management | Yes
10 | UNIT CONTROLS
11 | Integrated Bed Controls for all function | Yes, all integrated
12 | Nurse Controls
13 | Side Rails Nurse Controls :-
14 | Both side rails controls | Yes
15 | Head movement | Yes
16 | Foot movement | Yes
17 | Trendelenburg and Reverse | Yes
18 | Bed height adjustment | Yes
19 | Patient rotation hold | Yes
20 | Cardiac Chair | Yes
21 | Foot End Nurse Controls :-
22 | Integrated Foot board controls | Yes
23 | All of the above side rails bed motion functions | Yes
24 | scale at the foot end control panel | Yes
25 | Bed Exit alram | Yes
26 | All Therapies Control Panel | Yes
27 | warmer controls | Yes
28 | Patient Controls :-
29 | Integrated Side-Rails Controls | Yes
30 | Head section adjustment | Yes
31 | Foot section adjustment | Yes
32 | Integrated nurse call button | Yes
33 | TV/Radio Controls buttons | Yes
34 | BUILT-IN PATIENT WARMER
35 | Integrated Patient Warmer in mattress | option
36 | Adjustable temperature Level Control | option
37 | Steps of 3 degrees centigrates increament above rooms tempreture, up to nine | option
38 | INTEGRATES SLEEPING SURFACE
39 | integrated mattress with all controls within the foot board bed controls | Yes
40 | Deep air cylinders of not less than 11 Inches | Yes, specify
41 | Minimum of Sixteen Air Cylinders | Yes, Specify
42 | Lowest Patient Interface pressure | Yes
43 | Air Cylinders and cover material should be highly breathable | Yes, Specify
44 | Air cylinders and cover material should be highly bacterial barrier | Yes, Specify
45 | Air cylinders and cover material should be very smooth laminated surface to reduce friction and shearing forces | Yes, Specify
46 | Air Cylinders and cover material should have very high Moisture Vapor Transmission Rate | Yes, Specify
47 | THERAPIES System should provide the following therapies | Yes
48 | Low Air Loss Therapy | Yes
49 | Pressure Relief Therapy with minimal patient interface pressure | Yes
50 | Pulsation Therapy
51 | Pulsation interface pressure should be below 20 mmHg | Yes
52 | Automated Turning and Positioning
53 | Lateral Rotation of the mattress only but not the whole frame | Yes
54 | Controlled and Integrated mattress controls in the foot board controls | Yes
55 | 20 Degrees rotation on each side | Yes
56 | Ability to program lateral rotation including selection of the followings:- | Yes
57 | Period on each side individually | Yes
58 | Angle | Yes
59 | Hold time on each side individually | Yes
60 | Continuously changes pressure points to reduce interface pressure | Yes
61 | Minimises pain and discomfort associated with turning | Yes
62 | Shear Redcution | Yes
63 | Cushion Deflate functions
64 | Head Cushions Integrated Control to facilitate intubation | Yes
65 | Seat Cushions Deflate function to facilitate patient egress and bed pan | Yes
66 | Four Zone Air Suspension | Yes
67 | Four individually controlled sections to maximise comfort and meet the individual needs of the patient. | Yes
68 | Built-in Scale | Yes
69 | Accurate patient weight, weight trends and facilitate weight loss or gain | Yes
70 | Weighing scale controls from foot baord controls | Yes
71 | Instaflate | Yes
72 | Inflates mattress to its maximum firmness for ease of patient placement and care. | Yes
73 | CPR
74 | Fast movement to enable evacuating air from the mattress. | Yes
75 | Alarms
76 | Adiuable and Visual bed alram with indication of alarm identifiaction | Yes
77 | Bed function articulation
78 | Fowler 0-60? | Yes
79 | Knee Gatch 0-40? | Yes
80 | Trendelenburg +8? | Yes
81 | Reverse Trendelenburg -8? | Yes
82 | Reverse Trendelenburg -8? | Yes
83 | Bed Extender to accommodate tall patients | Yes
84 | Electrical Features
85 | Power Required, VAC | 220 V
86 | Frequency | 50 - 60 Hz
87 | Number of motors | Specify
88 | Double Insulation Frame | Yes
89 | Grounded | Yes
90 | Non-Conductive Side Rails | Yes
91 | Isolated IV Pole | Yes
92 | Other Specifications :-Built-in invertor as power back-up and during transportation | Yes
Company Stamp |  |  |  | Signature

## ITEM 153
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
153 | MWD10004 | 0 | BED ELECTRICAL PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Adjustable height | Yes
2 | Adjustable head | Yes
3 | Bed use as a stretcher | Yes
4 | Adjusts from flat to chair position | Yes
5 | Storage areas | Yes
6 | Paper roll dispenser | Yes
7 | Mattress Well padded and easy to clean | Yes
8 | Table chassis resembles in shape friendly object of natures suitable for pediatric patients | Yes
9 | Frame should be stainless steel . | Yes
10 | Handgrips removable | Yes
11 | Table should have castors, with brakes (diameter 15cm) | Yes
12 | Dimensions that suit pediatric patient approx. (50 hieght * 110 long )cm . | Specify
13 | New modern | Yes
14 | Patient weight above 30 KG . | Yes
15 | Capacity Scale . | Yes
16 | Infantometer . | Yes
17 | Working Surface of Vinyle . | Yes
18 | Three Drawers and Cupboard Below . | Yes
19 | CE or FDA approved . | Yes
Company Stamp |  |  |  | Signature

## ITEM 154
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
154 | MWD10006 | 0 | BED INTENSIVE CARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Intensive Care Bed/ ICU adult
2 | PATIENT CONTROLS | YES
3 | Type | specify
4 | Functions | Head High/ Low , Knee up / Down, Trendelenburg & Reverse trendelenburg, electrical flexafoot
5 | NURSE CONTROLS | YES
6 | Patient control lockout | YES
7 | Full low indicator | YES
8 | CPR MODE | Dual sided
9 | CPR control | YES
10 | graphical caregiver interface with touch screen for bed and mattress functions | YES
11 | AUTOMATIC CONTOUR | YES
12 | CHAIR MODE | Full Chair & Full Chair front Egress Position
13 | ALARM SYSTEM | Head angle alarm 30 and 45 degrees, bed Exit alarm, Brake off alarm
14 | TRENDELENBURG GAUGE | YES
15 | BackUp Function | Electro-hydraulic or other specify
16 | BED SURFACE TYPE | specify
17 | C-ARM APPLICATION | YES
18 | FRACTURE FRAME RECEPTACLES | YES
19 | ELECTRIC FOOT EXTENDER / RETRACTABLE | Yes / Electrical
20 | DRAINAGE BAG HOLDER | YES
21 | MAX. PATIENT WEIGHT, Kg | 230 kg or more
22 | OVERALL DIMENSIONS
23 | L x W, cm | L210 235x W93 -102 cm approx.
24 | Height, cm | 45 cm – 92 cm approx
25 | SIDE RAILS LENGTH
26 | Fraction of overall length | YES specify
27 | CASTERS
28 | SIZE (Diameter) cm | 12.5 cm or more specify
29 | FUNCTION | Swivel, brake, steer
30 | CENTRAL BREAK SYSTEM | YES
31 | BUMPERS | Corner/ Full perimeter
32 | REMOVAL HEADBOARD & FOOTBOARD | YES
33 | IV POLE MOUNT | MINIMUM 3 PROVISIONS
34 | INTEGRATED BED SCALE With patient weight statistics for up to7 days or more | YES
35 | Power supply | 220 Volts, 60 Hz
36 | Battery (Back-up) | OPTIONAL ITEMIZED
37 | Double insulation | YES
38 | Electric Shock Protection | Class 1, Type B
39 | Number of motors | 4 or more specify
40 | GROUNDED | YES
41 | Isolated (motor ground | YES
42 | Non-Conductive Side rails | YES
43 | AIR MATTRESS | Anti-bacterial, Anti-static, Moisture proof, fire resist & chemical-resist, tough & high durable
44 | INTEGRATED WITH THE BED | YES
45 | Dimensions 20cm or more height | Compatible with bed dimensions
46 | Controls are imbedded within the side rails | YES
47 | AUTOMATIC DEFLATE | Full inflation once CPR mode is pressed
48 | Max inflate function | YES
49 | X-ray Sleeve inside the mattress | YES
50 | other specifications | FDA approved CE marked
51 | Printout for the item offered with all technical parameters exactly matches the soft copy in the CD. | YES
52 | Supplier is an authorized representative of the manufacturer by SFDA and/or an official agent by Ministry of Trade. | Yes (Attach official documents)
53 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Signature

## ITEM 155
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
155 | MWD10007 | 0 | BED INTENSIVE CARE PEDIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | General FDA OR CE
2 | NUMBER OF SECTIONS | 4
3 | FRAMEWORK | Tubular steel
4 | PATIENT CONTROLS
5 | Type | Pendant / Rail Embedded Specify
6 | Number | 5 Appproximate
7 | Functions preset | Height High & Low, Head up & down , Knee up & Down, and Fowler
8 | NURSE CONTROLS
9 | Patient control lockouts | Yes
10 | CPR MODE | Dual sided
11 | POSITION CONTROL
12 | Trendelenburg | Specify
13 | Reverse Trendelenburg | Specify
14 | Auto Contour button | Yes
15 | INTEGRAL BED EXTENSION | Yes
16 | Removable head and footboard that give a 360 degree access. | yes
17 | MAX. PATIENT WEIGHT, kg. | 70 Kg or more
18 | OVERALL DIMENSIONS | aprox. 130 cm x 65 cm
19 | INTEGRATED BED SCALE | yes
20 | Side rail designed for easy cleaning, and visibility and for protecting baby from falling down | Yes
21 | Fraction of overall | length OR ful length
22 | CASTOR
23 | Diameter, cm | 12.5 cm - 15 cm
24 | Brake | Yes
25 | BUMPERS | Corners
26 | UNDER BED CLEARANCE | 150 mm
27 | IV POLE MOUNTS | Yes, 4
28 | ELECTRICAL FEATURES
29 | Power supply | 220-230 Volts
30 | Frequency | 60 Hz
31 | Double insulation | Yes
32 | Frame grounded | Yes
33 | Electric Shock Protection | Class 1, Type B
34 | Battery (Back-up) | yes, Specify
35 | Isolated IV pole | Yes
36 | Isolated transformer | Yes
37 | SAFETY STANDARD | IEC
38 | MATTRESS
39 | Application | Pediatric patient bed
40 | Mattress Ticking | On both sides
41 | Antibacterial | Yes
42 | Anti Static | Yes
43 | Moisture proof | Yes
44 | Fire Retardant | Yes
45 | Resists Puncture | Yes
46 | Tearing and chemicals resistant | Yes
47 | Dimensions | Compatible with bed dimensions
48 | OTHER SPECIFICATIONS | specify
Company Stamp |  |  |  | Signature

## ITEM 156
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
156 | MWD10013 | 0 | CABINET BEDSIDE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONSTRUCTION | Wooden laminated
2 | COLOUR & FINISH | Match with approved Hospital bed
3 | STORAGE DRAWER | 1
4 | NUMBER OF DOWN CABINET DOORS | 1 or 2, specify
5 | BASE & FRAME MATERIAL | Natural wood or steel, specify
6 | CASTORS | Yes, lockable
7 | DIMENSION IN CMS | 75(H) x 55(W) x 45(D)
8 | OTHER SPECIFICATION | FDA or CE
Company Stamp |  |  |  | Signature

## ITEM 157
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
157 | MWD10016 | 0 | CRIB NEWBORN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Baby Crib | Yes
2 | Configuration | Mobile
3 | Elevation Mechanism | Yes at least 20 Degree Up
4 | Material | High Quality S.S Frame
5 | Body ( Bed ) Material | High Quality Plastic or better
6 | Noise while moving | Less than 48 dB
7 | MATRESS | Yes
8 | STORAGE SHELF | Yes
9 | drwaers | Yes
10 | Castores | 4 two must be lockable
Company Stamp |  |  |  | Signature

## ITEM 158
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
158 | MWD10021 | 0 | STRETCHER EMERGENCY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Area of Use | Emergency
2 | Stretcher Type | Hydraulic
3 | Position Control | Dual sided foot pedal
4 | Trendelenburg | 12 degree APPROX.
5 | Reverse Trendelenburg | 12 degree APPROX.
6 | Emergency Trendelenburg Foot or Handle pedal | Yes, specify.
7 | Height Adjustment | Yes, specify range
8 | Back Rest | Adjustable Gas Spring Assisted
9 | Restraining straps | Yes
10 | Frame Material | Rigid tubular framework with epoxy painted steel.
11 | Patient platform | X-Ray translucent
12 | Safety Side rails | Integral Fold Down
13 | X-Ray cassette access | Yes
14 | Oxygen Cylinder Holder with Cylinder | Accept Size E
15 | Bumper | Yes
16 | Central Locking Bracking System | Four Corner Activation
17 | IV Pole | Yes
18 | Spare IV rod | Yes
19 | Sockets number
20 | Large Storage Area | Yes
21 | Maximum Patient Weight | 220kg OR more
22 | Four 20 cm Casters | All Lockable Antistatic
23 | Mattres | Yes
24 | Thickness cms | 6.cm or more
25 | Monitor Shelf | Yes
26 | 5th Wheel | Yes
Company Stamp |  |  |  | Signature

## ITEM 159
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
159 | MWD10022 | 0 | STRETCHER RECOVERY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | AREAS OF USE | Emergency / Re-covery bay
2 | FRAME MATERIAL | Epoxy painted steel
3 | STRETCHER TYPE | Hydraulic
4 | FOOT PEDAL OPERATED | Yes, both sides
5 | HEIGHT ADJUSTMENT | Yes
6 | TRENDELENBURG | 15 approx.
7 | REVERSE TRENDELENBURG | 15 approx.
8 | HEAD ADJUSMENT | Yes
9 | STRETCHER SURFACE | Radio translucent
10 | X-RAY CASSETTE HOLDER | Yes, full length
11 | SIDERAILS | Yes
12 | OXYGEN TANK HOLDER | Yes
13 | RESTRAINING STRAPS | Yes
14 | BUMPERS | Yes, all corners
15 | IV POLE SOCKETS | 2 or 4
16 | CASTORS | 4 or 5, lockable
17 | DIAMETER | 20 cm approx.
18 | MAXIMUM PATIENT WEIGHT | Approx 200 Kg
19 | ACCESSORIES | All neceessory accessories
20 | MATTRESS | Yes, durable and easy to clean
21 | PLATFORM DIMENSION IN MM | 2000 X 600 approx.
22 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 160
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
160 | MWD10023 | 0 | STRETCHER TRANSPORT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | AREAS OF USE | Emergency / Re-covery bay
2 | FRAME MATERIAL | Epoxy painted steel
3 | STRETCHER TYPE | Hydraulic
4 | FOOT PEDAL OPERATED | Yes, both sides
5 | HEIGHT ADJUSTMENT | Yes
6 | TRENDELENBURG | 15 degree approx.
7 | REVERSE TRENDELENBURG | 15 degree approx.
8 | HEAD ADJUSMENT | Yes
9 | STRETCHER SURFACE | Radio translucent
10 | X-RAY CASSETTE HOLDER | Yes, full length
11 | SIDERAILS | Yes
12 | OXYGEN TANK HOLDER | Yes
13 | RESTRAINING STRAPS | Yes
14 | BUMPERS | Yes, all corners
15 | IV POLE SOCKETS
16 | CASTORS | 4 or 5, lockable
17 | DIAMETER | 20 cm approx.
18 | MAXIMUM PATIENT WEIGHT | 220 Kg
19 | ACCESSORIES | All neceessory accessories
20 | MATTRESS | Yes, durable and easy to clean
21 | PLATFORM DIMENSION IN MM | 2000 X 800 approx.
22 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 161
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
161 | MWD10024 | 0 | TABLE OVERBED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MOVEMENT | Top tilting
2 | TOP MATERIAL | Laminated
3 | TOP SURFACE | Easily cleaned, laminate
4 | TOP SURFACE DIMENSION | 900 x 400 mm approx.
5 | SUPPORT ARM BENEATH TOP SURFACE | Yes
6 | TOP HEIGHT CONTROL | Spring
7 | HEIGHT ADJUSTMENT | 750 - 1100 mm approx.
8 | FRAME | Steel powder varnish
9 | CASTORS | Yes with lockable mechanism
10 | COLOUR | Color to match with the bed available at site.
11 | OTHER SPECIFICATION
Company Stamp |  |  |  | Signature

## ITEM 162
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
162 | MWD20002 | 0 | BED INTENSIVE CARE BARIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The Bed Shall Be Fully Motorized
2 | The Bed Controls Shall Be easily accessible
3 | It Shall Have One-Touch Control To Move The Patient Easily Into A Dining Chair Position. All Controls Shall Use International Symbols Where Applicable
4 | All Controls Shall Use International Symbols Where Applicable
5 | It Shall Have Shear-Less Pivot Mechanism (The Head Section Articulates Upward, The Back Section Moves Outwards) With Auto Contour (The Knee Section Moves Upward).
6 | It Shall Have A Central Braking System
7 | A Single Press Of The Brake Bar Shall Lock All The Wheels.
8 | An Auditory Alert Shall Sound When The Brakes Are Released And The Bed Is Still Plugged In
9 | It Shall Have Instant CPR Release Lever And Include A Dampened Descent, Gently Lowering The Patient And Reducing Caregiver Back Strain.
10 | It Shall Incorporate A Built In Digital Weight Scale
11 | It Shall Have Higher Headboard To Make Moving The Bed Easier
12 | It Shall Have Easy-To-Read Line Of Sight Angle Indicators For Degree Of Head Section And Trend And Reverse Trend.
13 | The Bed And Mattress Shall Be Designed To Address Main Entrapment Areas.
14 | It Shall Have An Isolation Transformer
15 | It Shall Have Embedded Caregiver Controls-Electric Adjustment For Height, Head, Knee, Foot, Trend/Reverse Trend And Dining Chair
16 | It Shall Have Patient Controls For Head/Knee With Nurse Call And Lighting
17 | It Shall Have Nurse Call-Embedded In Side Rail For Patient And Caregiver
18 | It Shall Have Foot Extender Pad With Fire Barrier
19 | It Shall Be Supplied With Standard IV Pole With Dual IV Hooks
20 | Rails And Bumpers Should Be Provided On The Bed To Avoid Impact And Abrasion Damage To The Bed And Wall Surfaces, Doorframes And Furniture.
21 | The Side Rails Shall Fold Down Or Tuck-Away For Easier Patient Transfer
22 | The Headboard And Footboard Shall Be Removed Easily For Quick Access To The Patient’S Head In An Emergency.
23 | There Shall Be A Minimum Of Four IV Receptacles And Should Be Provided On Both Sides Of The Bed.
24 | There Shall Be Multiple Drainage Bag Hooks Located On Each Side Of Bed.
25 | Electronic function lock-out control
26 | Frame Width: Approx. 100 Cm
27 | Frame Length With Head And Foot Board: Approx. 230 Cm
28 | Sleep Deck:
29 | Width: Approx. 110 Cm
30 | Normal Length: Approx. 200 Cm
31 | It Shall Be Supplied With A Comfortable Mattress With Fire Barrier
32 | The Mattress Shall Have The Following Features:
33 | The Mattress Shall Be Designed For Comfort And Protection.
34 | The Cover Shall Be Anti-Microbial, Anti-Fungal, And Promotes Skin Respiration
35 | It Shall Be Made Wider To Decrease The Gap Between The Mattress And The Side Rail
36 | The Surface Shall Also Have A Four-Way Stretch Cover To Prevent Shearing And To Protect The Skin
37 | Mattress Ticking Seams Shall Be Rf Welded For Added Durability
38 | Multi-Layered Mattress, Including Firm Perimeter Foam, For Precise Body Support
39 | It Shall Have A Pressure Redistribution Surface To Provide Low Air Loss Therapy, Static And Alteration Therapy Modes And Passive Massage Action To Aid In The Increase Of Capillary Blood Flow.
40 | The Bed Shall Have The Following Features:
41 | Bed Clearance Under Frame: Approx. 15 cm
42 | Low Position Floor To Deck: Approx. 50 cm
43 | High Position Floor To Deck: Approx. 93 cm
44 | Maximum Head Elevation: Approx. 70 Degrees
45 | Maximum Trend And Reverse Trend: Approx. 15 Degrees
46 | Caster Size: Approx 15 cm
47 | Maximum Weight Limit (Safe Working Load): Approx. 450 Kg
48 | Removable headboard
49 | Compliance with standards & legislation:
50 | The system must comply with the Electrical safety standards for electrical safety IEC-60601
51 | Should have a FDA approval and/or CE Mark & SFDA Registration, where applicable. List any other international standards (CE, UL, TUV, CSA), if any.
52 | All electrical connections and plugs should be hospital grade and follow international, local and hospital requirements.
53 | Provide hard/soft copies of the operation and maintenance manuals as per the tender terms and conditions
54 | All other basic accessories deemed necessary that are not mentioned in this specification but are required for full function and highest clinical outcome and output of the equipment must be included.
Company Stamp |  |  |  | Signature

## ITEM 163
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
163 | NMG10012 | 0 | ‫FREEZER 21 BODY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | 1-GENERAL :
2 | A-REFRIGERATOR, MORTUARY, 21 BODY | Delivery & installation of total of 21 body capacity mortuary Refrigerator.
3 | 1-Number of Unit | 1 (3 X 7)
4 | 2-Tier Quantity | 7
5 | 3-Tier Capacity | 3
6 | 4-Complete with lucks & keys for each chamber | Yes
7 | 5-End opening | Yes
8 | 6-Ambient Temp | 50º C or above
9 | 7-Temprature range | 0 to-20 adjustable. digitally display with 0.1 C accuracy
10 | 8-Compressor | 2 Remote compressor (according to site condition)
11 | 9-Compressor’s synchronised | Yes “Automatic change over switch” for both compressors
12 | 10-Full height door with inner plastic doors “Gasket” | Yes
13 | 11-Roller frame assembly | Yes
14 | 12-Cardholder | Yes
15 | 13-Frame heating | Yes
16 | 14-Doors | 21 single door easy replaceable hollow chamber seals.
17 | 15-Door opening angle, degrees | 180° and the direction depend on the site condition
18 | 16-Each unit shall have dual auto switchover system | Yes
19 | 17-Odor dissolving device mounted inside the unit. Chemical free – ionization principle | Yes
20 | 18-All surface finishes shall be resistant to abrasive chemical and detergent and corrosion resistant | Yes
21 | 19-Morgue, lifter and trolleys are compatible with the system | Yes
22 | 20-Odorless (ventilation Sys.) for mortuary building | Yes
23 | 21-Interior lighting – splash proof | Yes
24 | 22-Body Tray | 21 removable Body Tray’s are manufactured from stainless steel sheet with Polished pressed construction seamless shaped with tubular carrying handles with compatible drainage
25 | 23-Anti Ice | Yes
26 | 24-Refrigerant gas to be CFC free | Yes
27 | 2-MORGUE UNITS
28 | A-CONSTRUCTION
29 | 1-Exterior: 20 gauge polished stainless steel | Yes
30 | 2-Interior: 20 gauge polished stainless steel | Yes
31 | 3-Door: 20 gauge polished stainless steel | Yes
32 | 4-Inside Floor: 20 gauge polished stainless steel | Yes
33 | 5-Insulation: Foamed in place polyurethane insulation. Full 80 mm thick rigid polyurethane CFC free | Yes
34 | 6-Gasket: Extruded vinyl, welded corners, replaceable | Yes
35 | 7-handles with keylock Hardware: metal fittings and each door is provided with pull | Yes
36 | 8-Door heater: Perimeter anti-sweat device, replaceable | Yes
37 | 9-Frame: Roller frame assembly | Yes
38 | 10-Removable trays: Construction one piece, welded with tubular hand holds at each end. | Yes
39 | 11-Interior lighting | Yes
40 | 12-Inner safety release | Yes
41 | 13-Body tag holder on front door | Yes
42 | 14-Easy clean non-corroding surface | Yes
43 | B-REFRIGERATION
44 | 1-Refrigeration system shall be remote mounted in a mechanical room and shall include auto switchover | Yes, calculations are to be based on 30 meters or more (according to the mechanical room location) , and 50C ambient temperature.
45 | 2-The performance shall be at a maximum ambient | Yes
46 | 3-Compressor | 2 remote compressors semi-hermetic or hermetic compressors
47 | 1-Discharge vibration elemint | yes
48 | 2-Noisiness | 58Db or less
49 | 3-Air flow | 7200 m3 /h or more
50 | 4-Cooling capacity | 10 kw or more
51 | 4-It must accomplish the function of:
52 | 1-it should provide complete back up protection for the cooling system of the Refrigerator two complete systems with duplicate condensing units and evaporators (fan coil) | Yes
53 | 2-Each of these system are enough to cool the cabinet independently. So if one fails for any reason the other takes over automatically to maintain temperature | Yes
54 | C-TEMPRATURE ALARM SYSTEM | Complete temprature system which shall provide an audible and visual alarm
55 | 1-Alarm on each body chamber | Yes
56 | 2-Audio visual alarm for temperature and power failure. Digital display. | Yes
57 | Audio visual alarm for temperature and power failure. | Yes
58 | 3-Bright LED digital display with high resolution | Yes
59 | 4-Remote alarm connection | Yes
60 | 5-Temperature uniformity for all colding chambers | Yes
61 | D-RECORDER ON EACH REFRIGERATOR
62 | 1-7 day circular recording to meet regulatory standards | Yes
63 | 2-New 7-inches easy to read charts with temp Hi-Lite recording track area | Yes
64 | 3-Records temperature in degree Celsius and chart changes | Yes
65 | 4-Touch pad operation for calibration and chart changes | Yes
66 | 3-POWER SUPPLY | 380Volts – 60 Hz.
67 | 4-OTHER SPECIFICATIONS
68 | A-Technical Documentation | Original sets of manufacturer documentation including operation and service manual in English & Arabic
69 | B-Spare Parts | Availability of spare parts guaranteed for minimum period ten years…
70 | Compressor cover | Yes
Company Stamp |  |  |  | Signature

## ITEM 164
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
164 | NMG10028 | 0 | STRETCHER MORTUARY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MORTUARY TRANSPORT STRETCHER, CONCEALED CADAVER
2 | Stable torsion free, rectangular tube frame steel construction | Yes
3 | Upper part protected by disinfection proof custom fitted washable vinyl cover. | Yes complete body will be concealed
4 | Fixed mounted immersible torsion free steel frame to fit the vinyl cover | Yes
5 | Easy adjustable to hospital bed height | Yes
6 | hydraulic operated | Yes
7 | One-person operation | Yes
8 | Castors | Yes, 4 castors .Approx 2 with break, 2 swiffeling
9 | Maximum weight | 250 Kg approx.
10 | Body tray | removable stainless steel body tray 304 with drainage. Seamless shaped with carrying handles.
11 | Other Specifications | Compatible to morgue system
Company Stamp |  |  |  | Signature