## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | BODY LIFTER
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | FREEZER MORTUARY ROOM L20 W20 H4
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | PORTABLE MORTUARY RACK
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | BODY LIFTER MANUAL
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | FREEZER MORTUARY ROOM L20 W35 H7
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | FREEZER MORTUARY ROOM L28 W28 H5
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | FREEZER MORTUARY ROOM L20 W18 H4
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | FREEZER MORTUARY ROOM L20 W24 H4

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4229270300300 | BODY LIFTER |  | MEDICAL GROUP | 10
2 | 4226180601300 | FREEZER MORTUARY ROOM L20 W20 H4 |  | MEDICAL GROUP | 1
3 | 4226210000000 | PORTABLE MORTUARY RACK |  | MEDICAL GROUP | 458
4 | 4226210000100 | BODY LIFTER MANUAL |  | MEDICAL GROUP | 10
5 | 4226180601400 | FREEZER MORTUARY ROOM L20 W35 H7 |  | MEDICAL GROUP | 1
6 | 4226180601500 | FREEZER MORTUARY ROOM L28 W28 H5 |  | MEDICAL GROUP | 1
7 | 4226180601600 | FREEZER MORTUARY ROOM L20 W18 H4 |  | MEDICAL GROUP | 1
8 | 4226180601700 | FREEZER MORTUARY ROOM L20 W24 H4 |  | MEDICAL GROUP | 1
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | MEDICAL GROUP | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
|  | NON MEDICAL | 0 | 0 | [5 empty] | If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4229270300300 | 0 | BODY LIFTER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mortuary transport stretcher, concealed cadaver | Yes
2 | Stable torsion free, rectangular tube frame steel construction | Yes
3 | Resistant to corrosion and easy to clean | Yes
4 | Dimensions (MM) | Approx. 2300L X 700W
5 | weight
6 | Minimum loaded weight | 200 Kg.
7 | Maximum unloaded weight | approx. 100kg
8 | Height | Easy adjustable to hospital bed height
9 | Maximum Height | Up to the height position of the Rack
10 | Fork | matching with tray
11 | NO. Of Forks | 2
12 | Reversible forks for increased lifting height | Yes
13 | Castors
14 | 4 Castors, 2 with break, 2 swivelling
15 | Size of castors | Specify
16 | Operation | Hydaulic and Electric
17 | One-person operation | Yes
18 | Built in battery | yes
19 | Chargerable battery | Yes
20 | 220 V / 60HZ | Yes
21 | Upper part protected by disinfection proof custom fitted washable vinyl cover. | Yes complete body will be concealed
22 | Fixed mounted immersible torsion free steel frame to fit the vinyl cover | Yes
23 | Tray
24 | Body | Removable stainless steel body tray 304 with drainage. Seamless shaped with carrying handles.
25 | Other Specifications | Compatible to morgue system
Company Stamp |  |  |  | Company Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 4226180601300 | 0 | FREEZER MORTUARY ROOM L20 W20 H4 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | Storage of Bodies
2 | Dimension (please submit drawing based on your site visit and actual dimension status) | Yes
3 | Body structure
3.1 | Internal wall cladding and external wall of the frontal side | stainless steel
3.2 | Type of stainless steel | AISI 314, anti-bacterial, anti-corrosive, scratch resistance.
3.3 | Panel thickness | Metal sheet thickness at least 0.8 MM
3.4 | Insulation CDC free polyurethane rigid foam approx. 100mm | Yes
3.5 | Density 40 to 45 kg/ m3 foamed | Yes
3.6 | Panels fastened together by means of cam lock actuated by allen key | Yes
3.7 | Jointless corners to avoid leakage | Yes
4 | Door
4.1 | Heavy duty sliding door | Yes
4.2 | Construction similar to wall panels with insulation | Yes
4.3 | Pivot Type | Yes
4.4 | Door Handle & Latch assembly:- Automatic door - open cycle | Yes
4.5 | closure System with Internal Safety Release | Yes
4.6 | Mechanism on inside of room, preventing personnel from being locked from outside | Yes
4.7 | Door Panel Thickness | 100 MM Approx.
4.8 | Door Size | According to site and application required
4.9 | Door Lock | Yes, with failsafe feature/ sliding door
4.1 | Gasket with enough magnetic force to form positive airtight seal | vinyl type with magnetic steel core. Gasket on both side of door Adjustable rubber wiper gasket on the bottom
4.11 | Air Curtain | on the door
5 | Roof
5.1 | Installation method of roof panels should be specified in accordance to site condition | Submit detailed drawings
5.2 | System of installation to ensure air circulation around the cold room to avoid humidity and bacteria | Specify
6 | Floor
6.1 | Flooring should be made of anti bacteria material with re-inforcement to bear the load of racking system | Yes
6.2 | Non-slippery flooring | Yes
7 | Temperature
7.1 | Internal | 0C degree to - 20C degree
7.2 | External | ≥ + 50 C degree
7.3 | Accuracy | (+ or - 1 C)
8 | Compressor | Refrigeration system which are switched back & forth periodically on cycle timer to assure both systems are given equal running time & periodically tested to be functional in event of a failure, the stand-by system is automatically activated to maintain temperature of freezer room
8.1 | Refrigerant | Non-CFC
8.2 | Type | Specify ,should be provide data sheet of compressor
8.3 | Number of main compressor | According to site and tempreture
8.4 | Number of backup compressor | equal number of main compressor
8.5 | Operation | Synchronization
8.6 | Power efficiency | Specify
8.7 | Power consumption | Specify
8.8 | Defective system can be corrected without disturbing the operation of freezer room | YES
9 | Evaporator
9.1 | Electronically controlled evaporator | YES
9.2 | Integrated evaporator system , cooling copper coil, aluminium fin radiator , heating coil & motor , aluminium fan, enclosed in metal enclosure with hinged door. universal drain fitting. | Yes
10 | Condenser
10.1 | Horizontal or vertical air discharge, steel cabinet, motor with internal overload protection | Yes
10.2 | Highly efficient condenser with automatic condensate evaporating system | Yes
10.3 | air circulation | Forced air circulation to maintain chamber uniformity
11 | Display | Digital LED/LCD , Specify
12 | Control System
12.1 | Multifunction electronic control panel system controlled by solid-state programmable | Yes
12.2 | Microprocessor | Yes
12.3 | Touch screen interface control panel mounted at operators eye level on the front of Cold room | Yes
12.4 | Control for each cooling unit | Yes
12.5 | Control Panel Enclosure | Enclose operating controls, instrumentation, functional switches and control system in single control panel temperature record chart/memory card
12.6 | Compressor On/Off, status | Yes
12.7 | Defrost setting On/Off / status | Yes
12.8 | Switch off compressor when reach set temp | Yes
12.9 | Humidification system with digital indicator | Yes
12.1 | High pressure. indicator / controls | YES
12.11 | Low pressure. indicator / controls | YES
12.12 | Oil pressure. indicator / controls | YES
12.13 | Moisture indicating sight glass | YES
12.14 | Replaceable dryer | YES
12.15 | Chemical dryer | YES
13 | Main circuit breaker overload protection, for evaporator, heater, condenser | YES
14 | Filters | Odour
15 | Drain network for defrost water | Yes, outside room with pest prove system
16 | Lighting system | Yes
16.1 | Cool white cool led light | YES
16.2 | Minimum intensity 70 foot-candles measured 40 inch (1016 mm) above floor, uniform distribution of light inside the room . | Yes
17 | Audio visual alarms for malfunction of refrigeration system power failure alarm when temp. reach to set point | Yes
18 | Electrical power | 380V, 3 phase, 60Hz
19 | Installation & training | Installation by factory trained personals. At least one week training to operator & Technical staff of maintenance department.
20 | Operating manual | Yes
21 | Service (repair) manual | Yes, With wiring diagrams & schematic,& spare part list.
22 | Drawings: full and detailed drawings for the cold room showing locations of door, distribution of racking system and total number of bodies stored | Yes
23 | Other specifications | list of projects done in this field
24 | Dimension | L20 W20 H4
Company Stamp |  |  |  | Company Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 4226210000000 | 0 | PORTABLE MORTUARY RACK |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Designed for safety and reliability | Yes
1.1 | Material | Stainless steel
2 | Stable torsion free, rectangular tube frame steel construction | Yes
3 | No. of Tier | 4
4 | Resistant to corrosion and easy to clean | Yes
5 | Body Tray made of stainless steel quality 1.4301, seamless deep-drawn, rested on skids and at the small sides provided with bow handles | Yes
5.1 | 90% of the total body trays are for normal body size | Body tray dimensions L x W x H = 2000 x 600 x 100 mm approx.
5.2 | 10% of the total body trays are for obese body size | Body tray dimensions L x W x H = 21000 x 750 x 100 mm approx.
6 | Roller tracks: Per body one roller track with 5 rolls with maintenance-free ball bearings for an easy move of the tray in and out, at the rear side with an additional stop-roll which prevents shocks against the back wall. The rolls are made of corrosion resistant material, coated with plastic material | Yes
7 | Casters | 4 casters
7.1 | size of casters | Specify
7.2 | Brake of casters | All
8 | Rack load capacity | 1000 KG approx.
9 | One-person operation | Yes
10 | Other specifications | Compatible with morgue system
Company Stamp |  |  |  | Company Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 4226210000100 | 0 | BODY LIFTER MANUAL |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mortuary transport stretcher, concealed cadaver | Yes
2 | Stable torsion free, rectangular tube frame steel construction | Yes
3 | Resistant to corrosion and easy to clean | Yes
4 | Dimensions (MM) | Approx. 2300L X 700W
5 | weight
5.1 | loaded weight | APPROX. 250 Kg.
5.2 | Maximum unloaded weight | approx. 130kg
6 | Height | Easy adjustable to hospital bed height
6.1 | Maximum Height | Up to the height position of the Rack
7 | Fork | matching with tray
7.1 | NO. Of Forks | 2
7.2 | Reversible forks for increased lifting height | Yes
8 | Castors
8.1 | 4 Castors, 2 with break, 2 swivelling
8.2 | Size of castors | Specify
9 | Operation | Hydraulic
9.1 | Hydraulic foot operated | Yes
10 | Upper part protected by disinfection proof custom fitted washable vinyl cover. | Yes complete body will be concealed
11 | Fixed mounted immersible torsion free steel frame to fit the vinyl cover | Yes
12 | Tray
12.1 | Body | Removable stainless steel body tray 304 with drainage. Seamless shaped with carrying handles.
13 | Other Specifications | Compatible to morgue system
Company Stamp |  |  |  | Company Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 4226180601400 | 0 | FREEZER MORTUARY ROOM L20 W35 H7 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 60 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | Storage of Bodies
2 | Dimension (please submit drawing based on your site visit and actual dimension status) | Yes
3 | Body structure
3.1 | Internal wall cladding and external wall of the frontal side | stainless steel
3.2 | Type of stainless steel | AISI 314, anti-bacterial, anti-corrosive, scratch resistance.
3.3 | Panel thickness | Metal sheet thickness at least 0.8 MM
3.4 | Insulation CDC free polyurethane rigid foam approx. 100mm | Yes
3.5 | Density 40 to 45 kg/ m3 foamed | Yes
3.6 | Panels fastened together by means of cam lock actuated by allen key | Yes
3.7 | Jointless corners to avoid leakage | Yes
4 | Door
4.1 | Heavy duty sliding door | Yes
4.2 | Construction similar to wall panels with insulation | Yes
4.3 | Pivot Type | Yes
4.4 | Door Handle & Latch assembly:- Automatic door - open cycle | Yes
4.5 | closure System with Internal Safety Release | Yes
4.6 | Mechanism on inside of room, preventing personnel from being locked from outside | Yes
4.7 | Door Panel Thickness | 100 MM Approx.
4.8 | Door Size | According to site and application required
4.9 | Door Lock | Yes, with failsafe feature/ sliding door
4.1 | Gasket with enough magnetic force to form positive airtight seal | vinyl type with magnetic steel core. Gasket on both side of door Adjustable rubber wiper gasket on the bottom
4.11 | Air Curtain | on the door
5 | Roof
5.1 | Installation method of roof panels should be specified in accordance to site condition | Submit detailed drawings
5.2 | System of installation to ensure air circulation around the cold room to avoid humidity and bacteria | Specify
6 | Floor
6.1 | Flooring should be made of anti bacteria material with re-inforcement to bear the load of racking system | Yes
6.2 | Non-slippery flooring | Yes
7 | Temperature
7.1 | Internal | 0C degree to - 20C degree
7.2 | External | ≥ + 50 C degree
7.3 | Accuracy | (+ or - 1 C)
8 | Compressor | Refrigeration system which are switched back & forth periodically on cycle timer to assure both systems are given equal running time & periodically tested to be functional in event of a failure, the stand-by system is automatically activated to maintain temperature of freezer room
8.1 | Refrigerant | Non-CFC
8.2 | Type | Specify ,should be provide data sheet of compressor
8.3 | Number of main compressor | According to site and tempreture
8.4 | Number of backup compressor | equal number of main compressor
8.5 | Operation | Synchronization
8.6 | Power efficiency | Specify
8.7 | Power consumption | Specify
8.8 | Defective system can be corrected without disturbing the operation of freezer room | YES
9 | Evaporator
9.1 | Electronically controlled evaporator | YES
9.2 | Integrated evaporator system , cooling copper coil, aluminium fin radiator , heating coil & motor , aluminium fan, enclosed in metal enclosure with hinged door. universal drain fitting. | Yes
10 | Condenser
10.1 | Horizontal or vertical air discharge, steel cabinet, motor with internal overload protection | Yes
10.2 | Highly efficient condenser with automatic condensate evaporating system | Yes
10.3 | air circulation | Forced air circulation to maintain chamber uniformity
11 | Display | Digital LED/LCD , Specify
12 | Control System
12.1 | Multifunction electronic control panel system controlled by solid-state programmable | Yes
12.2 | Microprocessor | Yes
12.3 | Touch screen interface control panel mounted at operators eye level on the front of Cold room | Yes
12.4 | Control for each cooling unit | Yes
12.5 | Control Panel Enclosure | Enclose operating controls, instrumentation, functional switches and control system in single control panel temperature record chart/memory card
12.6 | Compressor On/Off, status | Yes
12.7 | Defrost setting On/Off / status | Yes
12.8 | Switch off compressor when reach set temp | Yes
12.9 | Humidification system with digital indicator | Yes
12.1 | High pressure. indicator / controls | YES
12.11 | Low pressure. indicator / controls | YES
12.12 | Oil pressure. indicator / controls | YES
12.13 | Moisture indicating sight glass | YES
12.14 | Replaceable dryer | YES
12.15 | Chemical dryer | YES
13 | Main circuit breaker overload protection, for evaporator, heater, condenser | YES
14 | Filters | Odour
15 | Drain network for defrost water | Yes, outside room with pest prove system
16 | Lighting system | Yes
16.1 | Cool white cool led light | YES
16.2 | Minimum intensity 70 foot-candles measured 40 inch (1016 mm) above floor, uniform distribution of light inside the room . | Yes
17 | Audio visual alarms for malfunction of refrigeration system power failure alarm when temp. reach to set point | Yes
18 | Electrical power | 380V, 3 phase, 60Hz
19 | Installation & training | Installation by factory trained personals. At least one week training to operator & Technical staff of maintenance department.
20 | Operating manual | Yes
21 | Service (repair) manual | Yes, With wiring diagrams & schematic,& spare part list.
22 | Drawings: full and detailed drawings for the cold room showing locations of door, distribution of racking system and total number of bodies stored | Yes
23 | Other specifications | list of projects done in this field
24 | Dimension | L20 W35 H7
Company Stamp |  |  |  | Company Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 4226180601500 | 0 | FREEZER MORTUARY ROOM L28 W28 H5 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 70 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | Storage of Bodies
2 | Dimension (please submit drawing based on your site visit and actual dimension status) | Yes
3 | Body structure
3.1 | Internal wall cladding and external wall of the frontal side | stainless steel
3.2 | Type of stainless steel | AISI 314, anti-bacterial, anti-corrosive, scratch resistance.
3.3 | Panel thickness | Metal sheet thickness at least 0.8 MM
3.4 | Insulation CDC free polyurethane rigid foam approx. 100mm | Yes
3.5 | Density 40 to 45 kg/ m3 foamed | Yes
3.6 | Panels fastened together by means of cam lock actuated by allen key | Yes
3.7 | Jointless corners to avoid leakage | Yes
4 | Door
4.1 | Heavy duty sliding door | Yes
4.2 | Construction similar to wall panels with insulation | Yes
4.3 | Pivot Type | Yes
4.4 | Door Handle & Latch assembly:- Automatic door - open cycle | Yes
4.5 | closure System with Internal Safety Release | Yes
4.6 | Mechanism on inside of room, preventing personnel from being locked from outside | Yes
4.7 | Door Panel Thickness | 100 MM Approx.
4.8 | Door Size | According to site and application required
4.9 | Door Lock | Yes, with failsafe feature/ sliding door
4.1 | Gasket with enough magnetic force to form positive airtight seal | vinyl type with magnetic steel core. Gasket on both side of door Adjustable rubber wiper gasket on the bottom
4.11 | Air Curtain | on the door
5 | Roof
5.1 | Installation method of roof panels should be specified in accordance to site condition | Submit detailed drawings
5.2 | System of installation to ensure air circulation around the cold room to avoid humidity and bacteria | Specify
6 | Floor
6.1 | Flooring should be made of anti bacteria material with re-inforcement to bear the load of racking system | Yes
6.2 | Non-slippery flooring | Yes
7 | Temperature
7.1 | Internal | 0C degree to - 20C degree
7.2 | External | ≥ + 50 C degree
7.3 | Accuracy | (+ or - 1 C)
8 | Compressor | Refrigeration system which are switched back & forth periodically on cycle timer to assure both systems are given equal running time & periodically tested to be functional in event of a failure, the stand-by system is automatically activated to maintain temperature of freezer room
8.1 | Refrigerant | Non-CFC
8.2 | Type | Specify ,should be provide data sheet of compressor
8.3 | Number of main compressor | According to site and tempreture
8.4 | Number of backup compressor | equal number of main compressor
8.5 | Operation | Synchronization
8.6 | Power efficiency | Specify
8.7 | Power consumption | Specify
8.8 | Defective system can be corrected without disturbing the operation of freezer room | YES
9 | Evaporator
9.1 | Electronically controlled evaporator | YES
9.2 | Integrated evaporator system , cooling copper coil, aluminium fin radiator , heating coil & motor , aluminium fan, enclosed in metal enclosure with hinged door. universal drain fitting. | Yes
10 | Condenser
10.1 | Horizontal or vertical air discharge, steel cabinet, motor with internal overload protection | Yes
10.2 | Highly efficient condenser with automatic condensate evaporating system | Yes
10.3 | air circulation | Forced air circulation to maintain chamber uniformity
11 | Display | Digital LED/LCD , Specify
12 | Control System
12.1 | Multifunction electronic control panel system controlled by solid-state programmable | Yes
12.2 | Microprocessor | Yes
12.3 | Touch screen interface control panel mounted at operators eye level on the front of Cold room | Yes
12.4 | Control for each cooling unit | Yes
12.5 | Control Panel Enclosure | Enclose operating controls, instrumentation, functional switches and control system in single control panel temperature record chart/memory card
12.6 | Compressor On/Off, status | Yes
12.7 | Defrost setting On/Off / status | Yes
12.8 | Switch off compressor when reach set temp | Yes
12.9 | Humidification system with digital indicator | Yes
12.1 | High pressure. indicator / controls | YES
12.11 | Low pressure. indicator / controls | YES
12.12 | Oil pressure. indicator / controls | YES
12.13 | Moisture indicating sight glass | YES
12.14 | Replaceable dryer | YES
12.15 | Chemical dryer | YES
13 | Main circuit breaker overload protection, for evaporator, heater, condenser | YES
14 | Filters | Odour
15 | Drain network for defrost water | Yes, outside room with pest prove system
16 | Lighting system | Yes
16.1 | Cool white cool led light | YES
16.2 | Minimum intensity 70 foot-candles measured 40 inch (1016 mm) above floor, uniform distribution of light inside the room . | Yes
17 | Audio visual alarms for malfunction of refrigeration system power failure alarm when temp. reach to set point | Yes
18 | Electrical power | 380V, 3 phase, 60Hz
19 | Installation & training | Installation by factory trained personals. At least one week training to operator & Technical staff of maintenance department.
20 | Operating manual | Yes
21 | Service (repair) manual | Yes, With wiring diagrams & schematic,& spare part list.
22 | Drawings: full and detailed drawings for the cold room showing locations of door, distribution of racking system and total number of bodies stored | Yes
23 | Other specifications | list of projects done in this field
24 | Dimension | L28 W28 H5
Company Stamp |  |  |  | Company Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 4226180601600 | 0 | FREEZER MORTUARY ROOM L20 W18 H4 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 80 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | Storage of Bodies
2 | Dimension (please submit drawing based on your site visit and actual dimension status) | Yes
3 | Body structure
3.1 | Internal wall cladding and external wall of the frontal side | stainless steel
3.2 | Type of stainless steel | AISI 314, anti-bacterial, anti-corrosive, scratch resistance.
3.3 | Panel thickness | Metal sheet thickness at least 0.8 MM
3.4 | Insulation CDC free polyurethane rigid foam approx. 100mm | Yes
3.5 | Density 40 to 45 kg/ m3 foamed | Yes
3.6 | Panels fastened together by means of cam lock actuated by allen key | Yes
3.7 | Jointless corners to avoid leakage | Yes
4 | Door
4.1 | Heavy duty sliding door | Yes
4.2 | Construction similar to wall panels with insulation | Yes
4.3 | Pivot Type | Yes
4.4 | Door Handle & Latch assembly:- Automatic door - open cycle | Yes
4.5 | closure System with Internal Safety Release | Yes
4.6 | Mechanism on inside of room, preventing personnel from being locked from outside | Yes
4.7 | Door Panel Thickness | 100 MM Approx.
4.8 | Door Size | According to site and application required
4.9 | Door Lock | Yes, with failsafe feature/ sliding door
4.1 | Gasket with enough magnetic force to form positive airtight seal | vinyl type with magnetic steel core. Gasket on both side of door Adjustable rubber wiper gasket on the bottom
4.11 | Air Curtain | on the door
5 | Roof
5.1 | Installation method of roof panels should be specified in accordance to site condition | Submit detailed drawings
5.2 | System of installation to ensure air circulation around the cold room to avoid humidity and bacteria | Specify
6 | Floor
6.1 | Flooring should be made of anti bacteria material with re-inforcement to bear the load of racking system | Yes
6.2 | Non-slippery flooring | Yes
7 | Temperature
7.1 | Internal | 0C degree to - 20C degree
7.2 | External | ≥ + 50 C degree
7.3 | Accuracy | (+ or - 1 C)
8 | Compressor | Refrigeration system which are switched back & forth periodically on cycle timer to assure both systems are given equal running time & periodically tested to be functional in event of a failure, the stand-by system is automatically activated to maintain temperature of freezer room
8.1 | Refrigerant | Non-CFC
8.2 | Type | Specify ,should be provide data sheet of compressor
8.3 | Number of main compressor | According to site and tempreture
8.4 | Number of backup compressor | equal number of main compressor
8.5 | Operation | Synchronization
8.6 | Power efficiency | Specify
8.7 | Power consumption | Specify
8.8 | Defective system can be corrected without disturbing the operation of freezer room | YES
9 | Evaporator
9.1 | Electronically controlled evaporator | YES
9.2 | Integrated evaporator system , cooling copper coil, aluminium fin radiator , heating coil & motor , aluminium fan, enclosed in metal enclosure with hinged door. universal drain fitting. | Yes
10 | Condenser
10.1 | Horizontal or vertical air discharge, steel cabinet, motor with internal overload protection | Yes
10.2 | Highly efficient condenser with automatic condensate evaporating system | Yes
10.3 | air circulation | Forced air circulation to maintain chamber uniformity
11 | Display | Digital LED/LCD , Specify
12 | Control System
12.1 | Multifunction electronic control panel system controlled by solid-state programmable | Yes
12.2 | Microprocessor | Yes
12.3 | Touch screen interface control panel mounted at operators eye level on the front of Cold room | Yes
12.4 | Control for each cooling unit | Yes
12.5 | Control Panel Enclosure | Enclose operating controls, instrumentation, functional switches and control system in single control panel temperature record chart/memory card
12.6 | Compressor On/Off, status | Yes
12.7 | Defrost setting On/Off / status | Yes
12.8 | Switch off compressor when reach set temp | Yes
12.9 | Humidification system with digital indicator | Yes
12.1 | High pressure. indicator / controls | YES
12.11 | Low pressure. indicator / controls | YES
12.12 | Oil pressure. indicator / controls | YES
12.13 | Moisture indicating sight glass | YES
12.14 | Replaceable dryer | YES
12.15 | Chemical dryer | YES
13 | Main circuit breaker overload protection, for evaporator, heater, condenser | YES
14 | Filters | Odour
15 | Drain network for defrost water | Yes, outside room with pest prove system
16 | Lighting system | Yes
16.1 | Cool white cool led light | YES
16.2 | Minimum intensity 70 foot-candles measured 40 inch (1016 mm) above floor, uniform distribution of light inside the room . | Yes
17 | Audio visual alarms for malfunction of refrigeration system power failure alarm when temp. reach to set point | Yes
18 | Electrical power | 380V, 3 phase, 60Hz
19 | Installation & training | Installation by factory trained personals. At least one week training to operator & Technical staff of maintenance department.
20 | Operating manual | Yes
21 | Service (repair) manual | Yes, With wiring diagrams & schematic,& spare part list.
22 | Drawings: full and detailed drawings for the cold room showing locations of door, distribution of racking system and total number of bodies stored | Yes
23 | Other specifications | list of projects done in this field
24 | Dimension | L20 W18 H4
Company Stamp |  |  |  | Company Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 4226180601700 | 0 | FREEZER MORTUARY ROOM L20 W24 H4 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 80 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application | Storage of Bodies
2 | Dimension (please submit drawing based on your site visit and actual dimension status) | Yes
3 | Body structure
3.1 | Internal wall cladding and external wall of the frontal side | stainless steel
3.2 | Type of stainless steel | AISI 314, anti-bacterial, anti-corrosive, scratch resistance.
3.3 | Panel thickness | Metal sheet thickness at least 0.8 MM
3.4 | Insulation CDC free polyurethane rigid foam approx. 100mm | Yes
3.5 | Density 40 to 45 kg/ m3 foamed | Yes
3.6 | Panels fastened together by means of cam lock actuated by allen key | Yes
3.7 | Jointless corners to avoid leakage | Yes
4 | Door
4.1 | Heavy duty sliding door | Yes
4.2 | Construction similar to wall panels with insulation | Yes
4.3 | Pivot Type | Yes
4.4 | Door Handle & Latch assembly:- Automatic door - open cycle | Yes
4.5 | closure System with Internal Safety Release | Yes
4.6 | Mechanism on inside of room, preventing personnel from being locked from outside | Yes
4.7 | Door Panel Thickness | 100 MM Approx.
4.8 | Door Size | According to site and application required
4.9 | Door Lock | Yes, with failsafe feature/ sliding door
4.1 | Gasket with enough magnetic force to form positive airtight seal | vinyl type with magnetic steel core. Gasket on both side of door Adjustable rubber wiper gasket on the bottom
4.11 | Air Curtain | on the door
5 | Roof
5.1 | Installation method of roof panels should be specified in accordance to site condition | Submit detailed drawings
5.2 | System of installation to ensure air circulation around the cold room to avoid humidity and bacteria | Specify
6 | Floor
6.1 | Flooring should be made of anti bacteria material with re-inforcement to bear the load of racking system | Yes
6.2 | Non-slippery flooring | Yes
7 | Temperature
7.1 | Internal | 0C degree to - 20C degree
7.2 | External | ≥ + 50 C degree
7.3 | Accuracy | (+ or - 1 C)
8 | Compressor | Refrigeration system which are switched back & forth periodically on cycle timer to assure both systems are given equal running time & periodically tested to be functional in event of a failure, the stand-by system is automatically activated to maintain temperature of freezer room
8.1 | Refrigerant | Non-CFC
8.2 | Type | Specify ,should be provide data sheet of compressor
8.3 | Number of main compressor | According to site and tempreture
8.4 | Number of backup compressor | equal number of main compressor
8.5 | Operation | Synchronization
8.6 | Power efficiency | Specify
8.7 | Power consumption | Specify
8.8 | Defective system can be corrected without disturbing the operation of freezer room | YES
9 | Evaporator
9.1 | Electronically controlled evaporator | YES
9.2 | Integrated evaporator system , cooling copper coil, aluminium fin radiator , heating coil & motor , aluminium fan, enclosed in metal enclosure with hinged door. universal drain fitting. | Yes
10 | Condenser
10.1 | Horizontal or vertical air discharge, steel cabinet, motor with internal overload protection | Yes
10.2 | Highly efficient condenser with automatic condensate evaporating system | Yes
10.3 | air circulation | Forced air circulation to maintain chamber uniformity
11 | Display | Digital LED/LCD , Specify
12 | Control System
12.1 | Multifunction electronic control panel system controlled by solid-state programmable | Yes
12.2 | Microprocessor | Yes
12.3 | Touch screen interface control panel mounted at operators eye level on the front of Cold room | Yes
12.4 | Control for each cooling unit | Yes
12.5 | Control Panel Enclosure | Enclose operating controls, instrumentation, functional switches and control system in single control panel temperature record chart/memory card
12.6 | Compressor On/Off, status | Yes
12.7 | Defrost setting On/Off / status | Yes
12.8 | Switch off compressor when reach set temp | Yes
12.9 | Humidification system with digital indicator | Yes
12.1 | High pressure. indicator / controls | YES
12.11 | Low pressure. indicator / controls | YES
12.12 | Oil pressure. indicator / controls | YES
12.13 | Moisture indicating sight glass | YES
12.14 | Replaceable dryer | YES
12.15 | Chemical dryer | YES
13 | Main circuit breaker overload protection, for evaporator, heater, condenser | YES
14 | Filters | Odour
15 | Drain network for defrost water | Yes, outside room with pest prove system
16 | Lighting system | Yes
16.1 | Cool white cool led light | YES
16.2 | Minimum intensity 70 foot-candles measured 40 inch (1016 mm) above floor, uniform distribution of light inside the room . | Yes
17 | Audio visual alarms for malfunction of refrigeration system power failure alarm when temp. reach to set point | Yes
18 | Electrical power | 380V, 3 phase, 60Hz
19 | Installation & training | Installation by factory trained personals. At least one week training to operator & Technical staff of maintenance department.
20 | Operating manual | Yes
21 | Service (repair) manual | Yes, With wiring diagrams & schematic,& spare part list.
22 | Drawings: full and detailed drawings for the cold room showing locations of door, distribution of racking system and total number of bodies stored | Yes
23 | Other specifications | list of projects done in this field
24 | Dimension | L20 W24 H4
Company Stamp |  |  |  | Company Signature