## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder. |  |  | RIYADH | JEDDAH | DAMMAM | Total
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | Medical Fridge +2 + 8 degrees with digital temperature and recording system 500 Liters | 12 | 10 | 9 | 31
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | USB Temperature Data logger -80 degree celcius re-usable | 8 | 6 | 4 | 18
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | USB Temperature Data logger +2 +8 degree Celcius re-usable | 12 | 12 | 11 | 35
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | Rotronic remote record temperature monitoring +2+8 (control tower) | 12 | 12 | 11 | 35
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | Rotronic remote record temperature monitoring -80 (Control tower) | 8 | 6 | 4 | 18

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 4455170010024 | Medical Fridge +2 + 8 degrees with digital temperature and recording system 500 Liters |  | MEDICAL GROUP | 31
2 | 4455330010141 | USB Temperature Data logger -80 degree celcius re-usable |  | MEDICAL GROUP | 18
3 | 4455150010390 | USB Temperature Data logger +2 +8 degree Celcius re-usable |  | MEDICAL GROUP | 35
4 | 4455330010059 | Rotronic remote record temperature monitoring +2+8 (control tower) |  | MEDICAL GROUP | 35
5 | 4455150010058 | Rotronic remote record temperature monitoring -80 (Control tower) |  | MEDICAL GROUP | 18
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | MEDICAL GROUP | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
|  | NON MEDICAL | 0 | 0 | [5 empty] | If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1625 or send email to Direct-Purchase@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 4455170010024 | 0 | Medical Fridge +2 + 8 degrees with digital temperature and recording system 500 Liters |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Floor standing
2 | CONFIGURATION | Storage of pharmaceuticals under safe, controlled conditions
3 | CAPACITY IN LITERS | 500 L approximatlly
4 | SHELVES | 3 or 4 heavy gauge zinc plated /powder coated
5 | DOORS | 1 or 2
6 | DOOR TYPE | Solid door
7 | DOOR SEALING SYSTEM | Yes, para-magnetic
8 | DOOR LOCK | Yes
9 | TEMPERATURE CONTROL | 2 degree Centigrate TO 8 degree Centigrate
10 | INTERIOR | Stainless steel
11 | EXTERIOR | Powder coated enemal paint
12 | TEMP-ALARM . And Door Open Alarm | Yes, High & Low / and Yes for Door open Alarm
13 | TEMPERATURE INDICATOR | Yes, Digital display
14 | TEMP. CHART RECORDER | Yes
15 | POWER FAILURE ALARM, AUDIO & VISUAL | Yes
16 | RAPID TEMP. RECOVERY AFTER DOOR OPENINGS | Yes
17 | INSULATION | CFC free
18 | INSIDE LIGHT | Yes
19 | CASTERS | Yes 4, (two lockable)
20 | UPS | Optional (suitable power)
21 | POWER SUPPLY | 220 V, 60 Hz
22 | OTHER SPECIFICATION | FDA Or CE
23 | Printout for the item offered with all technical parameters exactly matches the soft copy in the CD. | YES
24 | Supplier is an authorized representative of the manufacturer by SFDA and/or an official agent by Ministry of Trade. | Yes (Attach official documents)
25 | Attach MDMA Certificate | Yes (Give the MDMA Number)
Company Stamp |  |  |  | Company Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 4455330010141 | 0 | USB Temperature Data logger -80 degree celcius re-usable |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To monitor temperature & humidity in environmental Chambers, Chemical Storage area, Clean Rooms, Incubators, Refrigerators, Freezers & Store Room
2 | Measurement range for Temperature probe of ultra freezers -80°C | -90 °C..+260 °C
3 | Measurement range of lab wall sensors | 0 to 100%RH, -40 °C to 60°C
4 | Temperatures Calibrated Accuracy | Dry block calibrator: ±0.5 °C for core range:-30 °C..+30 °C Humidity generator: ± 1.0 °C for expanded range: -100 °C..-31 °C and +31 °C..+150 °C
5 | Transmitter Calibrated Accuracy | Temperature Calibration: Humidity generator: ± 1.0 °C for expanded range: -100 °C..-31 °C and +31 °C..+150 °C Humidity Calibration: As Found ± 5.0% RH
6 | Memory of the Measurement module | Internal non-volatile memory 10,000 measurements
7 | Logging interval | in the system monitor its configurable from 1 minute to 60 minutes
8 | Number of channels of the Measurement module | 4
9 | REUSABLE | YES
10 | USB PORT | YES
11 | LED Indicators in the Communication Module | Indicators to monitor: • The module Status(Activity) • Power Supply Status • Ethernet Status
12 | LED Indicators in the Measurement Module | Indicators to monitor: • The module Status(Activity). • Power Supply Status. • Sensors configuration Status.
13 | Material of the modules | ABS: Acrylonitrile Butadiene Styrene
14 | The environmental Ingress Protection (IP) rating | IP20
15 | Operating Environment | 0 °C to 40 °C, 10 %RH to 90 %RH, non-condensing
16 | Safety standard | UL, IEC, CSA or other international Standard, (Specify in offer).
17 | REGULATORY COMPLIANCES | - The System must be compliant with the FDA chapter 21 part 11 regulatory requirements - Subjected to CE marking - The manufacturer guarantees that this product complies with the relevant guidelines: EN61000-6-2 : 2001 and EN 61000-6-4 : 2001 - FCC ID: Z45-E11645398 - IC: 9954A-E11645398
Company Stamp |  |  |  | Company Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 4455150010390 | 0 | USB Temperature Data logger +2 +8 degree Celcius re-usable |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | To monitor temperature & humidity in environmental Chambers, Chemical Storage area, Clean Rooms, Incubators, Refrigerators, Freezers & Store Room
2 | Measurement range for Temperature probe of refrigerators 2-8°C and freezers -20°C | -50 °C to +105 °C
3 | Temperatures Calibrated Accuracy | Dry block calibrator: ±0.5 °C for core range:-30 °C..+30 °C Humidity generator: ± 1.0 °C for expanded range: -100 °C..-31 °C and +31 °C..+150 °C
4 | Transmitter Calibrated Accuracy | Temperature Calibration: Humidity generator: ± 1.0 °C for expanded range: -100 °C..-31 °C and +31 °C..+150 °C Humidity Calibration: As Found ± 5.0% RH
5 | Memory of the Measurement module | Internal non-volatile memory 10,000 measurements
6 | Logging interval | in the system monitor its configurable from 1 minute to 60 minutes
7 | Number of channels of the Measurement module | 4
8 | REUSABLE | YES
9 | USB PORT | YES
10 | LED Indicators in the Communication Module | Indicators to monitor: • The module Status(Activity) • Power Supply Status • Ethernet Status
11 | LED Indicators in the Measurement Module | Indicators to monitor: • The module Status(Activity). • Power Supply Status. • Sensors configuration Status.
12 | Material of the modules | ABS: Acrylonitrile Butadiene Styrene
13 | The environmental Ingress Protection (IP) rating | IP20
14 | Operating Environment | 0 °C to 40 °C, 10 %RH to 90 %RH, non-condensing
15 | Safety standard | UL, IEC, CSA or other international Standard, (Specify in offer).
16 | REGULATORY COMPLIANCES | - The System must be compliant with the FDA chapter 21 part 11 regulatory requirements - Subjected to CE marking - The manufacturer guarantees that this product complies with the relevant guidelines: EN61000-6-2 : 2001 and EN 61000-6-4 : 2001 - FCC ID: Z45-E11645398 - IC: 9954A-E11645398
Company Stamp |  |  |  | Company Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 4455330010059 | 0 | Rotronic remote record temperature monitoring +2+8 (control tower) |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Lab use for temperature measurement
2 | RANGE | (-20 to +15) degrees
3 | DISPLAY | Visual
4 | UNIT DISPLAYED | Centigrade
5 | DIVISIONS | 1 Degree
6 | IMMERSION DEPTH | Partial/Total
7 | TOTAL LENGTH | Not longer than 12 inches / (76mm)
8 | FINE PORE CAPILLARY IN STEM TO PERMIT VARTING IMMERSION DEPTH | Yes
9 | PERMANENTLY FUSED MARKINGS FOR LONG-TERM USE | Yes
10 | SAFETY | Non toxic, Blue biodegradable liquid (Isoamyle benzoate) filling
11 | FIXATION | Non-roll top, suspension ring
12 | ACCESSORIES | Yes, All Standard Acc.
13 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Company Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 4455150010058 | 0 | Rotronic remote record temperature monitoring -80 (Control tower) |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 60 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATION | Lab use for temperature measurement
2 | RANGE | (-100 to +50) degrees
3 | DISPLAY | Visual
4 | UNIT DISPLAYED | Centigrade
5 | DIVISIONS | 1 Degree
6 | IMMERSION DEPTH | Partial/Total
7 | TOTAL LENGTH | Not longer than 12 inches / (76mm)
8 | FINE PORE CAPILLARY IN STEM TO PERMIT VARTING IMMERSION DEPTH | Yes
9 | PERMANENTLY FUSED MARKINGS FOR LONG-TERM USE | Yes
10 | SAFETY | Non toxic, Blue biodegradable liquid (Isoamyle benzoate) filling
11 | FIXATION | Non-roll top, suspension ring
12 | ACCESSORIES | Yes, All Standard Acc.
13 | OTHER SPECIFICATIONS
Company Stamp |  |  |  | Company Signature