## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | SUITE ANGIOGRAPHY
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | FLUOROSCOPY MACHINE
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | AUTOMATED FLOW CYTOMETER ANALYZER 100–240 V
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | VENTILATOR TRANSPORT
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | ANESTHESIA WORKSTATION MRI FABIUS
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 5.5X65
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 4X65
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | ICU VENTILATOR MRI COMPATIBLE
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | PORTABLE VENTILATOR WITH TURBINE
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | VENTILATOR NEONATAL BASE UNIT SERVO-N WITH HFOV
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | AMBULATORY HOLTER MONITORING SYSTEM WITH 10 RECORDERS
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | TRACTION UNIT PHYSICAL THERAPY
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | MONITOR PATIENT-WORN M300 TELEMETRY DEVICE
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | UV DISINFECTOR FOR ULTRASOUND PROBES WITH STORAGE CABINET
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | ANALYZER SYSTEM FOR VOICE AND SPEECH CLINICAL ASSESSMENT AND BIOFEEDBACK
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | OPHTHALMOSCOPE / OTOSCOPE SET
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | EXAMINATION TABLE POWERED
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | PATIENT MONITOR VITAL SIGNS DIAGNOSTIC PHYSICAL EXAM SETS, CONNEX INTEGRATED WALL SYSTEM, WELCH ALLYN, MODEL 85NTVXC-4
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | MONITOR PATIENT PHYSIOLOGICAL FOR ANAESTHESIA
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | UNIVERSAL CHARGER OTO/OPHT SCOPE COMPLETE W/TWO 3.5V RECHARGEABLE HANDLES W/O HEADS
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | CHAIR DIALYSIS
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | BED PATIENT REGULAR
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | SPHYGMOMANOMETER HAND ANEROID W/ADULT CUFF
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | THERMOMETER DIGITAL FOR CLINICAL USE
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | LIGHT PROCEDURE PORTABLE F300 ON STAND WITH ARTICULATING ARM
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | SCALE WEIGHT DIGITAL ADULT SCALETRONIX 5002 STAND-ON
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | WARMING UNIT HYPOTHERMIA
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | DEVICE CONTINUOUS MONITORING SYSTEM CARDIAC OUTPUT
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | MECHANICAL CIRCULATORY SUPPORT DEVICE CARDIOHELP with component and acessories
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | ROTAFLOW CONSOLE WITH ICU KIT RFC 230 V / 50 HZ with component and acessories
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | MACHINE ANESTHESIA ADVANCED VENTILATION WITH MONITORING SYSTEM
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | MONITOR BODY COMPOSITION (BCM)
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | LAMP HEATING
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | AIRWAY CLEARANCE LUNG EXPANSION SYSTEM THREE IN ONE THERAPY
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | Endoscopic Video System For Difficult Intubation

## Summary
SN | CODE | ITEM DESCRIPTION | Requisition No. | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR) (including vat if applicable) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | Delivery 1 | Delivery 2
1 | 411556300 | ANGIOGRAPHY BIPLANE | 1123278 | Medical Eqpt | 1 | [10 empty] | 0
2 | 406320103 | FLUOROSCOPY MACHINE | 1123278 | Medical Eqpt | 2 | [10 empty] | 0
3 | 417857201 | FLOW CYTOMETER | 1133623 | Medical Eqpt | 1 | [10 empty] | 0
4 | 92007 | VENTILATOR TRANSPORT | 1154753 | Medical Eqpt | 2 | [10 empty] | 0
5 | 414268800 | ANESTHESIA MACHINE MRI COMPATIBLE | 1155685 | Medical Eqpt | 2 | [10 empty] | 0
6 | 411633102 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 5.5X65 | 1156343 | Medical Eqpt | 1 | [10 empty] | 0
7 | 411633101 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 4X65 | 1156343 | Medical Eqpt | 1 | [10 empty] | 0
8 | 411605301 | ICU VENTILATOR MRI COMPATIBLE | 1157157 | Medical Eqpt | 1 | [10 empty] | 0
9 | 411605203 | PORTABLE VENTILATOR WITH TURBINE | 1157159 | Medical Eqpt | 4 | [10 empty] | 0
10 | 411605311 | VENTILATOR ICU NEONATAL with HFOV | 1157796 | Medical Eqpt | 3 | [10 empty] | 0
11 | 411531400 | AMBULATORY HOLTER MONITORING SYSTEM WITH 10 RECORDERS | 1160697 | Medical Eqpt | 1 | [10 empty] | 0
12 | 105139 | TRACTION UNIT PHYSICAL THERAPY | 1163670 | Medical Eqpt | 2 | [10 empty] | 0
13 | 414348904 | TELEMETRY SYSTEM | 1168012 | Medical Eqpt | 2 | [10 empty] | 0
14 | 1160687 | UV DISINFECTOR FOR ULTRASOUND PROBES WITH STORAGE CABINET | 1160687 | Medical Eqpt | 1 | [10 empty] | 0
15 | 417666001 | ANALYZER SYSTEM FOR VOICE AND SPEECH CLINICAL ASSESSMENT AND BIOFEEDBACK | 1168762 | Medical Eqpt | 2 | [10 empty] | 0
16 | 414867100 | OPHTHALMOSCOPE / OTOSCOPE SET | 1191576 | Medical Eqpt | 1 | [10 empty] | 0
17 | 102508 | EXAMINATION TABLE POWERED | 1191576 | Medical Eqpt | 3 | [10 empty] | 0
18 | 442181901 | ELECTRONIC VITAL SIGN monitor with builtin OPHTHALMOSCOPE / OTOSCOPE SET | 1193204 | Medical Eqpt | 2 | [10 empty] | 0
19 | 411532308 | MONITOR PATIENT PHYSIOLOGICAL FOR ANAESTHESIA | 1193204 | Medical Eqpt | 1 | [10 empty] | 0
20 | 314867101 | UNIVERSAL CHARGER OTO/OPHT SCOPE COMPLETE W/TWO 3.5V RECHARGEABLE HANDLES W/O HEADS | 1193204 | Medical Eqpt | 4 | [10 empty] | 0
21 | 112471 | CHAIR DIALYSIS | 1193204 | Medical Eqpt | 1 | [10 empty] | 0
22 | 112466 | BED PATIENT REGULAR | 1193204 | Medical Eqpt | 3 | [10 empty] | 0
23 | 108263 | SPHYGMOMANOMETER ANEROID MOBILE | 1193204 | Medical Eqpt | 2 | [10 empty] | 0
24 | 33915 | THERMOMETER DIGITAL WALL MOUNT | 1193204 | Medical Eqpt | 3 | [10 empty] | 0
25 | 102142 | LIGHT EXAM MOBILE | 1193204 | Medical Eqpt | 2 | [10 empty] | 0
26 | 92029 | SCALE PATIENT WITH HEIGHT | 1193204, 1191576 | Medical Eqpt | 4 | [10 empty] | 0
27 | 347866100 | BLANKET WARMING SYSTEM AIR | 1193598 | Medical Eqpt | 5 | [10 empty] | 0
28 | 411530210 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE | 1193758 | Medical Eqpt | 5 | [10 empty] | 0
29 | 410434502 | Advanced and Portable ECMO Machine | 1194260 | Medical Eqpt | 1 | [10 empty] | 0
30 | 310434534 | ECMO MACHINE | 1194271 | Medical Eqpt | 1 | [10 empty] | 0
31 | 414268810 | MACHINE ANESTHESIA ADVANCED VENTILATION WITH MONITORING SYSTEM | 1194570 | Medical Eqpt | 1 | [10 empty] | 0
32 | 452169101 | MONITOR BODY COMPOSITION (BCM) | 1194845 | Medical Eqpt | 2 | [10 empty] | 0
33 | 100611 | LAMP HEATING | 1200563 | Medical Eqpt | 5 | [10 empty] | 0
34 | 442271600 | AIRWAY CLEARANCE LUNG EXPANSION SYSTEM THREE IN ONE THERAPY | 1210914 | Medical Eqpt | 3 | [10 empty] | 0
35 | 1157162 | Endoscopic Video System For Difficult Intubation | 1157162 | Medical Eqpt | 1 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 411556300 | 1123278 | ANGIOGRAPHY BIPLANE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Manufaturer | Specify
2 | Model Name | Specify
3 | Configuration
4 | Design type | Biplane
5 | FDA Approval | Yes
6 | FDA Approval Year | Specify
7 | Generator
8 | The generator must have the following characteristics and capabilities: Output waveform = 3-phase, 12-pulse, high frequency or constant potential. | Yes
9 | The generator should be 100 kW units nominally rated as follows: a) 1,000 mA at 80 KVp b) 800 mA at 100 KVp | Yes
10 | The generator should have radiographic range of 40 to 125 KVp or better | Yes
11 | Minimum Switching Time(ms) | 2 or less
12 | Shortest exposure time should be 1 msec or less | Yes
13 | Maximum continuoes power | Please specify (best preferred)
14 | The generators must include fully integrated controls for the following functions: Digital fluoroscopy Digital pulsed fluoroscopy Fluoroscopic last image hold | Yes
15 | The generator must incorporate fluoroscopic automatic brightness/dose control, varying mA as the primary control and kV as the secondary control. The generator must also include an integrating fluoroscopic timer to record fluoroscopic time from 0.1 to 99.9 minutes. | Yes
16 | System must display all technique parameters, system status, percentage of tube heat load, actual radiographic exposure time | Yes
17 | Digital Radiography (DR) Mode
18 | kVp range | 40-125 or better
19 | mA range | up to 1,000 or better
20 | AEC | yes
21 | Exposure time range (sec) | 0.001 - 1.5 or better
22 | Anatomical programming | yes
23 | Continuous Fluoroscopy Mode
24 | kVp range | 40-125
25 | Power for 10 minute fluoro | 4.5-5 KW
26 | Power for 20 minute fluoro | 4-4.5 KW
27 | mA range | up to 10 or better
28 | ABS control | yes (kVp and mA)
29 | Maximum exposure rate (R/min) | 10 or better
30 | Pulsed Fluoroscopy Mode | yes
31 | Pulses per sec | up to 30 or better
32 | X-ray Tube (Frontal)
33 | Tube Power Rating (kW) | 100 kW @ 100 kVp or greater
34 | Nominal kVp | 125 or greater
35 | Anode Heat Storage Capacity (MHU) | 6 or greater(higher preferred)
36 | Anode heat cooling Rate (MHU/min) | 1.7 or higher (best preferred)
37 | Power for 10 minute fluoro | 4.5-5 KW
38 | Power for 20 minute fluoro | 4-4.5 KW
39 | Max heat storage assembly(MHU) | 7 or greater( higher preferred)
40 | Thermal Safety switch | Yes
41 | Pre filters (0.1 to 1) mm,copper equivqlent | Yes (please specify)
42 | target angle | Please specify
43 | Maximum Field width on SID =120 | Please specify (best preferred)
44 | Maximum Field width on SID =100 | Please specify (best preferred)
45 | Maximum Field width on SID =70 | please specify (best preferred)
46 | cooling method | Please specify (best preferred)
47 | Grid Controlled | yes
48 | Focal Spot | Minimum 2 Focal Spots sizes
49 | Micro (mm) | Please specify.
50 | Small(mm) | Please specify
51 | Large(mm) | Please specify
52 | X-ray Tube (Lateral)
53 | Tube Power Rating (kW) | 100 kW @ 100 kVp or greater
54 | Nominal kVp | 125 or greater
55 | Anode Heat Storage Capacity (MHU) | 6 or greater(higher preferred)
56 | Anode heat cooling Rate (MHU/min) | 1.7 or higher (best preferred)
57 | Power for 10 minute fluoro | 4.5-5 KW
58 | Power for 20 minute fluoro | 4-4.5 KW
59 | Max heat storage assembly(MHU) | 7.5 or greater higher preferred
60 | Thermal Safety switch | Yes
61 | Pre filters (0.1 to 1) mm,copper equivqlent | Yes (please specify)
62 | Target angle | Please specify
63 | cooling method | Please specify (best preferred)
64 | Maximum Field width on SID =120 | Please specify (best preferred)
65 | Maximum Field width on SID =100 | Please specify (best preferred)
66 | Maximum Field width on SID =70 | please specify (best preferred)
67 | Grid Controlled | yes
68 | Focal Spot | Minimum 2 Focal Spots sizes
69 | Micro (mm) | specify.
70 | Small(mm) | Please specify
71 | Large(mm) | Please Specify
72 | Flat Panel X-ray Detector (Frontal)
73 | (FOV) diagnol | ≥48 cm
74 | Swivel or rotatable | yes
75 | Resolution | 2,500 x 2,000 or better
76 | Pixel Size (µm) | 160 µm or lower(best preferred)
77 | Semiconductor material | a-si
78 | Scintillator | csI
79 | Acquisition depth (pixel depth) | 16 bit
80 | Maximum acquisition speed (im/s) | up to 30
81 | Maximum diagonal size (cm) including housing. | Please Specify
82 | Input fields (cm) | Please specify
83 | Digital acquisition zoom fields(cm) | 42, 37, 31, 27, 23, 19, 16 Diagonal formats
84 | DQE at 0lp/mm | 77% or better (Please specify resolution (lp/mm) at 10%, 20% & 50%,70%)
85 | MTF at 1 lP/mm | 60% or better( Please specify resolution (lp/mm) at 10%, 20% & 50%)
86 | Flat Panel X-ray Detector (Lateral)
87 | Size (cm) | Please specify
88 | Swivel or rotatable | Yes
89 | Matrix | 1.5Kx1.5K or better(best preferred)
90 | Pixel Size (µm) | 200µm or lower(best preferred)
91 | Detector Bit | 16 Bit
92 | Resolution | 2.7 lp/mm or better
93 | DQE at 0lp/mm | 68 or better (Please specify resolution (lp/mm) at 10%, 20% & 40%,60%)
94 | MTF at 1lp/mm | 60 or better( Please specify resolution (lp/mm) at 10%, 20% & 50%)
95 | Image Acquisition
96 | Digital archiving capabilities and formats including number of images at 1K,2K matrix | Please Specify
97 | Matrix Size | 2,048 x 2,048 or better
98 | Resolution (pixel depth) | 12 bit or better
99 | Fluoroscopy pulse rates | 0.5, 1, 2, 3, 4, 6, 7.5, 10, 15, 30 p/s
100 | Frame Rates (fps) | 0.5-7.5 f/s. 7.5, 10, 15, and 30 f/s with cardiac option; 60 f/s with pediatric option
101 | Continuous Fluoroscopy Mode | 30 or better
102 | Pulsed Fluoroscopy Mode | up to 15 or better
103 | 512 x 512 | 60 or better
104 | 1,024 x 512 | 30 or better
105 | 1,024 x 1,024 | 15 and 30 or better
106 | 2,048 x 2,048 | Please specify
107 | Please state fastest possible acquisition rate for “Angiographic CT” | Please specify
108 | Real-time image processing during acquisition. | Yes
109 | Immediate review of acquired images | Yes
110 | Digital panning. | Yes
111 | Digital runoffs. | Yes
112 | Control panel for all functions in control room | Yes
113 | Quantitative analysis of percent stenosis | Yes
114 | The system must be capable of Digital Subtraction during gantry rotation (often referred to as dynamic rotational angiography). | Yes
115 | System must be capable of doing subtracted bolus chase. | Yes
116 | The digital system should be capable of pulsed progressive acquisition and should include scan conversion of progressively acquired images back into interlaced mode for viewing. | Yes
117 | The digital system should include a tableside control for image review in forward and reverse direction, stepwise or in slow motion. | Yes
118 | The digital system's control functions should include,Roadmap capabilities,Window width and window level adjustment,Image inversion ,Pixel shift ,Pan and zoom,Electronic shutters | Yes
119 | The system should have on-line DSA capabilities in 1024 x 1024 matrix with acquisition frame rate of 1 frame/sec to 6 frames/sec or more 2D roadmapping capabilities | Yes
120 | The system must be capable of 3D and CT imaging | Preferred
121 | The system should have integrated computer controlled (preferably automatic) X-Ray Beam filtering with copper filters of various sizes from 0.2 mm to 0.9 mm in fluro and acquisition mode. Please list the special filters available. | Yes
122 | The system should have the ability to process images (from a previous case or earlier sequence of the current case) while simultaneously displaying roadmap and/or information of the current case (i.e. the ability to process continuously, regardless of the case currently occupying the room. | Preferred
123 | Dose Area Product (DAP)
124 | Resolution | 0.01 µGym2
125 | Maximum reading | 10,000 µGym2
126 | Dose rate linearity | < 2% deviation
127 | DAP/ Cumulative Dose readout display | Yes
128 | Full details regarding patient radiation dose reduction strategies, such as userselectable beam filtration, noise reduction and ability to adjust beam collimators must be provided. | Yes
129 | Foot Switch,
130 | Wired | Yes
131 | Wireless | Yes
132 | Collimators
133 | The collimator must be remote controllable, automatic and Motor Driven | Yes
134 | The collimator should have optional Aluminum Filtration | Yes
135 | The collimator should have optional Copper Filtration | Yes
136 | Table.
137 | Tabletop motion | 4 way
138 | The patient table must be a special-purpose angiographic catheterization table. | Yes
139 | The table controls must be on a console attached to the patient table. | Yes
140 | The patient table must have a motorized adjustable tabletop. | yes
141 | The patient table must permit free selection of gantry angulations while maintaining isocentric positioning. | Yes
142 | Longitudinal travel (cm) | 170 or better
143 | Lateral travel (cm) | ± 20
144 | Vertical travel (cm) | 90 to 116
145 | The stepping table or gantry should include a minimum of four steps with five filming positions | Preffered
146 | Stepping movements must be controlled manually or by computer programmed times, and should be instigated from either the control room or tableside | Preferred
147 | The patient table must have floating, swivel and tilting movements and the configuration must allow maximum patient accessibility. | Yes
148 | The patient table must have a Trendelenburg tilting capability (please indicate range). | Preferred
149 | The tabletop must be capable of supporting a 400-pound patient without deflection at maximum tabletop extension. | Yes
150 | User Interface
151 | Touch Screen interface on examination table with the ability to | Yes
152 | Perform image processing ,acquisition setting ,view basic adjustments,images(live/reference and other parameters(like field size , table positioning,dose information,main monitor control, etc. | Yes
153 | Mount on table | Yes
154 | C-arm Gantry
155 | Configuration | C-Arm
156 | Depth (cm) | 90 or better
157 | SID range (cm) | 90 - 120
158 | Motorized/Manual | Motorized
159 | Rotation speedof FPD from Portrait to L.S(sec) | 3 or less
160 | Rotation range
161 | RAO Projection (degrees) | 120
162 | LAO Projection (degrees) | 120
163 | CC Angulation (degrees) | ± 50
164 | Number of user defined positions | 30 or greater
165 | The imaging gantry must be a C-arm multidirectional gantry suitable for use when catheterizing a patient in the supine position | Yes
166 | The gantry must permit free selection of multiple angulations while maintaining free access to the patient on all sides. | Yes
167 | Motion of the C portion of the gantry must permit variation of the SID | Yes
168 | Gantry construction must be mechanically stable and rigid to prevent vibration during operation. | Yes
169 | Gantry design must provide for rapid positioning and adjustment of positioning by means of motorized operation. | Yes
170 | Safety devices (interlocks) must prevent movements resulting in collision of the gantry with other subsystems and components. | Yes
171 | When not in use, the gantry must be capable of being parked. | Yes
172 | The gantry must incorporate a display of all positioning parameters, including C-arm angulations and SID. | Yes
173 | The overall gantry and table design must permit head-to-toe fluoroscopy (i.e., the imaging chain must travel the entire length of the table). | Yes
174 | The system must be capable of 3D reconstruction DSA. | Yes
175 | Gantry controls must be mounted in a separate mobile console or on a console attached to the patient table. | Yes
176 | Controls must include SID adjustment, emergency stop, and digital stepping if applicable. | Yes
177 | Auto positioning of the image stand and the table must be possible. | Yes
178 | The system should be equipped with 3D Road Mapping | Yes
179 | All motions of the gantry must be isocentric and must maintain orthogonal projections of the central x-ray beam such that the faces of the x-ray tube collimator and the detector remain parallel at all times. | Yes
180 | Lateral L-ARC Gantry
181 | ISO Center to Floor (cm). | 115
182 | Motorized/Manual | Motorized
183 | L-arc rotation speed | (6-10 )deg/sec
184 | Motor-drive angulation | Caudial/Cranial=<50 degree
185 | Multi speed control to position the beam longitudinally in the region of interest: 5 -10 cm /sec | Preferred
186 | Auto-stop in Iso-center. | Yes
187 | Motorized or manual longitudinal movement for parking or positioning. | Yes
188 | Source-image distance (cm) | 90 to130
189 | Motor-driven rotation | .-30 RAO to -110° RAO for lateral arc N or 0° LAO to 90°LAO for lateral arc CN
190 | Display Monitor (Control Rm)
191 | One or Two | Two
192 | CRT or Flat Panel | Flat Panel
193 | Diagonal Size (cm) | 46 cm or higher
194 | Black/White or Color | Color
195 | luminance ratio | Conform with latest ACR requirements
196 | Maximum brightness (Cd/m2) | Conform with latest ACR requirements
197 | Calibration | Conform with DICOM GSDF
198 | Display Monitor (Exam Rm)
199 | The monitor display should be offered with single TFT monitor of 56” or minimum 8 megapixel resolution. Facility of simultaneously display of at least 8 images source inputs analog/digital) should be available to display reference image,3D acquisition imaging and images from other sources like CT/MR. The monitor display system should be ceiling suspended and should allow flexibility in having different image layouts of desired/different format sizes. | Yes
200 | luminance ratio | Conform with latest ACR requirements
201 | Maximum brightness (Cd/m2) | Conform with latest ACR requirements
202 | Calibration | Conform with DICOM GSDF
203 | Image Viewing/Processing
204 | Last image hold | yes
205 | Frame averaging | yes
206 | 3D rotational angiography | yes (3D workstation)
207 | Subtracted rotational angiography | yes
208 | Fluoroscopy recording | yes
209 | 3D processing | yes
210 | 3D road mapping | yes
211 | 2D road mapping | yes
212 | Subtracted peripheral bolus chasing | yes
213 | Land marking | yes
214 | Remasking | yes
215 | Distance & stenosis analysis | yes
216 | Ventricular analysis | yes
217 | Arterial analysis | yes
218 | ROI selection | yes
219 | ROI histogram | yes
220 | Pixel shift | yes (manual & automatic)
221 | Automatic dose reductions | yes
222 | Edge enhancement | yes
223 | Digital real-time filters | yes
224 | Reregistration | yes
225 | Noise reduction algorithms | yes
226 | Sliding weighted averaging | yes
227 | Multi-modality image viewing | optional
228 | Digital zoom | yes
229 | Angiography CT | yes
230 | Image Storage
231 | DICOM image format | yes
232 | Image capture and export to DVD | yes
233 | Image storage capacity
234 | 512 x 512 | Please specify, Larger capacity is preferred
235 | 1,024 x 1,024 | Please specify, Larger capacity is preferred
236 | 2,048 x 2,048 | Please specify, Larger capacity is preferred
237 | Image Compression | DICOM lossless
238 | DICOM Functionality
239 | DICOM work list enabled | yes
240 | DICOM verify/store/send/print | yes
241 | DICOM storage commitment | yes
242 | DICOM media exchange | yes
243 | DICOM MPPS | yes
244 | Other
245 | High End Color Doppler ultrasound scanner with all general/3D vascular capability/Dicom features and probes. | Yes
246 | Integrated ultrasound scanner | (Optional)
247 | US image display on monitors | Yes (in examination room)
248 | Control unit portable trolley | Yes
249 | Portable trolley cable length (m) | 5
250 | Ceiling mounted examination lamp | Yes
251 | Portable automatic injector | Yes
252 | Ceiling mounted Pb shielding panels | Yes (two)
253 | Side-table mounted Pb aprons | Yes (two)
254 | Radiation lowering/dose control features
255 | Patient dose reduction system | yes
256 | Movable shield | yes
257 | Lead aprons | yes
258 | Lead eye googles | yes
259 | Ceiling suspended radiation protection lead glass | yes
260 | Table mounted radiation protection | yes
261 | Ability for X-rays to be turned on/off at table side in the exam room and also in the control room | yes
262 | Dicom structured report should include factors for assessment of patient skin dose (DAP, Cumulative Dose (CD) at reference point, Source to table distance and field size) | yes
263 | Please include a listing and detailed description of all main features and performance specifications, including the following:
264 | Standard 1024 x 1024 functions. | Please Specify (Best preferred)
265 | Functions available from the in-room control | Please Specify (Best preferred)
266 | Selections of dose rates | Please Specify (Best preferred)
267 | Time required between acquisitions and viewing (1024) | Please Specify (Best preferred)
268 | Selections of mask image run delay | Please Specify (Best preferred)
269 | Anatomical land marking capabilities (bone insertion) and selection of percentage un-subtracted | Please Specify (Best preferred)
270 | Image processor including memory specifications | Please Specify (Best preferred)
271 | Ability to selectively move images and/or patients from disk storage to a separate file for review or deletion | Please Specify (Best preferred)
272 | Digital archiving capabilities and formats including number of images at 1K,2K matrix | Please Specify (Best preferred)
273 | EQUIPMENT WARRANTY | The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale. The vendor must indicate clearly the warranty period of all parts, labor and glassware. The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year. The warranty period is to commence after the official institution acceptance testing document has been completed. As per statement 15.4, the institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM 3.0 issues and so forth
274 | ARTICLE 1: GENERAL SUMMARY | 1.1 Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and that:
275 | ARTICLE 1: GENERAL SUMMARY | · the equipment and its accessories are approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and
276 | ARTICLE 1: GENERAL SUMMARY | · the equipment meets the standard of electromagnetic compatibility
277 | ARTICLE 1: GENERAL SUMMARY | 1.2 One (1) original plus five (5) copies including CDs of the tender response shall be provided. Each copy is to include their own set of brochures.
278 | ARTICLE 1: GENERAL SUMMARY | 1.3 The manufacturer will confirm the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
279 | ARTICLE 2: TECHNOLOGY STANDARD | 2.1 All equipment must conform to the site existing DICOM standard and DICOM Worklist must be installed, enabled and seamless.
280 | ARTICLE 2: TECHNOLOGY STANDARD | 2.2 All PACS networking UTP cabling is to be of the CAT6 shielded type.
281 | ARTICLE 2: TECHNOLOGY STANDARD | 2.3 Must provide and install anti-virus that works with the Hospital network.
282 | ARTICLE 2: TECHNOLOGY STANDARD | 2.4 CD ROM & USB port
283 | ARTICLE 2: TECHNOLOGY STANDARD | 2.5 Provide external hard disk for archiving (size: 2 TB; quantity: 2)
284 | ARTICLE 2: TECHNOLOGY STANDARD | 2.6 HITA requirements :
285 | ARTICLE 2: TECHNOLOGY STANDARD | 1) Full connectivity with PACS system (DICOM worklist & sends images to PACS storage)
286 | ARTICLE 2: TECHNOLOGY STANDARD | 2) CDRAM&USB
287 | ARTICLE 2: TECHNOLOGY STANDARD | 3) Capability with hospital antivirus
288 | ARTICLE 2: TECHNOLOGY STANDARD | 4) Capability to send Dose report to radiation server
289 | ARTICLE 2: TECHNOLOGY STANDARD | 5) Q /R from PACS storage
290 | ARTICLE 2: TECHNOLOGY STANDARD | 6) For portable machines ensure it has wireless feature
291 | ARTICLE 2: TECHNOLOGY STANDARD | 7) Should have the MPPS feature to be able to send live status of the procedure.
292 | ARTICLE 2: TECHNOLOGY STANDARD | 8) Can calculate the Radiation dose and send a report to the radiation designated server
293 | ARTICLE 2: TECHNOLOGY STANDARD | 9) Having USB port and CD Drive
294 | ARTICLE 2: TECHNOLOGY STANDARD | 10) Any client software must be supported for Microsoft Windows 7 with IE8 and up to the latest Microsoft Windows Client Platform ,Service Pack & Internet Explorer (Windows 8.1 with IE11)
295 |  | 11) All Power cords must meet Hospital Standard Cords & Plugs
296 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.1 The equipment proposed shall incorporate any “latest generation” hardware and software technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
297 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.2 The vendor will guarantee an official notification to the institution of any equipment updates or model changes occurring immediately before the processing of this order. The vendor shall notify the institution accordingly and grant the institution the option to purchase the upgrade or revised models.
298 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.3 Any components not proposed but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as an option.
299 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.4 Software upgrades that become available within the warranty period shall be incorporated at no additional charge.
300 | ARTICLE 4: EQUIPMENT PRICING | 4.1 Equipment price must be clearly identified and guaranteed for a minimum of three (3) months from the date of receiving vendor quotation.
301 | ARTICLE 4: EQUIPMENT PRICING | 4.2 All prices must include delivery and installation.
302 | ARTICLE 4: EQUIPMENT PRICING | 4.3 If applicable, the vendor must also include the price to dismantle and remove any previously installed equipment.
303 | ARTICLE 4: EQUIPMENT PRICING | 4.4 Prices shall include all accessories as listed in the specifications and as necessary for a complete, fully functional installation. In addition, any optional equipment items listed must be itemized and individually priced in the same fashion.
304 | ARTICLE 4: EQUIPMENT PRICING | 4.5 All training sessions, visits and factory level training fees must be quoted (itemized) in Saudi Riyals.
305 | ARTICLE 4: EQUIPMENT PRICING | 4.6 An itemized price quotation for service contract after the warranty period should be provided for the duration of four (4) years.
306 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.1 The institution reserves the right to accept all or part of any proposal received and thus for
307 | ARTICLE 5: PROPOSAL ACCEPTANCE | RFP proposals that include more than one piece of equipment there can be multiple vendor awarding.
308 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.2 The institution reserves the right not to accept the lowest quoted price.
309 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.3 The institution reserves the right to accept the best suitable offer according to the medical operations requirements as tertiary health care institution
310 | ARTICLE 6: APPLICATION TRAINING | 6.1 Included with each proposal shall be an explanation of the application training provided to department personnel.
311 | ARTICLE 6: APPLICATION TRAINING | 6.2 Instruction and demonstrations will be provided by a qualified application specialist, per the end user’s standards and approval.
312 | ARTICLE 6: APPLICATION TRAINING | 6.3 The institution reserves the right not to accept an application specialist, and the vendor must provide a replacement at no additional cost to the institution.
313 | ARTICLE 6: APPLICATION TRAINING | 6.4 The vendor must provide two (2) application training sessions: the first visit just after equipment clinical acceptance and the second visit within six (6) months. The date of instruction will be designated by mutual agreements of the vendor and the Department of Radiology.
314 | ARTICLE 6: APPLICATION TRAINING | 6.5 The vendor must adhere to the Hijra calendar and KFSH&RC working hours when booking application training sessions. One (1) visit = five (5) KFSH&RC working days: Sunday through to Wednesday from 7:30 a.m. until 5:00 p.m. and Thursday from 7:30 a.m. until 4:00 p.m.
315 | ARTICLE 6: APPLICATION TRAINING | 6.6 The vendor will specify, with advance notification, their requirements for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
316 | ARTICLE 7: SYSTEM SOFTWARE UPGRADES | 7.1 The vendor shall guarantee software upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the institution.
317 | ARTICLE 8: REFERENCES | The vendor must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operate the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities.
318 | ARTICLE 9: INSTITUTION DOCUMENTATION | All plans and specifications must be returned to the owner within one (1) week of notification of award. The successful vendor(s) may retain these documents for reference.
319 | ARTICLE 10: QC TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | In addition to the requested equipment, the vendor shall purchase and deliver the QC test tools, manuals and/or software as per attached Article 29 and Article 30.
320 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.1 For each piece of equipment that is purchased the vendor shall provide advanced, abroad user training for TWO (2) Physicians, one (1) Medical Physicist, TWO (2) Technologists and One (1) factory level training for Clinical Engineering covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation.
321 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.2 Training must be provided during the contract duration and warranty period.
322 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.3 The choice of international offsite training is left to the decision of the institution and the vendor must fulfill the sponsorship befor end of warranty/final acceptance.
323 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.4 All arrangements must be made by the vendor in advance. Selection of Advance Training dates will be designated by the Department of Radiology.
324 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.5 If the vendor fails to provide user training, manuals, software, and/or schedule advanced abroad user training then, the warranty will be extended until all training and/or items are delivered.
325 | ARTICLE 12: EXCEPTIONS TO TERMS AND CONDITIONS | 12.1 Any exceptions to the terms and conditions of the following equipment specifications must be clearly stated and submitted in writing at the time of proposal.
326 | ARTICLE 12: EXCEPTIONS TO TERMS AND CONDITIONS | 12.2 Per statement 12.1, the vendor shall state all exceptions located on the CD’s “Exception to Terms & Condition Checklist” and submit a soft copy with their proposal, as per Technical Specification Sheet (Excel) Article 27.1.
327 | ARTICLE 14: VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | 14.1 The vendor in submitting his quotation represents that he has read and understood the proposal documentations and his quotation is made in accordance there within.
328 | ARTICLE 14: VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | 14.2 Should the vendor representative require any assistance in responding to the said RFP documentation they shall send an official letter to Contracts Management Department.
329 | ARTICLE 15: EQUIPMENT WARRANTY | 15.1 The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale.
330 | ARTICLE 15: EQUIPMENT WARRANTY | 15.2 The vendor must indicate clearly the warranty period of all parts, labor and glassware.
331 | ARTICLE 15: EQUIPMENT WARRANTY | 15.3 The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year.
332 | ARTICLE 15: EQUIPMENT WARRANTY | 15.4 The warranty period is to commence after the official institution acceptance testing document has been completed.
333 | ARTICLE 15: EQUIPMENT WARRANTY | 15.5 As per statement 15.4, the institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM 3.0 issues and so forth.
334 | ARTICLE 16: SCHEDULING AND COORDINATION OF EQUIPMENT INSTALLATION | 16.1 The Projects Management Department (PMD) will assign one (1) staff member to manage the equipment installation, and the vendor project manager and/or service field engineer should henceforth provide all installation relative communications to that assigned individual.
335 | ARTICLE 16: SCHEDULING AND COORDINATION OF EQUIPMENT INSTALLATION | 16.2 The delivery of equipment to the site of installation is subject to prior approval of the owner with at least one (1) month notice by company letter.
336 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.1 Prior to shipment, the vendor shall make suitable arrangements with the owner to examine room(s) accommodating the installation.
337 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.2 Vendor personnel shall familiarize themselves with the features and characteristics of the building which will affect the delivery and/or installation of their equipment. The vendor shall check space, height, accessibility to all applicable areas to ensure ease of installation. Vendors must advise if the proposed equipment will NOT fit through a doorway.
338 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.3 The vendor shall report inadequacies to the owner immediately to facilitate remedial measures prior to scheduling installation.
339 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.4 The vendor shall accept responsibility for error or neglect to familiarize with access, space and service conditions. There will be no storage of furniture or equipment on the job site, without prior authorization in writing from the owner.
340 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.1 The vendor to provide all relevant drawings and/or site preparation documentation.
341 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.2 Vendor should prepare the site to be suitable for the GA cases.
342 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.3 The vendor will be responsible for comprehensive scope of work for site modifications including radiation shielding work.
343 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.1 Delivery authorization will be provided by the owner at the appropriate time. Installation may commence when the finished space is ready to receive equipment and approval has been granted by the owner. The owner will advise the successful vendor of the scheduled commencement of installation, so as to permit the vendor to provide technical personnel in a timely manner. Medical physicist and RSO will be present during the construction phase and all subsequent phases.
344 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.2 The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
345 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.3 The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc.) for each employee they will have on site so that appropriate hospital gate passes can be issued.
346 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.5 Installation shall be done with close cooperation with the assigned Project Manager.
347 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.7 Power supply to the suites will be performed by the institution.
348 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.8 All wiring, wherever possible, is to be contained within the room of installation.
349 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.9 It is the vendor’s responsibility for networking the installed equipment to the nearest PACS edge switch.
350 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.1 All PACS networking UTP cabling is to be of the CAT6 shielded type.
351 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.11 The vendor shall be responsible for all additional work for complete installation beyond the services, etc., specified on the shop drawings as being the responsibility of the owner.
352 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.12 The vendor shall request a final inspection, in writing, of his work by the owner, attend the final inspection and participate in the preparation of a deficiency list, if necessary. The acceptance certificate will not be issued without such an inspection having been completed to the owners’ satisfaction.
353 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.13 The Radiation Safety Officer (RSO) will be present during the construction phase and installation of lead shielding. During the installation of the lead shielding, the vendor shall inform Project Management Department.
354 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.14 The medical physicist will be present during installation and initial testing by the manufacturer.
355 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.15 Project schedule with commencement and completion dates will be submitted to the owner taking into consideration required infection control procedures, sub-contractor holidays, etc.
356 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.16 Project completion including site renovation, delivery, Installation and vendor commissioning must be within 6 weeks.
357 | ARTICLE 20: MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | 20.1 All mobile equipment (and powered accessories) delivered to the institution (e.g., contrast injectors) must be dual voltage 110-220 VAC/60 Hz power (to be operated by 220v as well as 110v as applicable, otherwise 110v is accepted) and power plug/outlet should be American style, medical grade.
358 | ARTICLE 20: MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | 20.2 Mobile equipment not satisfying statement 20.1 will not be accepted by the institution and will ultimately delay clinical acceptance and final contract payment.
359 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.1 The vendor shall keep the premises, including assigned rooms, all passages and auxiliary spaces free from accumulation of waste material or rubbish caused by him at the end of each day.
360 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.2 The vendor shall assume full responsibility for all damages to the building plant and previously placed equipment, including gouges and smudges to the owner’s property made by their own personnel and/or sub-contractors.
361 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.3 Any resulting repairs by the vendor shall match the adjacent existing finishes and materials. The repairs are to be done by appropriate trade contractors without any additional cost to KFSH&RC.
362 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.4 The vendor shall comply with pre-infection control procedures of the hospital.
363 | ARTICLE 22: DOCUMENTATION | 22.1 Two (2) complete operational manuals for hardware and software.
364 | ARTICLE 22: DOCUMENTATION | 22.2 Two (2) service manuals.
365 | ARTICLE 22: DOCUMENTATION | 22.3 Two (2) spare part manuals (if applicable).
366 | ARTICLE 22: DOCUMENTATION | 22.4 Completed schematic diagrams with documentation of onsite modifications (if required).
367 | ARTICLE 22: DOCUMENTATION | 22.5 All above documentation must be provided in English.
368 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.1 Upon completion of installation and prior to first clinical utilization of the equipment, the entire imaging system shall undergo acceptance testing by the Medical Physicists.
369 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.2 The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
370 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Sales Literature;
371 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin;
372 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
373 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Other international regulatory agencies (i.e.: FDA, ACR, AAPM, NEMA, IAEA)
374 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.3 Any deficiencies identified upon acceptance testing must be rectified by the vendor to the institution’s satisfaction prior to commencing clinical acceptance and application training.
375 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.4 The medical physicist shall have access to the same diagnostic and testing software, and QC testing phantoms as the vendor engineer.
376 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | Vendor must provide the following information as it applies to their service organization:
377 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.1 Local service provider (i.e. office address and contact information).
378 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.2 Service response time for routine and after hour calls. The vendor shall define the term ‘Response Time’ and provide their minimum and maximum time limits.
379 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.3 For each piece of equipment specified within this document, list parts that are critical for normal equipment operation but are not available from a local service center and would take more than seven (7) working days to be delivered to the site.
380 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.4 Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have or will be factory trained on the specified equipment.
381 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.5 The vendor shall guarantee to provide to this institution with technical support and the supply of spare parts for a period of five (5) years after the date of last equipment production. The vendor shall also inform the institution of the end of production date when they are announced.
382 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.6 The vendor shall provide for each piece of equipment purchased a full-service contract for a period of four (4) years after the warranty period has expired.
383 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.7 Equipment downtime (DT) shall not exceed 3% of the total operation time per annum; downtime computations will assume 52 weeks per year and 44 work hours per week, including Ramadan and Hajj holidays.
384 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.8 Service response time for routine calls should be one (1) hour, including Ramadan and Hajj holidays. Penalties may ensue should the engineer not respond in a timely manner (3% DT).
385 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.9 Service engineer must provide a service report each time they evaluate and/or work on the equipment. Penalties will ensue should the engineer not comply (2% DT).
386 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.10 All empty boxes, tools, spare parts, defective parts, replacement parts and the like, must be removed from the premises immediately and/or temporarily stored in the RIS/PACS storage room as long as a property pass is in process. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc…
387 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.11 Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized, clutter-free and nothing on the floors).
388 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | The Bi-Plane Angiography System must be equipped with state-of-the-art medical imaging technology including advanced perfusion and fully integrated processing features.
389 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | In addition to the requested equipment, the vendor shall purchase, deliver, and install (if applicable) the following items for each system:
390 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | 25.1 Original software license to be provided and a back-up copy of the software. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge after the warranty period.
391 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | 25.2 The vendor should provide listing (with documentation) of all available software packages and should indicate clearly which software applications are offered as standard packages and which software is optional. Itemized software package price listing should also be provided.
392 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | 25.3 One (1) UPS system to protect the system from unexpected power surges.
393 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | 25.4 Sufficient cabinets (with shelving) should be installed in the examination room to store: equipment accessories, QC test phantoms, manuals (or other documents) and spare parts.
394 | ARTICLE 25 : Bi-Plane Angiography System (Quantity: 1) | 25.5 The imaging room should be well designed, and ceiling visual therapy provided.
395 | ARTICLE 28: SITE PREPARATION | COMPREHENSIVE SCOPE OF WORK - PROJECT MANAGEMENT (PMD) - Refer to CD submitted along with the RFP for PMD Requirements – Comprehensive SOW
396 | ARTICLE 29: QC TOOLS | TOR 18FG Leeds Test Objects
397 | ARTICLE 29: QC TOOLS | TOR CDR Leeds Test Objects
398 | ARTICLE 29: QC TOOLS | TO AEC Leeds Test Objects
399 | ARTICLE 29: QC TOOLS | PIX 13 Leeds Test Objects
400 | ARTICLE 29: QC TOOLS | PMMA/Cu Attenuator Leeds Test Objects
401 | ARTICLE 29: QC TOOLS | RaySafe DXR+ RaySafe
402 | ARTICLE 29: QC TOOLS | RaySafe P Fluoro Phantom RaySafe
403 | ARTICLE 29: QC TOOLS | RaySafe Pro-Digi Radiography Phantom V
404 | ARTICLE 29: QC TOOLS | RaySafe Pro-Slit
405 | ARTICLE 29: QC TOOLS | RaySafe Pro-Stand
406 | ARTICLE 29: QC TOOLS | X-Check® DSA Test Object, PTW
407 | ARTICLE 29: QC TOOLS | DIAVOLT UNIVERSAL All-in-one X-ray Meter for Clinical Engineering (QTY:2)
408 | Lead Apron two pieces ( top & skirt ), Attached thyroid shield | QTY. 15
409 | Color Printer ( Hospital standard ) | QTY. 1
410 | Regular Printer (Hospital standard ) | QTY. 1
411 | PC Copmuter ( Hospital standard ) | QTY. 4
412 | Lead Glasses | QTY. 10
413 | Lead Aprons hangers | QTY. 3
414 | Wide X-ray Mobile Barrier | QTY. 2
415 | Employees metal lockers ( single door lockers )tilted from above, each one single door contain 4 lockers | QTY. 3
416 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 406320103 | 1123278 | FLUOROSCOPY MACHINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Manufacture | Specify
2 | Model Name | Specify
3 | FDA Approval year | Specify
4 | Application and feature
5 | Gastroenterology | Yes
6 | Skeleton | Yes
7 | Thorax and lungs | Yes
8 | Urology/Gynecology | Yes
9 | Emergency | Yes
10 | Digital angiography | Optional
11 | Pediatrics | Yes
12 | Configuration
13 | Design type | Over table
14 | One system with dual tube /chest stand unit with wireless detector | Yes
15 | Generator
16 | Type | 3 phase, high frequency
17 | Power | 80KW or better
18 | Digital Radiography Mode
19 | kVp range | 40 - 150 or better
20 | No. of kVp increments | .+/- 1kv
21 | mA range | up to 800 or better
22 | No. of mA increments | Please specify
23 | mAs range | Please specify
24 | No. of mAs increments | Please specify
25 | AEC | yes
26 | Exposure time range (sec) | 0.001 - 1.5 or better
27 | No. of time increments | Please specify
28 | Anatomical programming | yes
29 | Continuous Fluoroscopy Mode
30 | Aqyuisition mode continuous fluoroscopy | up to 30 or better
31 | kVp range | 40 - 125 or better
32 | mA range | up to 10 or better
33 | ABS control | yes (kVp and mA)
34 | Maximum exposure rate (R/min) | 10 or better
35 | Pulsed Fluoroscopy Mode
36 | Acquisition Mode Pulsed Fluoroscopy | 20 or better
37 | X-ray Tube (Fluoroscopy)
38 | Tube Power Rating (kW) | 80 or greater.
39 | Nominal kVp | 125 or greater.
40 | Anode Heat Storage Capacity (MHU) | 783,000HU
41 | Heat Dissipation Rate (kHU/min) | 300 or greater.
42 | Grid switch mechanism for pulse fluoroscopy | Yes
43 | Cooling System Type | please specify
44 | Focal Spot | Dual
45 | Micro (mm) | Please specify
46 | Small (mm) | Please specify
47 | Large (mm) | Please specify
48 | X-ray Tube (Over Head)
49 | Tube Power Rating (kW) | 80 or greater.
50 | Nominal kVp | 150 or better
51 | Anode Heat Storage Capacity (MHU) | 783000 HU or better
52 | Heat Dissipation Rate (kHU/min) | 300 or greater.
53 | Cooling System Type | Please specify
54 | Focal Spot | Dual
55 | Micro (mm) | Please specify
56 | Small (mm) | Please specify
57 | Large (mm) | Please specify
58 | Anode Speed (RPM) | 5000 or higher
59 | Flat Panel X-ray Detector(Main)
60 | Manufacture | Specify
61 | Size (cm) | 43x43
62 | Actual Physical Detector Size including the housing (cm) in x,y directions | Please specify
63 | Active Imaging Area (cm) in x and y directions | (Please specify), the detector with larger active imaging area will be considered
64 | Number of Detector Elements in x and y directions | specify, larger matrix size will be considered
65 | DQE for the range of frequencies (lp/mm); Provide amplitude and frequency pairs from 0 to Nyquist frequency. | (≥77%@0LP/mm) or Specify DQE curve for the entire frequency range, Best performer will be considered
66 | MTF for the range of frequencies (lp/mm); Provide amplitude and frequency pairs from 0 to Nyquist frequency. | Specify MTF curve for the entire frequency range, Best performer will be considered
67 | Swival or rotatable | yes
68 | Resolution | 2,500 x 2,500 ( Best prefered)
69 | Pixel Size (mm) | 148 micro (lower preferable)
70 | Semiconductor material | a-si
71 | Scintillator | CSI
72 | Acquisition depth (pixel depth) | 16 bit
73 | Maximum acquisition speed (im/s) | up to 30
74 | Maximum diagonal size (cm) | 43cm x 43cm
75 | Input field (cm) | Please specify
76 | Digital acquisition zoom fields | Please specify
77 | Image Quality
78 | Spatial resolution (lp/cm) | 3LP/MM(Higher preferable)
79 | Flat Panel X-ray Detector(Wall stand)
80 | Size (cm) | 14x17 inch
81 | Resolution | 3 lp/mm
82 | Pixel Size (mm) | 148 or less
83 | Semiconductor material | a-si
84 | Scintillator | CSI
85 | wireless | prefered
86 | Battery operation time(hrs) | 4 or better
87 | DQE for the range of frequencies (lp/mm); Provide amplitude and frequency pairs from 0 to Nyquist frequency. | 75% or better @0lp/mm
88 | Detector battery charging time(hrs) | 3 or less
89 | Battery indicator and charger | Yes
90 | weight with battery(kg) | 3 or less
91 | Acquisition depth (pixel depth) | 16 bit
92 | Maximum acquisition speed (im/s) | up to 30
93 | Image Acquisition
94 | Matrix Size | 2,880 pixel x 2,881 pixel
95 | Resolution (pixel depth) | 16 bit or better
96 | Frame Rates (fps) | up to 30 Frames / sec
97 | Continuous Fluoroscopy Mode | 30 or better
98 | Pulsed Fluoroscopy Mode | up to 30 or better
99 | 512 x 512 | 60 or better
100 | 1,024 x 512 | Please specify higher preferable
101 | 1,024 x 1,024 | 15 and 30 or better
102 | 2,048 x 2,048 | Please specify higher preferable
103 | Dose Area Product (DAP)
104 | Active area (mm) | Please specify
105 | Resolution | 0.01 mGym2
106 | Maximum reading | 10,000 mGym2
107 | Dose rate linearity | < 2% deviation
108 | DAP readout display | yes
109 | Fluoroscopiy Table
110 | Dimension(cm) | 225x73
111 | Radiotransparent area(cm) | 220x50
112 | Transversal motorized travel | please specify
113 | Material | Please specify
114 | Tabletop motion | 4 way
115 | Longitudinal travel (cm) | 120 or better
116 | Lateral travel (cm) | ± 20
117 | Vertical motorized travel (cm) | 90 to 116
118 | Patient weight capacity (kg) | up to 284 Kg
119 | Tabletop to floor minimum distance (CM) | Please specify (minimum prefered)
120 | Tilting | .+/-90 degree
121 | Tiliting speed/angle display | Yes
122 | local table function control | Yes
123 | Remote touch screen display | Yes
124 | Compression | Yes(Motorized)
125 | Maximum compression force (KG) | Up to 15 or better
126 | Safety | Yes(Specify)
127 | Collimator | Yes(Automatic)
128 | Touch Screen Geometry Console on table | Yes
129 | Display Monitor (Control Rm)
130 | One or Two | Two
131 | CRT or Flat Panel | flat panel
132 | Diagonal Size (cm) | 46 cm or higher
133 | Black/White or Color | color
134 | Image display | Please specify
135 | Contrast ratio | Please specify
136 | Maximum brightness (Cd/m2) | Please specify
137 | Display Monitor (Exam Rm)
138 | One or Two | Two
139 | CRT or Flat Panel | flat panel
140 | Diagonal Size (cm) | 46 cm or higher
141 | Black/White or Color | color
142 | Image display | Please specify
143 | Contrast ratio | Please specify
144 | Maximum brightness (Cd/m2) | Please specify
145 | Image Viewing/Processing
146 | Last image hold | yes
147 | Frame averaging | yes
148 | 3D processing | Optional
149 | Image Storage
150 | DICOM image format | yes
151 | Image capture and export to DVD | yes
152 | Image storage capacity | 512 x 512 (80,000 or higher)
153 |  | 1,024 x 1,024 (50,000 or higher)
154 |  | 2,048 x 2,048 (10,000 or higher)
155 | Image Compression | DICOM lossless
156 | DICOM Functionality
157 | DICOM worklist enabled | yes
158 | DICOM verify/store/send/print | yes
159 | DICOM storage commitment | yes
160 | DICOM media exchange | yes
161 | DICOM MPPS | yes
162 | Wall Stand Assembly
163 | Tilt Range (deg) | (-10 / + 90)
164 | Vertical Travel Range (cm) | (0 to 150 or wider)
165 | Swing (angulation) Range (cm) | (+/- 45)
166 | Control Keypad on Stand | Yes
167 | Rotatable Bucky | Optional
168 | No. of AEC Chambers | Three (3)
169 | No. of AEC Density Steps | Specify
170 | Retractable Detector Assembly | Yes
171 | Buck Synchronization | Yes
172 | Auto Positioning | Preferred
173 | Dimensions (cm) L x W x H | Specify
174 | Other features
175 | system incorporate video & microphone for modified barium swallow videofluoroscopy for speech evaluation that can be save on DVD. | Yes
176 | Machine should capable of having adjustable DID that can fit in stetcher with patient for modified barrium swallow (video fluoroscopy) | 115 to 180(cm) or better
177 | overhead tube and a wireless detector for cross table X-Ray | Yes(One system have this overhead tube configuration)
178 | Foot switch | Wireless=Yes, Wire=Yes
179 | Detail description of X-Ray dose reduction /control feature with comprehensive details, hence best software or hardware fetured device will be considered. | Yes
180 | Local Refernce site availabilty | Prefered
181 | Trained service enginners ( local) | Yes(please mention in numbers)
182 | Local Spare parts stock | Prefered
183 | Premium ultrasound system with vescular package. | Yes(1)
184 | Number of Recalls of offered model | Specify in numbers
185 | System must be incorporate with collision sensors in table and over head tube suspension. | Yes
186 | Autopositioning of over head suspansion with chest stand and to parking position | Preferred
187 | Integrated area dose calculator | Yes
188 | EQUIPMENT WARRANTY | The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale. The vendor must indicate clearly the warranty period of all parts, labor and glassware. The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year. The warranty period is to commence after the official institution acceptance testing document has been completed. As per statement 15.4, the institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM 3.0 issues and so forth
189 | ARTICLE 1: GENERAL SUMMARY | 1.1 Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and that:
190 | ARTICLE 1: GENERAL SUMMARY | · the equipment and its accessories are approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and
191 | ARTICLE 1: GENERAL SUMMARY | · the equipment meets the standard of electromagnetic compatibility
192 | ARTICLE 1: GENERAL SUMMARY | 1.2 One (1) original plus five (5) copies including CDs of the tender response shall be provided. Each copy is to include their own set of brochures.
193 | ARTICLE 1: GENERAL SUMMARY | 1.3 The manufacturer will confirm the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
194 | ARTICLE 2: TECHNOLOGY STANDARD | 2.1 All equipment must conform to the site existing DICOM standard and DICOM Worklist must be installed, enabled and seamless.
195 | ARTICLE 2: TECHNOLOGY STANDARD | 2.2 All PACS networking UTP cabling is to be of the CAT6 shielded type.
196 | ARTICLE 2: TECHNOLOGY STANDARD | 2.3 Must provide and install anti-virus that works with the Hospital network.
197 | ARTICLE 2: TECHNOLOGY STANDARD | 2.4 CD ROM & USB port
198 | ARTICLE 2: TECHNOLOGY STANDARD | 2.5 Provide external hard disk for archiving (size: 2 TB; quantity: 2)
199 | ARTICLE 2: TECHNOLOGY STANDARD | 2.6 HITA requirements :
200 | ARTICLE 2: TECHNOLOGY STANDARD | 1) Full connectivity with PACS system (DICOM worklist & sends images to PACS storage)
201 | ARTICLE 2: TECHNOLOGY STANDARD | 2) CDRAM&USB
202 | ARTICLE 2: TECHNOLOGY STANDARD | 3) Capability with hospital antivirus
203 | ARTICLE 2: TECHNOLOGY STANDARD | 4) Capability to send Dose report to radiation server
204 | ARTICLE 2: TECHNOLOGY STANDARD | 5) Q /R from PACS storage
205 | ARTICLE 2: TECHNOLOGY STANDARD | 6) For portable machines ensure it has wireless feature
206 | ARTICLE 2: TECHNOLOGY STANDARD | 7) Should have the MPPS feature to be able to send live status of the procedure.
207 | ARTICLE 2: TECHNOLOGY STANDARD | 8) Can calculate the Radiation dose and send a report to the radiation designated server
208 | ARTICLE 2: TECHNOLOGY STANDARD | 9) Having USB port and CD Drive
209 | ARTICLE 2: TECHNOLOGY STANDARD | 10) Any client software must be supported for Microsoft Windows 7 with IE8 and up to the latest Microsoft Windows Client Platform ,Service Pack & Internet Explorer (Windows 8.1 with IE11)
210 | ARTICLE 2: TECHNOLOGY STANDARD | 11) All Power cords must meet Hospital Standard Cords & Plugs
211 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.1 The equipment proposed shall incorporate any “latest generation” hardware and software technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
212 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.2 The vendor will guarantee an official notification to the institution of any equipment updates or model changes occurring immediately before the processing of this order. The vendor shall notify the institution accordingly and grant the institution the option to purchase the upgrade or revised models.
213 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.3 Any components not proposed but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as an option.
214 | ARTICLE 3: TECHNOLOGY ENHANCEMENTS | 3.4 Software upgrades that become available within the warranty period shall be incorporated at no additional charge.
215 | ARTICLE 4: EQUIPMENT PRICING | 4.1 Equipment price must be clearly identified and guaranteed for a minimum of three (3) months from the date of receiving vendor quotation.
216 | ARTICLE 4: EQUIPMENT PRICING | 4.2 All prices must include delivery and installation.
217 | ARTICLE 4: EQUIPMENT PRICING | 4.3 If applicable, the vendor must also include the price to dismantle and remove any previously installed equipment.
218 | ARTICLE 4: EQUIPMENT PRICING | 4.4 Prices shall include all accessories as listed in the specifications and as necessary for a complete, fully functional installation. In addition, any optional equipment items listed must be itemized and individually priced in the same fashion.
219 | ARTICLE 4: EQUIPMENT PRICING | 4.5 All training sessions, visits and factory level training fees must be quoted (itemized) in Saudi Riyals.
220 | ARTICLE 4: EQUIPMENT PRICING | 4.6 An itemized price quotation for service contract after the warranty period should be provided for the duration of four (4) years.
221 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.1 The institution reserves the right to accept all or part of any proposal received and thus for
222 | ARTICLE 5: PROPOSAL ACCEPTANCE | RFP proposals that include more than one piece of equipment there can be multiple vendor awarding.
223 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.2 The institution reserves the right not to accept the lowest quoted price.
224 | ARTICLE 5: PROPOSAL ACCEPTANCE | 5.3 The institution reserves the right to accept the best suitable offer according to the medical operations requirements as tertiary health care institution
225 | ARTICLE 6: APPLICATION TRAINING | 6.1 Included with each proposal shall be an explanation of the application training provided to department personnel.
226 | ARTICLE 6: APPLICATION TRAINING | 6.2 Instruction and demonstrations will be provided by a qualified application specialist, per the end user’s standards and approval.
227 | ARTICLE 6: APPLICATION TRAINING | 6.3 The institution reserves the right not to accept an application specialist, and the vendor must provide a replacement at no additional cost to the institution.
228 | ARTICLE 6: APPLICATION TRAINING | 6.4 The vendor must provide two (2) application training sessions: the first visit just after equipment clinical acceptance and the second visit within six (6) months. The date of instruction will be designated by mutual agreements of the vendor and the Department of Radiology.
229 | ARTICLE 6: APPLICATION TRAINING | 6.5 The vendor must adhere to the Hijra calendar and KFSH&RC working hours when booking application training sessions. One (1) visit = five (5) KFSH&RC working days: Sunday through to Wednesday from 7:30 a.m. until 5:00 p.m. and Thursday from 7:30 a.m. until 4:00 p.m.
230 | ARTICLE 6: APPLICATION TRAINING | 6.6 The vendor will specify, with advance notification, their requirements for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
231 | ARTICLE 7: SYSTEM SOFTWARE UPGRADES | 7.1 The vendor shall guarantee software upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the institution.
232 | ARTICLE 8: REFERENCES | The vendor must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operate the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities.
233 | ARTICLE 9: INSTITUTION DOCUMENTATION | All plans and specifications must be returned to the owner within one (1) week of notification of award. The successful vendor(s) may retain these documents for reference.
234 | ARTICLE 10: QC TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | In addition to the requested equipment, the vendor shall purchase and deliver the QC test tools, manuals and/or software as per attached Article 29 and Article 30.
235 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.1 For each piece of equipment that is purchased the vendor shall provide advanced, abroad user training for TWO (2) Physicians, one (1) Medical Physicist, TWO (2) Technologists and One (1) factory level training for Clinical Engineering covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation.
236 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.2 Training must be provided during the contract duration and warranty period.
237 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.3 The choice of international offsite training is left to the decision of the institution and the vendor must fulfill the sponsorship befor end of warranty/final acceptance.
238 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.4 All arrangements must be made by the vendor in advance. Selection of Advance Training dates will be designated by the Department of Radiology.
239 | ARTICLE 11: STAFF TECHNOLOGY EXCHANGE (ADVANCED TRAINING) | 11.5 If the vendor fails to provide user training, manuals, software, and/or schedule advanced abroad user training then, the warranty will be extended until all training and/or items are delivered.
240 | ARTICLE 12: EXCEPTIONS TO TERMS AND CONDITIONS | 12.1 Any exceptions to the terms and conditions of the following equipment specifications must be clearly stated and submitted in writing at the time of proposal.
241 | ARTICLE 12: EXCEPTIONS TO TERMS AND CONDITIONS | 12.2 Per statement 12.1, the vendor shall state all exceptions located on the CD’s “Exception to Terms & Condition Checklist” and submit a soft copy with their proposal, as per Technical Specification Sheet (Excel) Article 27.1.
242 | ARTICLE 14: VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | 14.1 The vendor in submitting his quotation represents that he has read and understood the proposal documentations and his quotation is made in accordance there within.
243 | ARTICLE 14: VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | 14.2 Should the vendor representative require any assistance in responding to the said RFP documentation they shall send an official letter to Contracts Management Department.
244 | ARTICLE 15: EQUIPMENT WARRANTY | 15.1 The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale.
245 | ARTICLE 15: EQUIPMENT WARRANTY | 15.2 The vendor must indicate clearly the warranty period of all parts, labor and glassware.
246 | ARTICLE 15: EQUIPMENT WARRANTY | 15.3 The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year.
247 | ARTICLE 15: EQUIPMENT WARRANTY | 15.4 The warranty period is to commence after the official institution acceptance testing document has been completed.
248 | ARTICLE 15: EQUIPMENT WARRANTY | 15.5 As per statement 15.4, the institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM 3.0 issues and so forth.
249 | ARTICLE 16: SCHEDULING AND COORDINATION OF EQUIPMENT INSTALLATION | 16.1 The Projects Management Department (PMD) will assign one (1) staff member to manage the equipment installation, and the vendor project manager and/or service field engineer should henceforth provide all installation relative communications to that assigned individual.
250 | ARTICLE 16: SCHEDULING AND COORDINATION OF EQUIPMENT INSTALLATION | 16.2 The delivery of equipment to the site of installation is subject to prior approval of the owner with at least one (1) month notice by company letter.
251 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.1 Prior to shipment, the vendor shall make suitable arrangements with the owner to examine room(s) accommodating the installation.
252 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.2 Vendor personnel shall familiarize themselves with the features and characteristics of the building which will affect the delivery and/or installation of their equipment. The vendor shall check space, height, accessibility to all applicable areas to ensure ease of installation. Vendors must advise if the proposed equipment will NOT fit through a doorway.
253 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.3 The vendor shall report inadequacies to the owner immediately to facilitate remedial measures prior to scheduling installation.
254 | ARTICLE 17: EXAMINATION OF SPACE AND SERVICES (when applicable) | 17.4 The vendor shall accept responsibility for error or neglect to familiarize with access, space and service conditions. There will be no storage of furniture or equipment on the job site, without prior authorization in writing from the owner.
255 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.1 The vendor to provide all relevant drawings and/or site preparation documentation.
256 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.2 Vendor should prepare the site to be suitable for the GA cases.
257 | ARTICLE 18: SITE PREPARATION / VENDOR SCOPE OF WORK (Refer to attached PMD specifications and requirements article 28) | 18.3 The vendor will be responsible for comprehensive scope of work for site modifications including radiation shielding work.
258 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.1 Delivery authorization will be provided by the owner at the appropriate time. Installation may commence when the finished space is ready to receive equipment and approval has been granted by the owner. The owner will advise the successful vendor of the scheduled commencement of installation, so as to permit the vendor to provide technical personnel in a timely manner. Medical physicist and RSO will be present during the construction phase and all subsequent phases.
259 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.2 The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
260 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.3 The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc.) for each employee they will have on site so that appropriate hospital gate passes can be issued.
261 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.5 Installation shall be done with close cooperation with the assigned Project Manager.
262 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.7 Power supply to the suites will be performed by the institution.
263 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.8 All wiring, wherever possible, is to be contained within the room of installation.
264 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.9 It is the vendor’s responsibility for networking the installed equipment to the nearest PACS edge switch.
265 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.1 All PACS networking UTP cabling is to be of the CAT6 shielded type.
266 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.11 The vendor shall be responsible for all additional work for complete installation beyond the services, etc., specified on the shop drawings as being the responsibility of the owner.
267 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.12 The vendor shall request a final inspection, in writing, of his work by the owner, attend the final inspection and participate in the preparation of a deficiency list, if necessary. The acceptance certificate will not be issued without such an inspection having been completed to the owners’ satisfaction.
268 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.13 The Radiation Safety Officer (RSO) will be present during the construction phase and installation of lead shielding. During the installation of the lead shielding, the vendor shall inform Project Management Department.
269 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.14 The medical physicist will be present during installation and initial testing by the manufacturer.
270 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.15 Project schedule with commencement and completion dates will be submitted to the owner taking into consideration required infection control procedures, sub-contractor holidays, etc.
271 | ARTICLE 19: DELIVERY AND INSTALLATION | 19.16 Project completion including site renovation, delivery, Installation and vendor commissioning must be within 6 weeks.
272 | ARTICLE 20: MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | 20.1 All mobile equipment (and powered accessories) delivered to the institution (e.g., contrast injectors) must be dual voltage 110-220 VAC/60 Hz power (to be operated by 220v as well as 110v as applicable, otherwise 110v is accepted) and power plug/outlet should be American style, medical grade.
273 | ARTICLE 20: MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | 20.2 Mobile equipment not satisfying statement 20.1 will not be accepted by the institution and will ultimately delay clinical acceptance and final contract payment.
274 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.1 The vendor shall keep the premises, including assigned rooms, all passages and auxiliary spaces free from accumulation of waste material or rubbish caused by him at the end of each day.
275 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.2 The vendor shall assume full responsibility for all damages to the building plant and previously placed equipment, including gouges and smudges to the owner’s property made by their own personnel and/or sub-contractors.
276 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.3 Any resulting repairs by the vendor shall match the adjacent existing finishes and materials. The repairs are to be done by appropriate trade contractors without any additional cost to KFSH&RC.
277 | ARTICLE 21: CLEANING UP AND MAKING GOOD DAMAGES | 21.4 The vendor shall comply with pre-infection control procedures of the hospital.
278 | ARTICLE 22: DOCUMENTATION | 22.1 Two (2) complete operational manuals for hardware and software.
279 | ARTICLE 22: DOCUMENTATION | 22.2 Two (2) service manuals.
280 | ARTICLE 22: DOCUMENTATION | 22.3 Two (2) spare part manuals (if applicable).
281 | ARTICLE 22: DOCUMENTATION | 22.4 Completed schematic diagrams with documentation of onsite modifications (if required).
282 | ARTICLE 22: DOCUMENTATION | 22.5 All above documentation must be provided in English.
283 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.1 Upon completion of installation and prior to first clinical utilization of the equipment, the entire imaging system shall undergo acceptance testing by the Medical Physicists.
284 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.2 The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
285 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Sales Literature;
286 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin;
287 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
288 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | • Other international regulatory agencies (i.e.: FDA, ACR, AAPM, NEMA, IAEA)
289 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.3 Any deficiencies identified upon acceptance testing must be rectified by the vendor to the institution’s satisfaction prior to commencing clinical acceptance and application training.
290 | ARTICLE 23: EQUIPMENT ACCEPTANCE TESTING | 23.4 The medical physicist shall have access to the same diagnostic and testing software, and QC testing phantoms as the vendor engineer.
291 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | Vendor must provide the following information as it applies to their service organization:
292 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.1 Local service provider (i.e. office address and contact information).
293 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.2 Service response time for routine and after hour calls. The vendor shall define the term ‘Response Time’ and provide their minimum and maximum time limits.
294 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.3 For each piece of equipment specified within this document, list parts that are critical for normal equipment operation but are not available from a local service center and would take more than seven (7) working days to be delivered to the site.
295 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.4 Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have or will be factory trained on the specified equipment.
296 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.5 The vendor shall guarantee to provide to this institution with technical support and the supply of spare parts for a period of five (5) years after the date of last equipment production. The vendor shall also inform the institution of the end of production date when they are announced.
297 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.6 The vendor shall provide for each piece of equipment purchased a full-service contract for a period of four (4) years after the warranty period has expired.
298 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.7 Equipment downtime (DT) shall not exceed 3% of the total operation time per annum; downtime computations will assume 52 weeks per year and 44 work hours per week, including Ramadan and Hajj holidays.
299 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.8 Service response time for routine calls should be one (1) hour, including Ramadan and Hajj holidays. Penalties may ensue should the engineer not respond in a timely manner (3% DT).
300 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.9 Service engineer must provide a service report each time they evaluate and/or work on the equipment. Penalties will ensue should the engineer not comply (2% DT).
301 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.10 All empty boxes, tools, spare parts, defective parts, replacement parts and the like, must be removed from the premises immediately and/or temporarily stored in the RIS/PACS storage room as long as a property pass is in process. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc…
302 | ARTICLE 24: MANUFACTURER’S SERVICE COMMITMENT | 24.11 Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized, clutter-free and nothing on the floors).
303 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | The Fluoroscopy System must be equipped with state-of-the-art medical imaging technology. One system must be dual tube /chest stand unit with wireless detector
304 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | In addition to the requested equipment, the vendor shall purchase, deliver, and install (if applicable) the following items for each system:
305 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | 25.6 Original software license to be provided and a back-up copy of the software. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge after the warranty period.
306 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | 25.7 The vendor should provide listing (with documentation) of all available software packages and should indicate clearly which software applications are offered as standard packages and which software is optional. Itemized software package price listing should also be provided.
307 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | 25.8 One (1) UPS system to protect the system from unexpected power surges.
308 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | 25.9 Sufficient cabinets (with shelving) should be installed in the examination room to store: equipment accessories, QC test phantoms, manuals (or other documents) and spare parts.
309 | ARTICLE 25 : Fluoroscopy System (Quantity: 2) | 25.10 The imaging room should be well designed, and ceiling visual therapy provided.
310 | ARTICLE 28: SITE PREPARATION | COMPREHENSIVE SCOPE OF WORK - PROJECT MANAGEMENT (PMD) - Refer to CD submitted along with the RFP for PMD Requirements – Comprehensive SOW
311 | ARTICLE 29: QC TOOLS | TOR 18FG Leeds Test Objects
312 | ARTICLE 29: QC TOOLS | TOR CDR Leeds Test Objects
313 | ARTICLE 29: QC TOOLS | TO AEC Leeds Test Objects
314 | ARTICLE 29: QC TOOLS | PIX 13 Leeds Test Objects
315 | ARTICLE 29: QC TOOLS | PMMA/Cu Attenuator Leeds Test Objects
316 | ARTICLE 29: QC TOOLS | RaySafe DXR+ RaySafe
317 | ARTICLE 29: QC TOOLS | RaySafe P Fluoro Phantom RaySafe
318 | ARTICLE 29: QC TOOLS | RaySafe Pro-Digi Radiography Phantom V
319 | ARTICLE 29: QC TOOLS | RaySafe Pro-Slit
320 | ARTICLE 29: QC TOOLS | RaySafe Pro-Stand
321 | ARTICLE 29: QC TOOLS | X-Check® DSA Test Object, PTW
322 | ARTICLE 29: QC TOOLS | DIAVOLT UNIVERSAL All-in-one X-ray Meter for Clinical Engineering (QTY:2)
323 | Lead Apron two pieces ( top & skirt ), Attached thyroid shield | QTY. 15
324 | Color Printer ( Hospital standard ) | QTY. 1
325 | Regular Printer (Hospital standard ) | QTY. 1
326 | PC Copmuter ( Hospital standard ) | QTY. 4
327 | Lead Glasses | QTY. 10
328 | Lead Aprons hangers | QTY. 3
329 | Wide X-ray Mobile Barrier | QTY. 2
330 | Employees metal lockers ( single door lockers )tilted from above, each one single door contain 4 lockers | QTY. 3
331 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 417857201 | 1133623 | FLOW CYTOMETER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MANUFACTURER | SPECIFY
2 | MODEL
3 | CE MARK/FDA approval | Required, provide certificate number
4 | RESEARCH APPLICATIONS | Wide variety of research applications , > 9 color panel Analysis
5 | CLINICAL RESEARCH | For Research use only
6 | SAMPLE VIALS | 12 X 75 mm tube required
7 | LIGHT SOURCE | 3 Lasers minimum.
8 | FLUORESCENCE | 8 Colors minimum and more is preferred.
9 | PARAMETERS | Forward scatter, 90 degrees scatter
10 | FLOW RATE | Adjustable dynamic range of sample flow rates preferred
11 | DETECTOR
12 | 1- RANGE nm | 300-1000 nm preferred
13 | 2- Sensitivity | < 120 MESF at FITC < 100 MESF at PE <100 MESF at APC
14 | 3- Minimum cell size | < 0.5 UM
15 | 4- Maximum analysis rate, cells/sec | >10,000
16 | STORAGE | 8 GB RAM, > 1 TB HD
17 | MICROPROCESSOR | 4G HZ
18 | WINDOWS | Windows 10
19 | CLEANING CYCLE | Required
20 | SPECIAL FEATURES | Please Explore
21 | installation | Pre-installation and site preparation (mechanical &electrical, except civil), is the responsibility of the equipment vendor. The vendor has the right to inspect the site and prior to quotation and he must coordinate with hospital clinical engineering, facility project management and the main contractor for any work at site before starting.
22 | WARRANTY | The vendor is responsible to provide a minimum of one (1) year full warranty commencing from the date of first case
25 | WARRANTY | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
26 | WARRANTY | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
27 | WARRANTY | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
28 | WARRANTY | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
29 | WARRANTY | The contractor will provide to hospital all the essential tools and test equipment required to maintain the equipment free of charge.
30 | WARRANTY | Pre-installation and site preparation (mechanical &electrical, except civil), is the responsibility of the equipment vendor. The vendor has the right to inspect the site and prior to quotation and he must coordinate with hospital clinical engineering, facility project management and the main contractor for any work at site before starting.
31 | WARRANTY | The contractor should provide the hospital with Installation, operation and performance qualification certificate at the time of installation.
32 | Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
33 | Software and upgrades | b) If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
35 | Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
36 | Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
37 | Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
39 | Power Requirement | a) Equipment must be 220 V AC.
40 | Power Requirement | b) The Frequency must be 60Hz.
41 | Power Requirement | c) The above applies to all equipment unless stated in its specification table.
42 | Power Requirement | d) UPS back up power.
44 | Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
45 | Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
47 | local Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
48 | local Training | b) Full local service training to be provided, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
49 | The training minimum contents are: | I.Theory of Operation including circuit description and schematics.
50 | The training minimum contents are: | II. Hardware installation, assembly, and disassembly.
51 | The training minimum contents are: | III. Software installation and upgrade procedures.
52 | The training minimum contents are: | IV.Troubleshooting procedures.
53 | The training minimum contents are: | V.Preventive Maintenance and Calibration procedures.
54 | The training minimum contents are: | VI.Overview of consumables, supplies, and accessories.
55 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 92007 | 1154753 | VENTILATOR TRANSPORT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Adult to Infant (0.3 Kg – 250 Kg) or Vt(2ml to 2000ml)
2 | PARAMETER ACCURACY
3 | Tidal Volume | ≤ ±10%
4 | Inspiratory flow, L/min | ≤ ±10%
5 | Inspiration pressure, cmH20 | ≤ ± 2 cmH2O
6 | Respiratory rate, breaths/min | ≤ ± 2
7 | Inspiratory time, sec | ≤ ± 0.2 Sec
8 | Expiratory time, sec | ≤ ± 0.2 Sec
9 | Inspiratory Pause/Plateau time | ≤ ± 0.2 Sec
10 | FiO2% | ≤ ± 2 above 60% & below 60 ≤ ±5%
11 | PEEP/CPAP, cmH20 | ≤ ± 2 cmH2O
12 | Pressure support, cmH2O | ≤ ± 2 cmH2O
13 | Nebulizer | Pneumatic
14 | Trigger Mechanism | Pressure
15 | Bias/base flow range, L/min | 1 to 20
16 | Pressure slop/ramp adjustment | Yes
17 | Sigh breath function | Yes
18 | Control Panel Lock | Yes
19 | Additional Controls | · Automatic flow adjustment · ATP correction · BTPS correction · Flow resistance compensation
20 | EQUIPMENT ALARMS
21 | Gas supply failure
22 | Power Failure
23 | Vent Inoperative
24 | Low Battery
25 | Self-diagnostics and reference error code
26 | Alarm should be Visual and audible.
27 | Alarm Flashing indicator must be visible from any 360 degree of viewing angle
28 | Technical Error with appropriate error code
29 | MISCELLANEOUS INFORMATION
30 | External device interfaces
31 | Remote alarm/display port
32 | Reporting (Vent alarms and patient status) | Preferred
33 | View report on display
34 | HARDWARE SPECIFICATION
35 | Internal Turbine (Flow up to 2500ml/s)
36 | O2 sensor
37 | Integrated ETCO2 Module | Preferred
38 | Additional Distal end or Y sensor for infants flow measurement | Preferred
39 | Y sensor autoclave able | Preferred
40 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
41 | Weight of whole system must not be more than 10Kg without Cart
42 | Altitude : - 300m to 3000m from the sea level
43 | Dimension | Specify
44 | Temperature : 15 to 40 C | Preferred
45 | Cart must contain at least four wheel in which two wheels must have built-in brakes
46 | ELECTRICAL SPECIFICATION
47 | Ventilator must have dual voltage operational compatibility at automatic selection of frequency.
48 | Voltage Range : 100 ~ 240 Volts
49 | Frequency : 50Hz / 60Hz
50 | Maximum Current : ≤ 5 Amp
51 | Hot swappable batteries feature for external battery system
52 | Battery backup time must be 6 hrs. at-least
53 | PNEUMATIC SPECIFICATION
54 | Operating pressure : 2 to 6.5 Bar OR 35 to 85 PSI
55 | External Oxygen port to deliver demanded FiO2 (Ohmeda Type)
56 | SAFETY STANDARD
57 | The system must compliance with the following safety standards
58 | - ISO 80601-1-12
59 | - ISO 80601-1-55
60 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
61 | - IEC 60801-1 Type BF – for accessories, CO2 analyzer, Y sensor, nebulizer
62 | - IEC 60801-1 Type CF – Defibrillation proof
63 | - IPX1 | Preferred
64 | ALARMS
65 | Low/high FiO2
66 | Low minute volume
67 | High minute volume
68 | Low inspiratory pressure
69 | High pressure
70 | Loss of PEEP
71 | Apnea
72 | Continuous high pressure/occlusion
73 | Inverse IE ratio
74 | High respiratory rate
75 | High PEEP
76 | Autoset Feature
77 | Breathing circuit disconnect
78 | Gas supply failure
79 | Power failure
80 | Vent inoperative | Preferred
81 | Low battery
82 | Self-diagnostics
83 | FEATURES / OTHERS
84 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
85 | - The equipment or any of its accessories must not be presently listed in any recall
86 | - System must have minimum 72 Hrs. Trending
87 | - Technical log pool must have capacity of 1000 entries | Preferred
88 | - Audible Alarm volume can be adjusted : range 35dB to 90dB
89 | - Warning message to replace battery
90 | - Battery indicator on screen with expected battery time
91 | - Preventive maintenance must not recommend to replace any kind of PCB (It must be limited to filters, batteries, membranes and small pneumatic parts)
92 | - System must not require PM before 1 year or 5000 operating Hrs
93 | - Oxygen adjustable (21% to 100%)
94 | - High Flow oxygen Therapy
95 | - Leak Compensation
96 | - Non invasive ventilation
97 | Warranty | Warranty (Two Years) including Labor, spare parts and PPM kits
98 | TERMS & CONDITIONS | Power Plug type - set of American and British
99 | TERMS & CONDITIONS | Gas probe must be ohmeda/american standard.
100 | TERMS & CONDITIONS | Vendor to provide service manual
101 | TERMS & CONDITIONS | The vendor shall provide Local service training for clinical engineers.
102 | TERMS & CONDITIONS | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production
103 | TERMS & CONDITIONS | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
104 | TERMS & CONDITIONS | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
105 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 414268800 | 1155685 | ANESTHESIA MACHINE MRI COMPATIBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Pipe line gas inlets: | 3 (O2, N2O, air)
2 | Gas cylinder yokes: | O2, N2O
3 | Vaporizers/Agents, supported: | Isoflurane, halothane, enflurane, desflurane, sevoflurane
4 | Number and Type of Vaporizer on each Unit: | 2, Sevoflurane and Desoflurane
5 | Desoflurane Ready and Supported: | Yes
6 | Interlock Mechanism: | Yes
7 | SUCTION SYSTEM: | Yes (air ejector type is preferred)
8 | O2 FAIL-SAFE: | Audible, visual, N2O shutoff
9 | HYPOXIC MIXTURE FAIL-SAF: | Yes (methods vary, please specify)
10 | Auxiliary O2 outlet: | Yes
11 | Auxiliary common gas outlet | Yes
12 | AUTOMATIC VENTILATOR | Yes, with absorber.
13 | Ventilator Type: | Electronically controlled, electrically driven
14 | Operating Modes: | Volume controlled ventilation pressure controlled ventilation, Pressure support, SIMV/PS, Manual ventilation, Spontaneous ventilation
15 | Tidal volume Range cc | 20-1400 mL in Volume control, 20 to 1100 mL in SIMV/PS
16 | Minute volume range: | 0.05-26 l/min
17 | Breathing Frequency, bpm | 4 to 60 bpm
18 | Inspiratory flow, L/min | Please Specify
19 | IE ratio | Please specify
20 | Inspiratory pause: | Yes
21 | Pressure limit, cm H2O | 5 to 65 (adjustable)
22 | PEEP, cm H2O | 0 to 20 adjustable
23 | Integrated Safety Functions | Please specify
24 | Auto Compliance Compensation | Yes
25 | SCAVENGING SYSTEM | Open reservoir. passive
26 | Airway pressure monitoring | Yes, Inspiratory & Expiratory,
27 | Prioritized alarms system: | Advisory, caution, alarm
28 | High-pressure alarm: | Yes
29 | Sub atmospheric pressure alarm: | Yes
30 | Continuing pressure alarm: | Yes
31 | Operating time w/ fully charged batteries: | Greater than 45 min
32 | EQUIPMENT WARRANTY | The vendor is responsible to provide (2) years full warranty commencing from the date of preliminary acceptance certificate.
33 | GENERAL terms and conditions | 1.1 The request for proposal contains Anesthesia machine for Department of Anesthesiology.
34 | GENERAL terms and conditions | 1.2 The Bid must include all consumables and servicing costs; both regularly scheduled preventive maintenance as well as emergency maintenance as per manufacturer’s requirement, available on a 24hour basis. The vendor will cover the cost of all replacement parts and factory upgrades.
35 | GENERAL terms and conditions | 1.3 Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and that:
36 | GENERAL terms and conditions | · the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and
37 | GENERAL terms and conditions | · the equipment meets the standard of electromagnetic compatibility
38 | GENERAL terms and conditions | 1.4 The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as
39 | GENERAL terms and conditions | 1.5 SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
40 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.1 The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
41 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.2 The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order.
42 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.3 The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models.
43 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.4 Any components not proposed but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released with same price. To be discussed with contract dept.
44 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 3.1 Included with each vendor offer shall include all details of the training scope, location, details provided to department personnel.
45 | ARTICLE 3: TRAINING | 3.2 The hospital reserves the right not to accept the training and the vendor must provide a replacement at no additional cost to hospital.
46 | ARTICLE 3: TRAINING | 3.3 The vendor must provide two application training sessions: the first visit just after equipment preliminary acceptance and the second visit within six (6) months.
47 | ARTICLE 3: TRAINING | 3.4 The vendor will specify, with advance notification, required for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
48 | ARTICLE 3: TRAINING | 3.5 The vendor shall provide a factory-level technical service training for (1) clinical engineer at the factory or authorized training facility covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital covering, but not limited to, the following areas as applicable:
49 | ARTICLE 3: TRAINING | · Theory of Operation including circuit description and schematics
50 | ARTICLE 3: TRAINING | · Hardware installation, assembly, and disassembly
51 | ARTICLE 3: TRAINING | · Software installation and upgrade procedures
52 | ARTICLE 3: TRAINING | · Troubleshooting procedures
53 | ARTICLE 3: TRAINING | · Preventive Maintenance and Calibration procedures
54 | ARTICLE 3: TRAINING | · Overview of consumables, supplies, and accessories
55 | ARTICLE 3: TRAINING | 4.1 Vendors shall guarantee system upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the hospital.
56 | ARTICLE 4: SYSTEM’s UPGRADES | 5.1 In addition to the requested equipment, the vendor shall provide the QC test tools, manuals and/or software should be delivered at the time of equipment installation.
57 | ARTICLE 5: QC TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | 6.1 The vendor is responsible to provide (2) years full warranty commencing from the date of preliminary acceptance certificate.
58 | ARTICLE 6: EQUIPMENT WARRANTY | 7.1 Vendor personnel shall familiarize themselves with the features and characteristics of the building which will affect the delivery and/or installation of their equipment. The vendor shall check space, height, accessibility to all applicable areas to ensure ease of installation. Vendors must advise if the proposed equipment will NOT fit through a doorway.
59 | ARTICLE 7: EXAMINATION OF SPACE AND SERVICES (when applicable) | 7.2 The vendor shall report inadequacies to the owner immediately to facilitate remedial measures prior to scheduling installation.
60 | ARTICLE 7: EXAMINATION OF SPACE AND SERVICES (when applicable) | 7.3 The vendor shall accept responsibility for error or neglect to familiarize with access, space and service conditions. There will be no storage of furniture or equipment on the job site.
61 | ARTICLE 7: EXAMINATION OF SPACE AND SERVICES (when applicable) | 8.1 Delivery and installation shall be coordinated with hospital in accordance with contractual completion period
62 | ARTICLE 8: DELIVERY AND INSTALLATION | 8.2 The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
63 | ARTICLE 8: DELIVERY AND INSTALLATION | 8.3 The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc.) for each employee they will have on site so that appropriate hospital gate passes can be issued.
64 | ARTICLE 8: DELIVERY AND INSTALLATION | 9.1 The vendor shall keep the premises, including assigned rooms, all passages and auxiliary spaces free from accumulation of waste material or rubbish caused by him at the end of each day.
65 | ARTICLE 9: CLEANING UP AND MAKING GOOD DAMAGES | 9.2 The vendor shall assume full responsibility for all damages to the building plant and previously placed equipment, including gouges and smudges to the owner’s property made by their own personnel and/or sub-contractors.
66 | ARTICLE 9: CLEANING UP AND MAKING GOOD DAMAGES | 9.3 Any resulting repairs by the vendor shall match the adjacent existing finishes and materials. The repairs are to be done by appropriate trade contractors who did the work originally and paid by the vendor.
67 | ARTICLE 9: CLEANING UP AND MAKING GOOD DAMAGES | 9.4 The vendor shall comply with the required pre-infection control procedures.
68 | ARTICLE 9: CLEANING UP AND MAKING GOOD DAMAGES | 10.1 Two (2) complete operational manuals for hardware and software.
69 | ARTICLE 10: DOCUMENTATION | 10.2 Two (2) service manuals for hardware and software.
70 | ARTICLE 10: DOCUMENTATION | 10.3 Two (2) spare part manuals (if applicable).
71 | ARTICLE 10: DOCUMENTATION | 10.4 Completed schematic diagrams with documentation of onsite modifications (if required).
72 | ARTICLE 10: DOCUMENTATION | 10.5 All above documentation must be provided in English.
73 | ARTICLE 10: DOCUMENTATION | 11.1 The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
74 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | • Sales Literature.
75 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin.
76 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
77 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | • NEMA specifications and/or other international regulatory agencies.
78 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | 11.2 Any deficiencies identified upon preliminary acceptance testing must be rectified by the vendor to the hospital’s satisfaction prior to commencing preliminary acceptance.
79 | ARTICLE 11: EQUIPMENT ACCEPTANCE TESTING | Vendor must provide the following information as it applies to their service organization:
80 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.1 Local service is imperative.
81 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.2 Local service location.
82 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.3 Service response time for routine and after hour calls. The vendor shall define the term ‘Response Time’ and provide their minimum and maximum time limits.
83 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.4 For each piece of equipment specified within this document, list percentage of parts that are critical for normal operation of the equipment that are not available from the local service center, and would take more than one (1) day to receive from another area.
84 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.5 Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment.
85 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.6 The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced.
86 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.7 Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), the contract warranty periods stated above will be extended to cover 3 times the actual downtime.
87 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.8 Service response time for routine calls should be one (1) hour.
88 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.9 Service engineer must provide a service report each time they evaluate and/or work on the equipment to Clinical Engineering Department.
89 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.10 All empty boxes, tools, spare parts, defective parts, replacement parts and the like, will be removed from the premises immediately. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc
90 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 12.11 Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized and clutter-free).
91 | ARTICLE 12: MANUFACTURER’S SERVICE COMMITMENT | 13.1 The vendor is requested to fill the technical specification sheets -. Not providing data requested can delay the vendor selection process wherein the hospital does reserve the right to ask a vendor to recompile and re submit the data.
92 | ARTICLE 13: GENERAL COMMENTS | 13.2 If data is not available for a certain requested data field the vendor should not leave the field blank but instead use the term ‘No Data’ in the field. If this term is not used, it will be assumed that the data sheet is incomplete and a resubmission may be requested. The vendor does have the right to reference the field to an attachment to better clarify the entry.
93 | ARTICLE 13: GENERAL COMMENTS | 13.3 The vendor must provide sufficient manufacture documentation to support the data quoted.
94 | ARTICLE 13: GENERAL COMMENTS | 13.4 The vendor should provide additional information that may not have been requested on the data specification sheets that can elaborate on the quoted technology capabilities and clinical applications.
95 | ARTICLE 14: POWER REQUIRMENTS: | 14.1 220V British style cable Hospital grade minimum 5m in length.
96 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 411633102 | 1156343 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 5.5X65 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Direction of View | 0 Deg
2 | Angle of View | 100 Deg
3 | Working Length | 65 cm
4 | Total Length | 94 cm
5 | Working Channel inner diameter | 2.1 mm
6 | Distal tip outer diameter | 5.5 mm
7 | Deflection up/down | 140 Deg/140 Deg
8 | High image resolution and video imaging | Yes
9 | EQUIPMENT WARRANTY | The vendor is responsible to provide (1) year full warranty commencing from the date of preliminary acceptance certificate. Vendor or supplier will guarantee the availability of a loaner unit in Riyadh whenever requested by Hospital during the repair of any defective equipment
10 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 411633101 | 1156343 | FLEXIBLE INTUBATION VIDEO ENDOSCOPE 4X65 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Direction of View | 0 Deg
2 | Angle of View | 100 Deg
3 | Working Length | 65 cm
4 | Total Length | 94 cm
5 | Working Channel inner diameter | 1.5 mm
6 | Distal tip outer diameter | 4.1 mm
7 | Deflection up/down | 140 Deg/140 Deg
8 | High image resolution and video imaging | Yes
9 | Endotracheal Tubes good gliding properties | Yes
10 | Manual/Machine disinfection | Yes
11 | Suitable for low temp reprocessing methods | Pls specify
12 | EQUIPMENT WARRANTY | The vendor is responsible to provide (1) year full warranty commencing from the date of preliminary acceptance certificate. Vendor or supplier will guarantee the availability of a loaner unit in Riyadh whenever requested by Hospital during the repair of any defective equipment
13 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 411605301 | 1157157 | ICU VENTILATOR MRI COMPATIBLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Adult to Infant (0.5 Kg – 250 Kg) or Vt(20ml to 2000ml)
2 | QUANTITY | As Requested
3 | PARAMETER ACCURACY
4 | Tidal Volume | ≤ ±10%
5 | Inspiratory flow, L/min | ≤ ±10%
6 | Inspiration pressure, cmH20 | ≤ ± 2 cmH2O
7 | Respiratory rate, breaths/min | ≤ ± 2
8 | Inspiratory time, sec | ≤ ± 0.2 Sec
9 | Expiratory time, sec | ≤ ± 0.2 Sec
10 | Inspiratory Pause/Plateau time | ≤ ± 0.2 Sec
11 | FiO2% | ≤ ± 2 above 60% & below 60 ≤ ±5%
12 | PEEP/CPAP, cmH20 | ≤ ± 2 cmH2O
13 | Pressure support, cmH2O | ≤ ± 2 cmH2O
14 | Nebulizer | Pneumatic
15 | Trigger Mechanism | Pressure
16 | Bias/base flow range, L/min | 1 to 20
17 | Pressure slop/ramp adjustment | Yes
18 | Sigh breath function | Yes
19 | Control Panel Lock | Yes
20 | Additional Controls | · Automatic flow adjustment · ATP correction · BTPS correction · Flow resistance compensation
21 | EQUIPMENT ALARMS
22 | Gas supply failure
23 | Power Failure
24 | Vent Inoperative
25 | Low Battery
26 | Self-diagnostics and reference error code
27 | Alarm should be Visual and audible.
28 | Alarm Flashing indicator must be visible from any 360 degree of viewing angle
29 | Technical Error with appropriate error code
30 | MISCELLANEOUS INFORMATION
31 | External device interfaces
32 | Remote alarm/display port
33 | Reporting (Vent alarms and patient status) | Preferred
34 | View report on display
35 | HARDWARE SPECIFICATION
36 | Internal Turbine (Flow up to 2500ml/s)
37 | O2 sensor
38 | Integrated ETCO2 Module | Preferred
39 | Proximal Flow Sensor
40 | Y sensor autoclave able
41 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
42 | Weight of whole system must not be more than 10Kg without Cart
43 | Altitude : - 300m to 3000m from the sea level
44 | Dimension | Specify
45 | Temperature : 15 to 40 C | Preferred
46 | Cart must contain at least four wheel in which two wheels must have built-in brakes
47 | MRI Compatibility
48 | MRI Conditional upto 1 meter for 3 T static magnetic field
49 | ELECTRICAL SPECIFICATION
50 | Ventilator must have dual voltage operational compatibility at automatic selection of frequency.
51 | Voltage Range : 100 ~ 240 Volts
52 | Frequency : 50Hz / 60Hz
53 | Maximum Current : ≤ 5 Amp
54 | Hot swap-able batteries
55 | Battery backup time must be 6 hrs. at-least
56 | PNEUMATIC SPECIFICATION
57 | Operating pressure : 2 to 6.5 Bar OR 35 to 85 PSI
58 | External Oxygen port to deliver demanded FiO2 (Ohmeda Type)
59 | SAFETY STANDARD
60 | The system must compliance with the following safety standards
61 | - ISO 80601-1-12
62 | - ISO 80601-1-55
63 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
64 | - IEC 60801-1 Type BF – for accessories, CO2 analyzer, Y sensor, nebulizer
65 | - IEC 60801-1 Type CF – Defibrillation proof
66 | - IPX1 | Preferred
67 | ALARMS
68 | Alarm Silence Manual Activation / Can be reset to Active Alarm
69 | Low/high FiO2
70 | Low minute volume
71 | High minute volume
72 | Low inspiratory pressure
73 | High pressure
74 | Loss of PEEP
75 | Apnea
76 | Continuous high pressure/occlusion
77 | Inverse IE ratio
78 | High respiratory rate
79 | High PEEP
80 | Autoset Feature
81 | Breathing circuit disconnect
82 | Gas supply failure
83 | Power failure
84 | Vent inoperative | Preferred
85 | Low battery
86 | Self-diagnostics
87 | FEATURES / OTHERS
88 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
89 | - The equipment or any of its accessories must not be presently listed in any recall
90 | - System must have minimum 72 Hrs. Trending
91 | - Technical log pool must have capacity of 1000 entries | Preferred
92 | - Audible Alarm volume can be adjusted : range 35dB to 90dB
93 | - Warning message to replace battery
94 | - Battery indicator on screen with expected battery time
95 | - Preventive maintenance must not recommend to replace any kind of PCB (It must be limited to filters, batteries, membranes and small pneumatic parts)
96 | - System must not require PM before 1 year or 5000 operating Hrs
97 | - Oxygen adjustable (21% to 100%)
98 | - High Flow oxygen Therapy
99 | - Leak Compensation
100 | - Non invasive ventilation
101 | Warranty | Warranty (Two Years) including Labor, spare parts and PPM kits
102 | TERMS & CONDITIONS | - Power Plug type - set of American and British
103 | TERMS & CONDITIONS | - Gas probe must be ohmeda/american standard.
104 | TERMS & CONDITIONS | - Vendor to provide service manual
105 | TERMS & CONDITIONS | - The vendor shall provide Factory level service training for ONE service engineer & Clinical factory level training for TWO Respiratory Therapists.
106 | TERMS & CONDITIONS | - The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production
107 | TERMS & CONDITIONS | - The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
108 | TERMS & CONDITIONS | - The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
109 | - Authority Clarification | SFDA Required
110 | CLINICAL SPECIFICATION | 0.10-4.0 or 5.0 sec. minimum range, To accommodate all patients sizes / needs
111 | CLINICAL SPECIFICATION | Adjustable Ti Ped through Adult (0.5 – 250kg), Ability
112 | CLINICAL SPECIFICATION | Patient Range\ Tidal Volume: to SELECT Range preferred (w/ associated safety features) 20 ml – 2500 ml Minimum Range
113 | CLINICAL SPECIFICATION | Internal Exhalation Valve Does not require circuits with exhalation valves (Disposable, balloon-valve systems not acceptable)
114 | Flow-Sensitivity | Ped through Adult, including “Y” sensor (or other similar external flow/volume sensor) for smaller targeted volumes. To accommodate smaller patients, provide maximum responsiveness to patient efforts (by Range, preferable).
115 | PEEP | 0 – 40 cmH2O Minimum Range
116 | Pressure | Minimum Range 0 – 80 cmH2O
117 | Respiratory Rate | 0 – 100 bpm Minimum Range
118 | NIV Leak Compensation | ≥ 60 LPM Minimum Range
119 | Inspiratory / Expiratory Hold | Manual activation, Time-limited
120 | Desired Modes | If not same designation, should provide similar function to:
121 | Desired Modes | Synchronized Intermittent Mandatory Ventilation – Volume Control
122 | Desired Modes | Synchronized Intermittent Mandatory Ventilation – Pressure Control
123 | Desired Modes | Pressure Regulated Volume Controlled/Targeted,
124 | Desired Modes | Pressure Control Ventilation
125 | Desired Modes | Volume Control Ventilation
126 | Desired Modes | Control Ventilation with Automode with combined Support Ventilation
127 | Desired Modes | Pressure Support Ventilation
128 | Desired Modes | Continuous Positive Airway Pressure
129 | Desired Modes | Bi-level / Airway Pressure-Release Ventilation
130 | Desired Modes | Apnea/Back Up Ventilation
131 | Desired Modes | Non-Invasive: As a selected mode, similar to patient selection / separate from actual mode selection
132 | Desired Modes | O2 Therapy: High Flow Oxygen Therapy for Adult and Pediatrics. Min Flow Range: 2-60 LPM, FiO2 0.21-1.0
133 | Desired Modes | Standby: Ability to detect patient connection and resume ventilation required.
134 | Start / Single Breath: | Activates single breath to current set parameters
135 | End-Tidal CO2 | Infant/Pediatric/Adult, modular-preferred, Mainstream type
136 | Flow Patterns: | Constant, Descending, Accelerating (Decelerating – mandatory)
137 | Monitoring | Peak / Pause / Plateau Pressure / Mean Airway Pressure
138 | Monitoring | PEEP (measured / total) and Auto PEEP
139 | Monitoring | I:E Ratio (measured / set)
140 | Monitoring | Patient Effort Indication
141 | Monitoring | Oxygen Concentration (inspired)
142 | Monitoring | End-Tidal CO2
143 | Monitoring | Resistance / Compliance (static & dynamic)
144 | Monitoring | Tidal Volume (Inspired / Expired)
145 | Monitoring | Minute Volume (Inspired / Expired)
146 | Monitoring | Respiratory Rate
147 | Monitoring | Tidal volume /kg
148 | Flow-Slope | >120 lpm, Flow-Slope adjustable
149 | Accessories | Include external F-P Heater/Humidifier with dual heated wire per ventilator
150 | Accessories | Include Support Arm
151 | Accessories | Include 10 Adult disposable dual heated wire circuits per ventilator, approved by KFSHRC
152 | Accessories | Include 10 Pediatric disposable dual heated wire circuits per ventilator, approved by KFSHRC
153 | Accessories | Include 2 Adult reusable EtCO2 Airway Adapters per ventilator, approved by KFSHRC
154 | Accessories | Include 2 Pediatric reusable EtCO2 Airway Adapters per ventilator
155 | Accessories | Consumable supplies are not to be proprietary to one manufacturer for items such as filters and ventilator circuits
156 | Accessories | arm support,
157 | Accessories | Humidifier holder,
158 | Service / Communications: | Software-Upgradable, Serial Port (RS-232), Data Output (USB-Like) for transferring data
159 | Service / Communications: | See enclosed further documents regarding device connectivity to HIS and requirements
160 | Service / Communications: | Must be 100% full integration to Cerner systems
161 | Physical | Patient Unit Easily Detachable from the cart, cart included with O2 cylinder holder.
162 | Physical | Appropriate handles to fit KFSHRC beds to hang the ventilator
163 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | 411605203 | 1157159 | PORTABLE VENTILATOR WITH TURBINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Adult to Infant (0.3 Kg – 250 Kg) or Vt(2ml to 2000ml)
2 | PARAMETER ACCURACY
3 | Tidal Volume | ≤ ±10%
4 | Inspiratory flow, L/min | ≤ ±10%
5 | Inspiration pressure, cmH20 | ≤ ± 2 cmH2O
6 | Respiratory rate, breaths/min | ≤ ± 2
7 | Inspiratory time, sec | ≤ ± 0.2 Sec
8 | Expiratory time, sec | ≤ ± 0.2 Sec
9 | Inspiratory Pause/Plateau time | ≤ ± 0.2 Sec
10 | FiO2% | ≤ ± 2 above 60% & below 60 ≤ ±5%
11 | PEEP/CPAP, cmH20 | ≤ ± 2 cmH2O
12 | Pressure support, cmH2O | ≤ ± 2 cmH2O
13 | Nebulizer | Pneumatic
14 | Trigger Mechanism | Pressure
15 | Bias/base flow range, L/min | 1 to 20
16 | Pressure slop/ramp adjustment | Yes
17 | Sigh breath function | Yes
18 | Control Panel Lock | Yes
19 | Additional Controls | • Automatic flow adjustment • ATP correction • BTPS correction • Flow resistance compensation
20 | EQUIPMENT ALARMS
21 | Gas supply failure
22 | Power Failure
23 | Vent Inoperative
24 | Low Battery
25 | Self-diagnostics and reference error code
26 | Alarm should be Visual and audible.
27 | Alarm Flashing indicator must be visible from any 360 degree of viewing angle
28 | Technical Error with appropriate error code
29 | MISCELLANEOUS INFORMATION
30 | External device interfaces
31 | Remote alarm/display port
32 | Reporting (Vent alarms and patient status) | Preferred
33 | View report on display
34 | HARDWARE SPECIFICATION
35 | Internal Turbine (Flow up to 2500ml/s)
36 | O2 sensor
37 | Integrated ETCO2 Module | Preferred
38 | Additional Distal end or Y sensor for infants flow measurement | Preferred
39 | Y sensor autoclave able | Preferred
40 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
41 | Weight of whole system must not be more than 10Kg without Cart
42 | Altitude : - 300m to 3000m from the sea level
43 | Dimension | Specify
44 | Temperature : 15 to 40 C | Preferred
45 | Cart must contain at least four wheel in which two wheels must have built-in brakes
46 | ELECTRICAL SPECIFICATION
47 | Ventilator must have dual voltage operational compatibility at automatic selection of frequency.
48 | Voltage Range : 100 ~ 240 Volts
49 | Frequency : 50Hz / 60Hz
50 | Maximum Current : ≤ 5 Amp
51 | Hot swappable batteries feature for external battery system
52 | Battery backup time must be 6 hrs. at-least
53 | PNEUMATIC SPECIFICATION
54 | Operating pressure : 2 to 6.5 Bar OR 35 to 85 PSI
55 | External Oxygen port to deliver demanded FiO2 (Ohmeda Type)
56 | SAFETY STANDARD
57 | The system must compliance with the following safety standards
58 | - ISO 80601-1-12
59 | - ISO 80601-1-55
60 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
61 | - IEC 60801-1 Type BF – for accessories, CO2 analyzer, Y sensor, nebulizer
62 | - IEC 60801-1 Type CF – Defibrillation proof
63 | - IPX1 | Preferred
64 | ALARMS
65 | Low/high FiO2
66 | Low minute volume
67 | High minute volume
68 | Low inspiratory pressure
69 | High pressure
70 | Loss of PEEP
71 | Apnea
72 | Continuous high pressure/occlusion
73 | Inverse IE ratio
74 | High respiratory rate
75 | High PEEP
76 | Autoset Feature
77 | Breathing circuit disconnect
78 | Gas supply failure
79 | Power failure
80 | Vent inoperative | Preferred
81 | Low battery
82 | Self-diagnostics
83 | FEATURES / OTHERS
84 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
85 | - The equipment or any of its accessories must not be presently listed in any recall
86 | - System must have minimum 72 Hrs. Trending
87 | - Technical log pool must have capacity of 1000 entries | Preferred
88 | - Audible Alarm volume can be adjusted : range 35dB to 90dB
89 | - Warning message to replace battery
90 | - Battery indicator on screen with expected battery time
91 | - Preventive maintenance must not recommend to replace any kind of PCB (It must be limited to filters, batteries, membranes and small pneumatic parts)
92 | - System must not require PM before 1 year or 5000 operating Hrs
93 | - Oxygen adjustable (21% to 100%)
94 | - High Flow oxygen Therapy
95 | - Leak Compensation
96 | - Non invasive ventilation
97 | Warranty | Warranty (Two Years) including Labor, spare parts and PPM kits
98 | TERMS & CONDITIONS | - Warranty (Two Years) including Labor, spare parts and PPM kits
99 | TERMS & CONDITIONS | - Power Plug type - set of American and British
100 | TERMS & CONDITIONS | - Gas probe must be ohmeda/american standard.
101 | TERMS & CONDITIONS | - Vendor to provide service manual
102 | TERMS & CONDITIONS | - The vendor shall provide Local service training for ONE service engineers
103 | TERMS & CONDITIONS | - The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production
104 | TERMS & CONDITIONS | - The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
105 | TERMS & CONDITIONS | - The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
106 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | 411605311 | 1157796 | VENTILATOR ICU NEONATAL with HFOV |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Patient Type | Neonatal to Infant (0.5 Kg – 25 Kg) or 2ml to 300ml
2 | PARAMETER ACCURACY
3 | Tidal Volume | ≤ ±10%
4 | Inspiratory flow, L/min | ≤ ±10%
5 | Inspiration pressure, cmH20 | ≤ ± 2 cmH2O
6 | Respiratory rate, breaths/min | ≤ ± 2
7 | Inspiratory time, sec | ≤ ± 0.2 Sec
8 | Expiratory time, sec | ≤ ± 0.2 Sec
9 | Inspiratory Pause/Plateau time | ≤ ± 0.2 Sec
10 | FiO2% | ≤ ± 2 above 60% & below 60 ≤ ±5%
11 | PEEP/CPAP, cmH20 | ≤ ± 2 cmH2O
12 | Pressure support, cmH2O | ≤ ± 2 cmH2O
13 | Nebulizer | Ultrasonic
14 | Trigger Mechanism | Pressure, Flow
15 | Bias/base flow range, L/min | 1 to 20
16 | Pressure slop/ramp adjustment | Yes
17 | Sigh breath function | Yes
18 | Control Panel Lock | Yes
19 | Additional Controls | • Automatic flow adjustment • ATP correction • BTPS correction • Flow resistance compensation
20 | EQUIPMENT ALARMS
21 | Gas supply failure
22 | Power Failure
23 | Vent Inoperative
24 | Low Battery
25 | Self-diagnostics and reference error code
26 | Alarm should be Visual and audible.
27 | Alarm Flashing indicator must be visible from any 360 degree of viewing angle
28 | Technical Error with appropriate error code
29 | MISCELLANEOUS INFORMATION
30 | External device interfaces
31 | Remote alarm/display port | Preferred
32 | Reporting (Vent alarms and patient status) | Preferred
33 | View report on display
34 | send reports to printer | Preferred
35 | Save data to external media
36 | Send data via network
37 | Non invasive ventilation therapy
38 | O2 therapy
39 | High Frequency Oscillation
40 | Frequency - 2 - 20Hz
41 | Automatic tube compensation
42 | Leakage Compensation
43 | DISPLAY
44 | Data displayed should be user customization
45 | Configurable display screen
46 | Display Size must be 15” or more
47 | HARDWARE SPECIFICATION
48 | O2 sensor (life not lesser than 2 years)
49 | flow transducer
50 | Integrated ETCO2 Module
51 | Built-in Or integrated Ultrasonic Nebulizer
52 | Additional Distal end or Y sensor for infants flow measurement
53 | Y sensor autoclave able
54 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION
55 | Modular System | Preferred
56 | Weight of whole system must not be more than 30Kg without Cart
57 | User Interface should be less than 10Kg
58 | Breathing Delivery Unit less than 25Kg
59 | Altitude : - 300m to 3000m from the sea level
60 | Screen Size 15” or more
61 | Touch Screen interface
62 | Dimension | Specify
63 | Temperature : 15 to 35 C | Preferred
64 | Cart must contain at least four wheel in which two wheels must have built-in brakes
65 | ELECTRICAL SPECIFICATION
66 | Ventilator must have dual voltage operational compatibility at automatic selection of frequency.
67 | Voltage Range : 100 ~ 240 Volts, 110 is acceptable
68 | Frequency : 50Hz / 60Hz
69 | Maximum Current : ≤ 5 Amp
70 | Hot swappable batteries feature | Preferred
71 | Battery backup time must be 1 hrs. at-least
72 | System must have Extended Battery backup feature
73 | Battery type should be Non Spill-able (medical grade), SLA is not acceptable
74 | PNEUMATIC SPECIFICATION
75 | Operating pressure : 2 to 6.5 Bar OR 35 to 85 PSI
76 | Hoses for gases and connectors must compliance with American standard (Ohmeda Type)
77 | SAFETY STANDARD
78 | The system must compliance with the following safety standards
79 | - ISO 80601-1-12
80 | - ISO 80601-1-55
81 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient
82 | - IEC 60801-1 Type BF – for accessories, CO2 analyzer, Y sensor, nebulizer
83 | - IEC 60801-1 Type CF – Defibrillation proof
84 | - IPX1 | Preferred
85 | ALARMS
86 | Low/high FiO2
87 | Low minute volume
88 | High minute volume
89 | Low inspiratory pressure
90 | High pressure
91 | Loss of PEEP
92 | Apnea
93 | Continuous high pressure/occlusion
94 | Inverse IE ratio
95 | High respiratory rate
96 | High PEEP
97 | Autoset Feature
98 | Breathing circuit disconnect
99 | Gas supply failure
100 | Power failure
101 | Vent inoperative
102 | Low battery
103 | Self-diagnostics
104 | OTHERS
105 | - Hospital Engineer can update/upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC.
106 | - The equipment or any of its accessories must not be presently listed in any recall
107 | - System must have minimum 72 Hrs. Trending
108 | - Technical log pool must have capacity of 1000 entries | Preferred
109 | - Audible Alarm volume can be adjusted : range 35dB to 90dB
110 | - Warning message to replace battery
111 | - Battery indicator on screen with expected battery time
112 | - Preventive maintenance must not recommend to replace any kind of PCB (It must be limited to filters, batteries, membranes and small pneumatic parts)
113 | - System must not require PM before 1 year or 5000 operating Hrs
114 | - Hard drive or memory device which keep the operating system, log files, licenses etc must be solid state semiconductor.
115 | - Secondary Display to view and control ventilator from outside the isolation room | preferred
116 | MEDICAL AIR COMPRESSOR
117 | Integrated | With offered Ventilator
118 | Pressure | upto 4 Bar
119 | Safety Relief Valve
120 | Battery operated
121 | Weight | ≤ 15Kg
122 | Heat Exchange | To provide dry air
123 | Autostart | When Needed
124 | Voltage | 100 – 250 V Auto Or Ventilator Powered
125 | Fit on the Ventilator Cart
126 | Power Consumption | ≤ 300 W
127 | Humidifier
128 | Voltage | 200 – 250 V
129 | Power | 150W max
130 | Supply Frequency | 60Hz
131 | Modes | Invasive & Non Invasive
132 | Temperature control Ranges | 35 – 40 C Airway Invasive Mode 28 – 34 C Non Invasive Mode
133 | Display
134 | Over Temperature Cutout | @ 120 ± 5 C
135 | Flow Measurement and Alarm
136 | High Humidity Alarm | Airway temp. exceed 43 Celsius
137 | Low Humidity Alarm | @34.5 C
138 | Sound Level | Above 50dBA
139 | Chamber Temp. measurement
140 | Airway Temp. measurement
141 | Warm up Time | Less than 30 Minutes
142 | warranty | The vendor is responsible to provide a minimum of THREE (3) year full warranty commencing from the date of first case.
143 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of THREE (3) year full warranty commencing from the date of first case.
144 | 1.01 Warranty and Technical Support | b) Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
145 | 1.01 Warranty and Technical Support | c) Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
146 | 1.01 Warranty and Technical Support | d) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
147 | 1.01 Warranty and Technical Support | e) The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
148 | 1.01 Warranty and Technical Support | f) The contractor will provide all the essential tools and test equipment required to maintain the equipment free of charge.
149 | 1.01 Warranty and Technical Support | g) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
150 | 1.02 Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
151 | 1.02 Software and upgrades
152 | 1.02 Software and upgrades | b) If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
153 | 1.03 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
154 | 1.03 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
155 | 1.03 Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
156 | 1.04 Power Requirement | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V)
157 | 1.04 Power Requirement | b) The Frequency must be 60Hz.
158 | 1.04 Power Requirement | c) Power plug should be American style medical grade.
159 | 1.05 Gas Requirement | a) Medical equipment to be used in the hospital will comply with following color code for gas hoses:
160 | 1.05 Gas Requirement | · Oxygen hose GREEN
161 | 1.05 Gas Requirement | · Medical Air hose YELLOW
162 | 1.05 Gas Requirement | b) Ventilator should be capable to work with gas Pressure: 30 to 70 PSI
163 | 1.05 Gas Requirement | c) Medical Gas probes for all gases should be American standard OHMEDA style.
164 | 1.06 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
165 | 1.06 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
166 | 1.07 Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
167 | 1.07 Training | b) The vendor shall provide factory-level service training on each modality for TWO service engineer, at the factory or authorized training facility covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation, with no additional cost to the Hospital. The training minimum contents are:
168 | 1.07 Training | I. Theory of Operation including circuit description and schematics.
169 | 1.07 Training | II. Hardware installation, assembly, and dis assembly.
170 | 1.07 Training | III. Software installation and upgrade procedures.
171 | 1.07 Training | IV. Troubleshooting procedures.
172 | 1.07 Training | V. Preventive Maintenance and Calibration procedures.
173 | 1.07 Training | VI. Overview of consumables, supplies, and accessories.
174 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | 411531400 | 1160697 | AMBULATORY HOLTER MONITORING SYSTEM WITH 10 RECORDERS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | pc workstation with interface sofware | QTY (1)
2 | recorder module | QTY (10)
3 | 1. Recorders, Electrocardiography
4 | Recorder PATIENT TYPE | Adult, pediatric
5 | Recorder Configuration | Retrospective, real-time, or combination is preferred (please identify).
6 | RECORDER ANNOTATION
7 | Entry of patient ID, | Yes (Key panel or via PC)
8 | Patient (event marker) | Yes
9 | Time | Yes
10 | Arrhythmia detection | Yes (For real-time systems)
11 | Pacemaker detection | Yes
12 | RECORDER Type | Digital
13 | No. of patient lead wires | 7-lead patient cable with color-coded leads
14 | Recorder Modes (can record 3 or 12 leads with same recorder) | Yes
15 | Recorder Medium | PCMCIA flashcard / SD
16 | Recorder Capacity, hr. | 24 or 48 Please specify
17 | Recorder Memory size, GB | Not less than 32
18 | Recorder Frequency resp, Hz | 0.05 to 50 Hz
19 | Recorder Sampling rate, samples/sec | Not less than 128
20 | Resolution, bits | Not less than 10
21 | Recorder Gain/sensitivity, mm/mV | Auto
22 | Calibration signal | Yes
23 | Recorder Signal output jack | Preferred
24 | Real-time clock (for event marking) | Yes
25 | No. of ECG channels | 2020-03-12
26 | FULL DISCLOSURE | Yes
27 | SOFTWARE UPGRADEABLE | Yes
28 | BATTERY Type (number) | Alkaline Any Standard Commonly used Type (please specify)
29 | Recorder Carrying case/pouch | Included
30 | built-in patient cable tester(built-in impedance check with lead quality indicator) | Yes
31 | low-battery warning | Yes
32 | Recorder Splash Proof Body | Yes
33 | Recorder Meets requirements of AAMI EC-11, IEC 601-1, ISO 9001, and UL 544. | Yes
34 | 2. Scanners, Electrocardiography
35 | HARDWARE/SOFTWARE | Both
36 | SYSTEM TYPE | Prospective, retrospective
37 | CHANNELS ANALYZED | Minimum 2 simultaneous or individual
38 | MAXIMUM NO. OF RECORDERS FOR ANALYSIS AND STORAGE | Unlimited
39 | SCAN DISPLAY Scanning modes | Automatic or manual
40 | SCAN DISPLAY Time | Yes
41 | Scan Display Heart rate | Low, mean, high
42 | FULL DISCLOSURE | Yes
43 | PACEMAKER ANALYSIS | Yes
44 | TOTAL BEATS/BEATS,PER HR All types, Normal, Ventricular and Supraventricular | Yes
45 | ARRHYTHMIA Detection Indicator | Yes
46 | ST-segment analysis | Yes
47 | Arrhythmia No. of classifications | Not less than 10
48 | PLAYBACK TYPE | Digital
49 | PRINTER | Yes, High Speed B/W Laser Printer with two trays
50 | PC COMPATIBLE | Yes
51 | UPGRADEABLE | Yes
52 | Processing Unit: | I7, 32 GB RAM 1 TB SSD HD plus one additional 6 TB HD for storage (or latest processor at the time of shipping).
53 | Operating System: | Windows 10 Ent / Pro 64 bit
54 | Download/scan Time: | 90 Seconds or less
55 | Long - Term Archive: | DVD recorder with built-in HD
56 | Interface to the Hospital Information System / Clinical Information System HIS at no additional cost to the hospital | Yes
57 | Extensive Holter analysis and report generation software | Included, full licensed unlimited access
58 | TREND GRAPHS R-R interval | Yes
59 | Trend Graphs ST-segment deviation | Yes
60 | Trend Graphs Ventricular beats | Yes
61 | REPORT Strip annotation | Time, patient ID, HR, strip annotation (e.g., arrhythmia annotate)
62 | STRIP EXPANSION ON FD REPORT | Yes
63 | DATA MANAGEMENT | Preferred
64 | HL7 interfaces | Yes
65 | Meets requirements of AAMI EC-11, ISO 9001, and UL 544. | Yes
66 | warranty | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.
67 | Terms and Conditions | 1-The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.
68 | Terms and Conditions | 2-Parts and software updates will be done F.O.C. during warranty period.
69 | Terms and Conditions | 3-The unit should be provided with a line (power) cord of acceptable durability, quality, length, and capacity and should be secured with adequate strain reliefs.
70 | Terms and Conditions | 4-Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance.
71 | Terms and Conditions | 5-The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recalls.
72 | Terms and Conditions | 6-The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period)
73 | Terms and Conditions | Full local technical service training to be provided by certified service engineer at the time of installation (or at a mutually agreed upon time free of charge) covering the following areas but not limited to:
74 | Terms and Conditions | a) theory of operation
75 | Terms and Conditions | b) familiarization with the device
76 | Terms and Conditions | c) operational procedures and operating controls
77 | Terms and Conditions | d) operational calibration procedures (if applicable)
78 | Terms and Conditions | e) identification of consumables and supplies (if applicable)
79 | Terms and Conditions | f) Software and servicing the software
80 | Terms and Conditions | 7-The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
81 | Terms and Conditions | 8-Demo and or loaner equipment might be checked and managed while at hospital before final purchase.
82 | Terms and Conditions | 9-Provide software revisions, corrections, and updates that affect the present operational functions of the equipment for the life of the equipment free of charge.
83 | Terms and Conditions | 10-The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by the hospital.
84 | Terms and Conditions | 11-The user in service training must be provided by a qualified clinical instructors who is not a sales personnel.
85 | Terms and Conditions | 12-Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required.
86 | Terms and Conditions | 13-The vendor shall provide a factory level training (for one service engineers/technicians) at the factory or authorized training facility covering all expenses (round trip tickets, accommodation, registration fees, transportation, medical insurance, and food….etc.). The detailed training contents should be submitted and agreed upon in advance.
87 | Terms and Conditions | 14-All items on quotation are itemized and individually priced. Discounts / VAT are listed. Quotation should include accessories, pm kits, service kit, calibration tools, Simulators and setup and consumables that can run the system for at least three months.
88 | Terms and Conditions | 15-A delivery schedule is provided with the quotation. Delivery is FOB hospital, paid for by the vendor.
89 | Terms and Conditions | 16-List of reference sites both locally and internationally should be submitted with the details of work performed.
90 | Terms and Conditions | 17- All quoted data specification should be supported with original manufacture certification and technical data sheet.
91 | Terms and Conditions | 18- power plug should be 220 v UK style hospital grade
92 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | 105139 | 1163670 | TRACTION UNIT PHYSICAL THERAPY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | A universal traction unit for lumbar and cervical tractions with electric height adjustment table/couch. | YES
2 | Traction unit:
3 | microprocessor controlled traction unit | YES
4 | LCD, digital control panel | YES
5 | Selected parameter: traction force, base force, base hold time, and treatment time | YES
6 | The traction force shall be electronically measured and constantly monitored | YES
7 | Mechanical maximum force limitation | YES
8 | Patient stop switch | YES
9 | Patient movement shall not influence the traction force | YES
10 | The traction unit shall have the following specifications:
11 | Types of treatments: static traction and intermittent traction | YES
12 | Traction force | YES,1.5-90 Kg OR better
13 | Treatment time | YES, 0-60 minutes with acoustic signal and automatic signal and automatic reducing of traction force.
14 | Intermittent traction | YES
15 | Traction hold time | YES,should be: 0-90 seconds
16 | Base hold time setting | YES,should be: 0-90 seconds
17 | Unit to be mounted on the traction table | YES
18 | Traction table/couch:
19 | 4-section, electric height adjustment table fully equipped for cervical and lumbar tractions. | YES
20 | Table shall offer both sitting and lying treatment positions | YES
21 | Table surface to prevent friction to hip and legs during treatment | YES
22 | Adjustable head rest | YES
23 | Breathing hole and plug | YES
24 | Easy to move on casters with braking system | YES
25 | Durable, hygienic, and washable upholstery | YES
26 | Complete with fixation rails for fitting thoracic and pelvic belts | YES
27 | Height adjustment SPEED | YES, 25 seconds from 450 to 950 mm, OR better
28 | Lifting capacity | YES, 135 kg or more
29 | It should completely supplied with standard accessories and to be listed | YES
30 | Additional accessories:
31 | Pulley | YES
32 | Mobile traction frame for cervical traction | YES
33 | Carabine hook, galvanized, oval, diameter 70mm | YES
34 | Flexi stool with separate height adjustment for each leg | YES
35 | Traction frame for the manumed special H/L (W: 340mm) | YES
36 | Thoracic and pelvic fixation belts with back-to-back fastening, canvas | YES
37 | Padded neck harness, complete with spreader bar | YES
38 | Clamp for straps (2 required) | YES
39 | warranty | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of Technical acceptance.
40 | Terms and Conditions | 1-The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of Technical acceptance.
41 | Terms and Conditions | 2-The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
42 | Terms and Conditions | 3-The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment .
43 | Terms and Conditions | 4-The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as,SFDA ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
44 | Terms and Conditions | 5-Voltage: 220V , 60 Hz, UK plug hospital grade
45 | Terms and Conditions | 6-The vendor to provide installation, set up, calibration, performance verification, preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
46 | Terms and Conditions | 7-One (1) complete operations manual, to be provided in both hard and soft copies if available.
47 | Terms and Conditions | 8-One (1) complete service manual (in English language) including schematics and spare parts list, (if applicable), the manual is to be provided in both hard and soft copies if available.
48 | Terms and Conditions | 9-Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
49 | Terms and Conditions | 10-Full local technical service training to be provided, by a certified trainer, at the time of installation (or at a mutually agreed upon time free of charge) covering the following areas as applicable: a) Theory of Operation including circuit description and schematics b) Hardware installation, assembly, and disassembly c) Software installation and upgrade procedures d) Troubleshooting procedures e) Preventive Maintenance and Calibration procedures f) Overview of consumables, supplies, and accessories
50 | Terms and Conditions | 11-If the above service training, manuals, and/or software are not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled.
51 | Terms and Conditions | 12-The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
52 | Terms and Conditions | 13-Equipment Downtime: maximum allowable time down is 3% of the total operational time, per annum (calculated based on 44 hrs/week X 52 weeks per annum). In case downtime exceeds this percentage, then the warranty will be extended automatically to twice of the actual down time.
53 | Terms and Conditions | 14-The equipment should have SFDA certificate
54 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | 414348904 | 1168012 | TELEMETRY SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | monitoring system central station wide (main component) | QTY (2)
2 | monitor patient-worn telemetry module (component) | QTY (7)
3 | Intended For | High-Acuity Patient Monitors
4 | PATIENT TYPE | Adult, pediatric, neonate
5 | MONITORING PARAMETER
6 | Parameters | Multichannel ECG
7 | Optional Parameters | Oximetry
8 | ECG Number of leads | 5 leads ECG system or better
9 | ECG HR range, bpm | 30-250
10 | HR Accuracy | <5% or 5 bpm
11 | 12 Lead ECG Disclosure
12 | Interpretation | Yes
13 | ST analysis | Yes
14 | 1 mV calibration
15 | Respiration Breath Rate, bpm | 0 - 30
16 | Respiration Method | ECG Impedance
17 | Respiration Waveform displayed
18 | Respiration Threshold control
19 | PULSE OXIMETRY | Yes
20 | PULSE OXIMETRY Probe type | Nellcor
21 | PULSE OXIMETRY Disposable/reusable | Reusable
22 | ALARMS
23 | HR Alarms
24 | IBP Alarms
25 | Arrhythmia detect | Full Arrythmia
26 | Visual ALARM | Yes
27 | Audible ALARM | Yes
28 | TRENDING | Yes
29 | TRENDING Parameters | All
30 | TRENDING, Graphical/tabular | Both
31 | NETWORK CONNECTIVITY - INTERFACE
32 | NETWORKING
33 | NETWORKING, Configuration | WIFI, access point (Preferred)
34 | NETWORKING Architecture | RF networking is also accepted but preference will give to WIFI connectivity between device and central station or server.
35 | HARDWARE SPECIFICATION
36 | Battery | Removable lithium ion is preferred but AA alkaline battery will be accepted
37 | Battery Slots | One or more
38 | Battery Type | Lithium Ion or Nickel Metal Hydride
39 | Battery Back up | At-least 48Hrs
40 | Battery Life | 500 Cycles or 3 Years (Preferred)
41 | Voltage Range : 100 ~ 240 Volts for central station/server
42 | Frequency : 50Hz / 60Hz
43 | Maximum Current : ≤ 5 Amp
44 | Lead Analyzed simultaneously | Four or more
45 | QRS Detection Range | 0.5 to 5 mV
46 | Water Resistance | IPX7
47 | Frequency Range for RF based System only | 420 – 460MHZ or 610 – 615Mhz (Preferred)
48 | Defibrillator Protection | Up to 5k VDC
49 | Lead Fail detection
50 | Parameter Wave Forms | Two per sector
51 | Sampling Rate | At least 200 /sec
52 | Network Coverage | All level at Private Wing
53 | Central Station Device Display | At least 8 on Each screen
54 | Dual Monitor Set up | One for all patient view and other for detailed view, and full disclosure
55 | External Speakers | For Alarming
56 | Restricted Alarm Suspend feature
57 | Restricted Alarm Volume level
58 | HL7 outbound
59 | Uninterruptable power supply | 1KVA with each central station
60 | UPS backup time | Min. 45 mins at 50% load
61 | Electrical requirement | 220V auto frequency 50/60Hz
62 | Control Bedside settings from central
63 | Maximum Current | ≤ 5A
64 | warranty | The vendor is responsible to provide a minimum of TWO (2) year full warranty commencing from the date of first case.
65 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of TWO (2) year full warranty commencing from the date of first case.
66 | 1.01 Warranty and Technical Support | b) Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
67 | 1.01 Warranty and Technical Support | c) Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
68 | 1.01 Warranty and Technical Support | d) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
69 | 1.01 Warranty and Technical Support | e) The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
70 | 1.01 Warranty and Technical Support | f) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
71 | 1.01 Warranty and Technical Support | g) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
72 | 1.01 Warranty and Technical Support | 1.02 Software and upgrades
73 | 1.01 Warranty and Technical Support | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
74 | 1.01 Warranty and Technical Support | b) Original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
75 | 1.01 Warranty and Technical Support | 1.03 Certification and equipment recalls
76 | 1.01 Warranty and Technical Support | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
77 | 1.01 Warranty and Technical Support | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
78 | 1.01 Warranty and Technical Support | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
79 | 1.01 Warranty and Technical Support | 1.04 Power Requirement
80 | 1.01 Warranty and Technical Support | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted.
81 | 1.01 Warranty and Technical Support | b) The Frequency must be 60Hz.
82 | 1.01 Warranty and Technical Support | c) Power plug should be British style medical grade.
83 | 1.01 Warranty and Technical Support | 1.05 Manuals
84 | 1.01 Warranty and Technical Support | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
85 | 1.01 Warranty and Technical Support | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
86 | 1.06 Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
87 | 1.06 Training | b) The vendor shall provide factory-level service training for TWO service engineers, at the factory or authorized training facility covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation, with no additional cost to the Hospital. The training minimum contents are:
88 | 1.06 Training | I. Theory of Operation including circuit description and schematics.
89 | 1.06 Training | II. Hardware installation, assembly, and disassembly.
90 | 1.06 Training | III. Software installation and upgrade procedures.
91 | 1.06 Training | IV. Troubleshooting procedures.
92 | 1.06 Training | V. Preventive Maintenance and Calibration procedures.
93 | 1.06 Training | VI. Overview of consumables, supplies, and accessories.
94 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | 1160687 | 1160687 | UV DISINFECTOR FOR ULTRASOUND PROBES WITH STORAGE CABINET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | High Level Disinfection by UV-c | Yes
2 | Time Required | 180sec
3 | Chemical free | Yes
4 | Effective | Bactericide,Viricide,Fungicide, Mycobactericide & Sporicide.
5 | Automatic | Yes
6 | Size | Cylindrical Enclosure
7 | Diameter | 55 cm
8 | Height | 202 cm
9 | Weight | 110 kg
10 | Power | 1430 Watt
11 | Supply Voltage | 230V @60HZ
12 | Intensity | 6.2 A
13 | NO CONSUMABLES | Yes
14 | Report print electronic version | Yes
15 | Wheels / Can be shifted easily | Yes
16 | Environmental Friendly | Yes
17 | Included:
18 | AUTOMATED DISINFECTION MACHINE UV-C FOR EXTERNAL ULTRASOUND PROBES AND ENDOCAVITARY BASED ON HIGH LEVEL DISINFECTION SYSTEM | QTY (1)
19 | spare LAMP UVC 325W | QTY (4)
20 | LAMP UVC 35W | QTY (2)
21 | TAG RADIO FREQUENCY IDENTIFICATION FOR TEE PROBE DISINFECTION SYSTEM | QTY (3)
22 | DISKETTE LABELS for UVC DISINFECTION MACHINE | QTY (10)
23 | CABINET TEE PROBES STORAGE | QTY (1)
24 | EQUIPMENT WARRANTY | The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale. The vendor must indicate clearly the warranty period of all parts, labor and glassware. The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year. However additional warranty period requirements should be fulfilled if requested in technical specification. The warranty period is to commence after the official institution acceptance testing document has been completed. The institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM issues and so forth
25 | TERMS AND CONDITIONS | Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and supporting documentation that:
26 | TERMS AND CONDITIONS | · the equipment and its accessories are approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.:NFPA or equivalent); and
27 | TERMS AND CONDITIONS | · the equipment meets the standard of electromagnetic compatibility
28 | TERMS AND CONDITIONS | One (1) original plus five (5) copies of the tender response shall be provided. Each copy of the documents is to include their own set of brochures.
29 | TERMS AND CONDITIONS | The manufacturer will confirm the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
30 | TERMS AND CONDITIONS | All equipment must conform to the site existing DICOM standard and DICOM Worklist must be installed, enabled and seamless.
31 | TECHNOLOGY STANDARD (As Applicable) | CD/DVD & USB port in the system
32 | TECHNOLOGY STANDARD (As Applicable) | The equipment proposed shall incorporate any “latest generation” hardware and software technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
33 | TECHNOLOGY ENHANCEMENTS | The vendor will guarantee an official notification to the institution of any equipment updates or model changes occurring immediately before the processing of this order. The vendor shall notify the institution accordingly and grant the institution the option to purchase the upgrade or revised models.
34 | TECHNOLOGY ENHANCEMENTS | Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as an option.
35 | TECHNOLOGY ENHANCEMENTS | Software upgrades that become available within the warranty period shall be incorporated at no additional charge.
36 | TECHNOLOGY ENHANCEMENTS | Equipment price must be clearly identified and guaranteed for a minimum of three (3) months from the date of receiving vendor quotation.
37 | EQUIPMENT PRICING | All prices must include delivery and installation.
38 | EQUIPMENT PRICING | If applicable, the vendor must also include the price to remove any previously installed equipment.
39 | EQUIPMENT PRICING | Prices shall include all accessories as listed in the specifications and as necessary for a complete, fully functional installation. In addition, any optional equipment items listed must be itemized and individually priced in the same fashion.
40 | EQUIPMENT PRICING | The institution reserves the right to accept all or part of any proposal received and thus for
41 | PROPOSAL ACCEPTANCE | RFP proposals that include more than one piece of equipment there can be multiple vendor awarding.
42 | PROPOSAL ACCEPTANCE | The institution reserves the right not to accept the lowest quoted price.
43 | PROPOSAL ACCEPTANCE | The institution reserves the right to accept the best suitable offer according to the medical operations requirements as tertiary health care institution
44 | PROPOSAL ACCEPTANCE | ARTICLE 6: APPLICATION /SERVICETRAINING
45 | PROPOSAL ACCEPTANCE | Included with each proposal shall be an explanation of the application training provided to department personnel.
46 | PROPOSAL ACCEPTANCE | Instruction and demonstrations will be provided by a qualified application specialist, per the end user’s standards and approval.
47 | PROPOSAL ACCEPTANCE | The institution reserves the right not to accept an application specialist, and the vendor must provide a replacement at no additional cost to the institution.
48 | PROPOSAL ACCEPTANCE | The vendor must provide two (2) application training sessions: the first visit just after equipment clinical acceptance, and the second visit within six (6) months. The date of instruction will be designated by mutual agreements of the vendor and the Non-Invasive Cardiology Lab.
49 | PROPOSAL ACCEPTANCE | The vendor must adhere to KFSH&RC working hours when booking application training sessions. One (1) visit = five (5) KFSH&RC working days: Sunday through to Wednesday from 7:30 a.m. until 5:00 p.m. and Thursday from 7:30 a.m. until 4:00 p.m.
50 | PROPOSAL ACCEPTANCE | The vendor will specify, with advance notification, their requirements for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
51 | PROPOSAL ACCEPTANCE | Full local Service Training to be provided at the time of installation.
52 | PROPOSAL ACCEPTANCE | The vendor shall guarantee software upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the institution.
53 | SYSTEM SOFTWARE UPGRADES(If Applicable) | If the proposed model is a first time offer to KFSH & RC the vendor must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operate the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities.
54 | REFERENCES | All plans and specifications must be returned to the owner within one (1) week of notification of award. The successful vendor(s) may retain these documents for reference.
55 | INSTITUTION DOCUMENTATION | Any exceptions to the terms and conditions of the following equipment specifications must be clearly stated and submitted in writing at the time of proposal.
56 | EXCEPTIONS TO TERMS AND CONDITIONS | The vendor in submitting his quotation represents that he has read and understands the proposal documentations and his quotation is made in accordance there within.
57 | VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | 13.2 Should the vendor representative require any assistance in responding to the said RFP documentation they shall send an official letter to Contracts Management Department.
58 | VENDOR REPRESENTATIVE AGREEMENT AND ASSISTANCE | The vendor must supply complete warranty details for all quoted equipment on the terms and conditions of sale.
59 | EQUIPMENT WARRANTY | The vendor must indicate clearly the warranty period of all parts, labor and glassware.
60 | EQUIPMENT WARRANTY | The minimum warranty period provided by the vendor that is acceptable to this institution shall be one (1) year. However additional warranty period requirements should be fulfilled if requested in technical specification.
61 | EQUIPMENT WARRANTY | The warranty period is to commence after the official institution acceptance testing document has been completed.
62 | EQUIPMENT WARRANTY | The institution reserves the right to delay clinical acceptance until such a time that all pending implementation issues have been resolved to the satisfaction of the institution. Implementation issues can be but not all inclusively include: acceptance testing troublesome findings, incomplete installations and/or equipment delivery, workstation and/or software activation delays due to license key delays, DICOM issues and so forth.
63 | EQUIPMENT WARRANTY | Vendor personnel shall familiarize themselves with the features and characteristics of the building which will affect the delivery and/or installation of their equipment. The vendor shall check space, height, accessibility to all applicable areas to ensure ease of installation. Vendors must advise if the proposed equipment will NOT fit through a doorway.
64 | ARTICLE 15: EXAMINATION OF SPACE AND SERVICES (when applicable) | The vendor shall report inadequacies to the owner immediately to facilitate remedial measures prior to scheduling installation.
65 | ARTICLE 15: EXAMINATION OF SPACE AND SERVICES (when applicable) | The vendor shall accept responsibility for error or neglect to familiarize with access, space and service conditions. There will be no storage of furniture or equipment on the job site, without prior authorization in writing from the owner.
66 | ARTICLE 15: EXAMINATION OF SPACE AND SERVICES (when applicable) | It is the responsibility of the vendor to remove old equipment (if any) from previously used examination room(s) and to prepare the site (site renovations) for new equipment installation.
67 | SITE PREPARATION / SCOPE OF WORK (As applicable) | The delivery of equipment to the site of installation is subject to prior approval of the owner with at least two (2) weeks’ notice by company letter.
68 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | Installation shall be done with close cooperation with the assigned KFSH & RC representatives.
69 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | Delivery authorization will be provided by the owner at the appropriate time. Installation may commence when the finished space is ready to receive equipment and approval has been granted by the owner. The owner will advise the successful vendor of the scheduled commencement of installation, so as to permit the vendor to provide technical personnel in a timely manner.
70 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
71 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc.) for each employee they will have on site so that appropriate hospital gate passes can be issued.
72 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | The Clinical Engineer will be present during installation and initial testing by the manufacturer.
73 | SCHEDULING AND COORDINATION OF DELIVERY AND INSTALLATION | 18.1 All mobile equipment (and powered accessories) delivered to the must be dual voltage 110-220 VAC/60 Hz power (to be operated by 220v as well as 110v as applicable) and power plug/outlet should be American or British style, medical grade, depending upon final location of equipment.
74 | MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | 18.2 Mobile equipment not satisfying statement 21.1 will not be accepted by the institution and will ultimately delay clinical acceptance and final contract payment.
75 | MOBILE EQUIPMENT POWER REQUIREMENTS (where applicable) | The vendor shall keep the premises, including assigned rooms, all passages and auxiliary spaces free from accumulation of waste material or rubbish caused by him at the end of each day.
76 | CLEANING UP AND MAKING GOOD ANY DAMAGES | The vendor shall assume full responsibility for all damages to the building plant and previously placed equipment, including gouges and smudges to the owner’s property made by their own personnel and/or sub-contractors.
77 | CLEANING UP AND MAKING GOOD ANY DAMAGES | Any resulting repairs by the vendor shall match the adjacent existing finishes and materials. The repairs are to be done by appropriate trade contractors without any additional cost to KFSH&RC.
78 | CLEANING UP AND MAKING GOOD ANY DAMAGES | The vendor shall comply with pre-infection control procedures of the hospital.
79 | CLEANING UP AND MAKING GOOD ANY DAMAGES | Two (2) complete operational manuals for hardware and software.
80 | ARTICLE 20: DOCUMENTATION | Two (2) service manuals hard and soft copies.
81 | ARTICLE 20: DOCUMENTATION | Two (2) spare part manuals (if applicable).
82 | ARTICLE 20: DOCUMENTATION | Completed schematic diagrams with documentation of onsite modifications (if required).
83 | ARTICLE 20: DOCUMENTATION | All above documentation must be provided in English.
84 | ARTICLE 20: DOCUMENTATION | Preventive maintenance procedure (softcopy).
85 | ARTICLE 20: DOCUMENTATION | Upon completion of installation and prior to first clinical utilization of the equipment, the entire imaging system shall undergo acceptance testing by the Clinical Engineering Department.
86 | EQUIPMENT ACCEPTANCE TESTING | The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
87 | EQUIPMENT ACCEPTANCE TESTING | • Sales Literature;
88 | EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin;
89 | EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
90 | EQUIPMENT ACCEPTANCE TESTING | • Other international regulatory agencies (i.e.: FDA, ACR, AAPM)
91 | EQUIPMENT ACCEPTANCE TESTING | Any deficiencies identified upon acceptance testing must be rectified by the vendor to the institution’s satisfaction prior to commencing clinical acceptance and application training.
92 | EQUIPMENT ACCEPTANCE TESTING | The medical physicist shall have access to the same diagnostic and testing software as the vendor engineer.
93 | EQUIPMENT ACCEPTANCE TESTING | The Clinical Engineering Department shall have access to the same diagnostic and testing phantoms as the vendor engineer.
94 | EQUIPMENT ACCEPTANCE TESTING | Vendor must provide the following information as it applies to their service organization:
95 | MANUFACTURER’S SERVICE COMMITMENT | Local service provider (i.e. office address and contact information).
96 | MANUFACTURER’S SERVICE COMMITMENT | Service response time for routine and after hour calls. The vendor shall define the term
97 | MANUFACTURER’S SERVICE COMMITMENT | ‘Response Time’ and provide their minimum and maximum time limits.
98 | MANUFACTURER’S SERVICE COMMITMENT | For each piece of equipment specified within this document, list parts that are critical for normal operation of the equipment that are not available from the local service center, and would take more than seven (7) days to receive from another area.
99 | MANUFACTURER’S SERVICE COMMITMENT | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment.
100 | MANUFACTURER’S SERVICE COMMITMENT | The vendor shall guarantee to provide to this institution with technical support and the supply of spare parts for a period of ten (10) years after the date of last equipment production. The vendor shall also inform the institution of the end of production date when announced.
101 | MANUFACTURER’S SERVICE COMMITMENT | Equipment downtime shall not exceed 3% of the total operation time per annum; downtime computations will assume 52 weeks per year and 44 work hours per week, including Ramadan and Hajj holidays.
102 | MANUFACTURER’S SERVICE COMMITMENT | Service response time for routine calls should be three (3) hours, including Ramadan and Hajj holidays. Penalties may ensue should the engineer not respond in a timely manner (3% DT).
103 | MANUFACTURER’S SERVICE COMMITMENT | Service engineer must provide a service report each time they evaluate and/or work on the equipment. Penalties will ensue should the engineer not comply (3% DT).
104 | MANUFACTURER’S SERVICE COMMITMENT | All empty boxes, tools, spare parts, defective parts, replacement parts and the like, will be removed from the premises immediately and/or temporarily stored in the RIS/PACS storage room as long as a property pass is in process. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc…
105 | MANUFACTURER’S SERVICE COMMITMENT | Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized and clutter-free). Penalties will ensue should the engineer not comply (3% DT).
106 | MANUFACTURER’S SERVICE COMMITMENT | The vendor is requested to fill the data specification sheets for each piece of equipment summarized below in the manner it is presented – please provide a soft copy of completed specification sheets, with all relevant clinical and technical documents, on a flash drive and submit five (5) copies with the offer. Not providing data requested in a soft copy format can delay the vendor selection process wherein the institution does reserve the right to ask a vendor to recompile and re-submit the data.
107 | GENERAL COMMENTS | If data is not available for a certain requested data field the vendor should not leave the field blank but instead use the term ‘No Data’ in the field. If this term is not used, it will be assumed that the data sheet is incomplete and a resubmission may be requested. The vendor does have the right to reference the field to an attachment to better clarify the entry.
108 | GENERAL COMMENTS | All references to a data specification sheet field must use the corresponding reference number.
109 | GENERAL COMMENTS | The vendor must provide sufficient manufacturer documentation to support the data quoted.
110 | GENERAL COMMENTS | The vendor should provide additional information that may not have been requested on the data specification sheets that can elaborate on the quoted technology capabilities and clinical applications.
111 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | 417666001 | 1168762 | ANALYZER SYSTEM FOR VOICE AND SPEECH CLINICAL ASSESSMENT AND BIOFEEDBACK |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Bill Of Quantities | Computerized Speech lab qty2 and Visi-Pitch qty 2
2 | COMPUTERIZD SPEECH LAB (CSL) | • Hardware/Software System Which Includes Professional Quality I/O System • External Module with Input Filter & Amplifier • PCI Interface To Host Computer • 4 Channels (24-Bit, Max:200 kHz per Channel) • High Fidelity Output with Studio Speaker • Professional Quality Headphones & Microphones • Balanced Microphone (XLR) & Line Inputs • AC & DC Coupling • Auto calibration Of Hardware for DC Offsets • Digital Control of Levels • High-Gain Pre-Amp with Overload Detection Includes but not limited to: • S/PDIF I/O for Digital Communication w/DAT & Other Digital Recorders and Players. • Latest Software Acquisition, Playback, Editing and Analysis. (Sonograms, LPC, Pitch, Intensity, FFT, Editing Tools, Digital, Filtering, IPA Transcription, Data Log Features, etc...) • Windows MDI and CUI Compliance with user-Configurable Workspace, Flexible Access to Commands & User-Definable Macro/Keys.
3 | PC | COMPUTER, Latest Version with Color Desk Jet Printer (ink has to be available in the local market) & Table
4 | Software | Providing Feedback for Speech Parameters and Voice Exercises. For Speech/Voice Therapy Involving Children.
5 | Multi-Dimensional Voice Program | Calculates Over 22 Different Voicing Parameters and Compares Results to Database.
6 | Multimedia pedagogical program | On the applied speech science for voice and resonance disorders. Program provides extensive graphic and descriptive information on various aspects of voicing and resonance. Requires sound card.
7 | REAL-TIME SPECTROGRAM (RTS) | Generating Real-Time Narrow and Wide-Band Spectrograms. For Visual Biofeedback Applications; Especially for Classroom Teaching Demonstrations
8 | Video Phonetics Program and CD-ROM Database. | Used for Loading and Displaying Digital Video Files (.avi) Within Multi-Speech. Audio Is Separated for Analysis. 200 Files of Video Speech Linked to IPS Symbols. Video Display May Include Lips, Jaw, Tongue, Pharynx, or The Glottis.
9 | PHONETIC DATABASE (CSL -SOFTWARE OPTION) | Database (One CD-ROM) consisting of More Than 4000 Separate Digitized Files of 45 Languages.
10 | Real time pitch | Yes
11 | SONA match | Yes
12 | Delayed Auditory feedback | Yes
13 | Motor speech profile | Yes
14 | Voice range profile | Yes
15 | condensing microphone | ASC Contains Three Unidirectional Sensors for Placement in Oropharynx & Hypo pharynx. Three Centimeter Distal Tip Included. Calibration Kit Provided.
16 | Nasal Canuula | Used for Viewing of Inspiration, Expiration and Apnea During Swallowing. Disposable but Reusable by The Same patient.
17 | Laryngeal Microphone and Powered Speaker Laryngeal Stethoscopes Microphone. (1500Hz Freq. Response- | For Acoustic Analysis of Swallow. Includes Powered Speaker to Enable Patient and Clinician to Monitor Swallowing and Breathing Aurally.
18 | System Cart | Yes
19 | Color Printer | Yes ink to be available in the local market
20 | warranty | The vendor is responsible to provide a minimum of Three (3) year full warranty commencing from the date of Technical acceptance
21 | Technical Terms and Conditions | 1. The vendor is responsible to provide a minimum of Three (_3_) year full warranty commencing from the date of Technical acceptance.
22 | Technical Terms and Conditions | 2. Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
23 | Technical Terms and Conditions | 3. Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
24 | Technical Terms and Conditions | 4. The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
25 | Technical Terms and Conditions | 5. a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
26 | Technical Terms and Conditions | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
27 | Technical Terms and Conditions | 6. The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert.
28 | Technical Terms and Conditions
29 | Technical Terms and Conditions | 7. Voltage:
30 | Technical Terms and Conditions | a) Dual (To be operated by 220V as well as 110V as applicable, otherwise 110V is accepted).
31 | Technical Terms and Conditions | b) Frequency: 60Hz.
32 | Technical Terms and Conditions | d) Power plug should be American style medical grade.
33 | Technical Terms and Conditions | 8. The vendor to provide:
34 | Technical Terms and Conditions | a) Installation.
35 | Technical Terms and Conditions | b) Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase).
36 | Technical Terms and Conditions | c) Set up.
37 | Technical Terms and Conditions | d) Calibration.
38 | Technical Terms and Conditions | e) Performance verification.
39 | Technical Terms and Conditions | f) Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
40 | Technical Terms and Conditions | h) UPS: if applicable
41 | Technical Terms and Conditions | 9. One (1) complete operations manual, to be provided preferably in both hard and soft copies.
42 | Technical Terms and Conditions | 10. One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
43 | Technical Terms and Conditions | 11. If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
44 | Technical Terms and Conditions | 12. Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
45 | Technical Terms and Conditions | 13. Full local technical service training to be provided, by a certified trainer, at the time of installation (or at a mutually agreed upon time free of charge) covering, but not limited to, the following areas as applicable:
46 | Technical Terms and Conditions | i) Theory of Operation including circuit description and schematics
47 | Technical Terms and Conditions | ii) Hardware installation, assembly, and disassembly
48 | Technical Terms and Conditions | iii) Software installation and upgrade procedures
49 | Technical Terms and Conditions | iv) Troubleshooting procedures
50 | Technical Terms and Conditions | v) Preventive Maintenance and Calibration procedures
51 | Technical Terms and Conditions | vi) Overview of consumables, supplies, and accessories
52 | Technical Terms and Conditions | a) If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided.
53 | Technical Terms and Conditions | 15. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
54 | Technical Terms and Conditions | 16. The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA).
55 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
| Company Stamp |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | 414867100 | 1191576 | OPHTHALMOSCOPE / OTOSCOPE SET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Ophthalmoscope:
2 | Lamp | LED
3 | FDA Clearance | Required, Provide Certificate number
4 | SFDA Clearance | Required, Provide Certificate number
5 | CE Mark | Required, Provide Certificate number
6 | Number of aperture/filter combinations | 18
7 | Number of focusing lenses | 28
8 | Range of diopters | -25 to +40
9 | (Small, Medium, Large Spot) Red-free, Slit and Cobalt blue, Corneal Magnifying Lens | Required
10 | Otoscope:
11 | Lamp | LED
12 | Type | Advance, high durable
13 | Magnification | 4.2X
14 | Viewing area | Nearly complete view of the tympanic membrane
15 | Adjustable focus wheel | Required
16 | Ear Speculum Type | Disposable computable with hospital standard
17 | Ejectable Tips | Required
18 | Insufflation port for pneumatic otoscopy | Required
19 | warranty | The vendor is responsible to provide a minimum of (1) year full warranty commencing from the date of first case
20 | Software and upgrades | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
21 | Certification and equipment recalls | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
22 | Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
23 | Power Requirement | Battery operated.
24 | Manuals | One (1) complete operations manual, to be provided preferably in both hard and soft copies.
25 | Manuals | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies If the vendor fails to provide manuals as stated above, then equipment acceptance will be put on hold until requirements provided.
26 | Manuals | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
27 | Manuals | The Equipment’s should have SFDA certification
28 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | 102508 | 1191576 | EXAMINATION TABLE POWERED |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA Clearance
2 | SFDA Clearance
3 | CE Mark
4 | 3 storage sections each with laminate interiors and 1 adjustable shelf.
5 | adjustable backrest
6 | Paper holder, (wooden/plastic bar)
7 | Remote control, foot and hand
8 | Weight capacity: | 150 Kg or more
9 | 2” high-density urethane foam top
10 | Each door shall be equipped with 1 adjustable laminate shelf.
11 | Concealed, self-closing hinges.
12 | It shall have laminated plywood base
13 | resistant vinyl upholstery finish.
14 | warranty | The vendor is responsible to provide a minimum of (3) year full warranty commencing from the date of first case
15 | Warranty and Technical Support | The vendor to provide: Installation, set up, Calibration, Performance verification, Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
16 | Software and upgrades | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
17 | Certification and equipment recalls | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
18 | Certification and equipment recalls | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
19 | Power Requirement | Voltage: Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted) Frequency: 60Hz.
20 | Power Requirement | Power plug should be British style medical grade.
21 | Manuals | One (1) complete operations manual, to be provided preferably in both hard and soft copies.
22 | Manuals | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies If the vendor fails to provide manuals as stated above, then equipment acceptance will be put on hold until requirements provided.
23 | Manuals | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
24 | Manuals | The Equipment’s should have SFDA certification
25 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | 442181901 | 1193204 | ELECTRONIC VITAL SIGN monitor with builtin OPHTHALMOSCOPE / OTOSCOPE SET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Comprehensive integrated wall diagnostic with advanced vital signs documentation in one convenient system | Required
2 | Ergonomic design, wall mounted ( removable ) | Required
3 | Simple color touchscreen interface | Required
4 | Electrical : 110-230V volts / 50- 60 Hz. | Required
5 | Vital Signs Monitor
6 | Disply | 8.9 inch color touchscreen display
7 | Parameters: NIBP, SPO2, Tempreture, HR | Required
8 | Manual measurements entry (weight, height, respiration rate, pain ) | Required
9 | Operation Mode : Spot, Office, Intervals, Continuous | Required
10 | Patient population ( Adult, Pediatric, Neonate ) | Required
11 | Monitoring: Intervals for NIBP, SPO2 | Required
12 | Device alarm : Programmable Visual & Aduible alarms for : BP, SPO2, Temp and Pulse | Required
13 | Back-up Battery type | Lithium-ion
14 | Automated Early Warning Scoring to help clinicians identify and respond to signs of patient deterioration (ie. MEWS, NEWS ) | Up to 6 protocols
15 | Custom modifiers for patient documentation | Up to 20
16 | Connect to digital weight scales, Scale Integration | YES
17 | Memory (# Reading) | 400 Readings
18 | Compitable to connect with Hospital Information System | Required
19 | Connectivity Type | WiFi , USB, Ethernet
20 | Integrated 3.5 V handles for Otoscope & Ophthalmoscope | Required
21 | Integrated ear speculum dispenser | Required
22 | Barcode Scanner | Required
23 | NIBP
24 | Systolic range | Adult: 30 to 260 mmHg; Pediatric: 30 to 260 mmHg; Neonate: 20 to 120 mm
25 | Diastolic rang | Adult: 20 to 220 mmHg; Pediatric: 20 to 220 mmHg; Neonate: 10 to 110 mm
26 | Blood pressure determination time | SureBP technology typical: 15 seconds; Maximum: 150 second
27 | Blood pressure accuracy | ± 5 mmHg mean error, 8 mmHg standard devian
28 | SPO2
29 | Type | Reusable
30 | Technology | Nellcor
31 | SpO2 cable length | 10'
32 | Measurement Range | 1-100%
33 | Pulse Rate Range | 25 – 240 bpm
34 | SpO2 Accuracy | ± 3 digits
35 | Pulse Rate Accuracy | ± 3 digits
36 | Tempreture
37 | Type | SureTemp Plus Oral / Axillary
38 | Measurement Range | 26.7 °C to 43.3 °C
39 | Probe cord length | 9'
40 | Calibration Accuracy | ±0.1°C
41 | Determination Time | Oral: 4-6 seconds Adult Axillary: 15 seconds Pediatric Axillary: 10 seconds
42 | Probe Cover | Disposable compitable with current avalible in hospital
43 | Ophthalmoscope
44 | Lamp | LED
45 | Type | Coaxial, shadow-free spot
46 | Number of aperture/filter combinations | 18
47 | Number of focusing lenses | 28
48 | Range of diopters | -25 to +40
49 | (Small, Medium, Large Spot) Red-free, Slit and Cobalt blue, Corneal Magnifying Lens | Required
50 | Otoscope
51 | Lamp | LED
52 | Type | Advance, high durable
53 | Magnification | 4.2X
54 | Viewing area | Nearly complete view of the tympanic membrane
55 | Adjustable focus wheel | Requried
56 | Ear Speculum Type | Disposable compitable with current avalible in hospital
57 | Ejectable Tips | Required
58 | Insuffl ation port for pneumatic otoscopy | Required
59 | warranty | The vendor is responsible to provide a minimum of ONE (3) year full warranty commencing from the date of Technical acceptance.
60 | Technical Terms and Conditions for | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
61 | Technical Terms and Conditions for | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
62 | Technical Terms and Conditions for | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
63 | Technical Terms and Conditions for | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
64 | power and voltage | Voltage: Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted)
65 | power and voltage | Frequency: 60Hz Power plug should be British style medical grade
66 | power and voltage | The vendor to provide: Installation with wall mounting that includes proper hardware, Performance verification
67 | Manuals | One (1) complete operations manual, to be provided preferably in both hard and soft copies.
68 | Manuals | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
69 | Manuals | If the vendor fails to provide manuals, as stated above, then equipment acceptance will be put on hold until requirements provided.
70 | Manuals | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
71 | Local service training | The vendor shall provide Local service training for one service engineer to be given locally. The training minimum contents are:
72 | Local service training | I. Theory of Operation including circuit description and schematics.
73 | Local service training | II. Hardware installation, assembly, and disassembly.
74 | Local service training | III. Software installation and upgrade procedures.
75 | Local service training | IV. Troubleshooting procedures.
76 | Local service training | V. Preventive Maintenance and Calibration procedures.
77 | Local service training | VI. Overview of consumables, supplies, and accessories.
78 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | 411532308 | 1193204 | MONITOR PATIENT PHYSIOLOGICAL FOR ANAESTHESIA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Intended For | High-Acuity Patient Monitors
2 | PATIENT TYPE | Adult, pediatric, neonate
3 | MONITORING PARAMETER
4 | Parameters | ECG, SpO2, NIBP, respiration, temperature, IBP, ETCO2
5 | Optional Parameters | BIS, SvO2, tcpO2, tcpCO2
6 | ECG Number of leads | 12 ECG system or better
7 | ECG HR range, bpm | 30-250
8 | HR Accuracy | <5% or 5 bpm
9 | 12 Lead ECG Disclosure
10 | Interpretation | Yes, with 12 SL 12-Lead ECG
11 | Arrythmia No. of leads analyzed | 4 Leads
12 | Arrhythmia Sampling rate: | at least 250 sample/second
13 | ST analysis | Yes
14 | ST No. of leads analyzed | 12 Leads
15 | Auto lead switch | Preferred
16 | 1 mV calibration
17 | Respiration Breath Rate, bpm | 0 - 30
18 | Respiration Method | ECG Impedance or etCO2
19 | Respiration Waveform displayed
20 | Respiration Threshold control
21 | IBP
22 | IBP No. of channels | 4 or More
23 | IBP Scales range, mm Hg | -30 to +340
24 | IBP Labels
25 | NIBP
26 | NIBP Measurement technique | Oscillometric
27 | NIBP Pressure Range | At least 20-300 mmHg
28 | NIBP Cuff size | adult, pediatric, neonate, thigh
29 | NIBP Hose connection | Any; Luer not recommended (please specify)
30 | PULSE OXIMETRY | Yes
31 | PULSE OXIMETRY Probe type | Nellcor, Masimo
32 | PULSE OXIMETRY Disposable/reusable | Reusable
33 | TEMPERATURE
34 | TEMPERATURE, No. of channels | 2 channels
35 | TEMPERATURE Probe type | Oral, axillary, rectal
36 | ETCO2
37 | ETCO2 accuracy | ± 5%
38 | ALARMS
39 | HR Alarms
40 | IBP Alarms
41 | Arrhythmia detect | Full Arrythmia
42 | Respiration Rate Alarm
43 | Systolic Pressure
44 | Diastolic Pressure
45 | Low SpO2 Saturation
46 | ETCO2 Alarm
47 | ALARMS Type | HR/PR, respiration, NIBP, SpO2, lethal arrhythmia
48 | Visual ALARM | Yes
49 | Audible ALARM | Yes
50 | TRENDING | Yes
51 | TRENDING Parameters | All
52 | TRENDING, Graphical/tabular | Both
53 | Trending Length of time, hr | Min. 72
54 | Retrieve Patient data After Discharging
55 | Silence active alarm | Yes
56 | ALARM Silence duration, sec | 60 or 120
57 | All-alarm suspend | Yes
58 | All-alarm suspend Silence duration, sec | userconfigured 60 or 120
59 | Alarm Volume Retriction feature
60 | Alarm Suspend Restriction
61 | NETWORK CONNECTIVITY - INTERFACE
62 | Interfaced to other devices, | ventilators, anesthesia units, intravenous pumps, pulse oximeters, and multiple medical-gas monitoring systems. (Preferred)
63 | NETWORKING
64 | NETWORKING, Configuration | Hardwired
65 | NETWORKING Communication | Ethernet, TCP/IP
66 | Network Architecture | Recommended Ethernet (if others, please specify)
67 | Max no. of devices on single network | up to 1024
68 | Bedside-to-bedside monitor communication | Yes
69 | Bedside-to-bedside Auto View on Alarm | Optional
70 | Networking Other clinical information systems | Hospital Information System with HL7 compliance
71 | Peripheral device interface
72 | DISPLAY SPECIFICATION
73 | Display Type | LCD, Medical Grade
74 | DISPLAY Size, cm (in) | 48.3 (19)
75 | Display Traces | 8, Configurable
76 | Display USER INTERFACE | Trim Knob, remote control, touchscreen
77 | security management system
78 | PERIPHERALS
79 | Full disclosure of ECG waveforms. | At least 72 hours
80 | Recorder/Printer | Yes, Thermal dot array
81 | Recorder Channels | Single
82 | Auto Print | Detect arrhythmias
83 | HARDWARE SPECIFICATION
84 | Battery
85 | Battery Slots | One or more
86 | Battery Type | Lithium Ion or Nickel Metal Hydride
87 | Battery Back up | Min. 1 hr
88 | Battery Life | 500 Cycles or 3 Years
89 | Monitor configuration | Modular Bedside - Wall Mounted
90 | Monitor must have dual voltage operational compatibility at automatic selection of frequency.
91 | Voltage Range : 100 ~ 240 Volts
92 | Frequency : 50Hz / 60Hz
93 | Maximum Current : ≤ 5 Amp
94 | The following Parameters will be dedicated for Neonatal monitoring:
95 | Transcotaneous blood gas Monitoring pO2/pCO2 | Yes
96 | Transcotaneous pO2 Range: | 0 – 750 mmHg or 0.0 – 100.0 kPa
97 | Transcotaneous pO2 Accuracy: | Better than ±1 mmHg ±0.1% of reading at 25C
98 | Transcotaneous PCO2 Channel Range: | 5 – 200 mmHg or 0.7 – 26.7 kPa
99 | Transcotaneous PCO2 Accuracy: | Better than ±1 mmHg ±1% of reading at 25C
100 | All required accessories for above functions, for Neonatal applications, should be included. | Yes(please specify the included needed accessories)
101 | Neonatal Cuffs, all sizes | Yes
102 | Rectal and skin neonatal sensors | Yes
103 | warranty | The vendor is responsible to provide a minimum of THREE (3) year full warranty commencing from the date of first case.
104 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of THREE (3) year full warranty commencing from the date of first case.
105 | 1.01 Warranty and Technical Support | b) Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
106 | 1.01 Warranty and Technical Support | c) Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
107 | 1.01 Warranty and Technical Support | d) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
108 | 1.01 Warranty and Technical Support | e) The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
109 | 1.01 Warranty and Technical Support | f) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
110 | 1.01 Warranty and Technical Support | g) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
111 | 1.02 Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
112 | 1.02 Software and upgrades | b) Original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
113 | 1.03 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
114 | 1.03 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
115 | 1.03 Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
116 | 1.04 Power Requirement | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted.
117 | 1.04 Power Requirement | b) The Frequency must be 60Hz.
118 | 1.04 Power Requirement | c) Power plug should be British style medical grade.
119 | 1.05 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
120 | 1.05 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
121 | 1.06 Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
122 | 1.06 Training | b) The vendor shall provide Local service training for ONE service engineers, at the authorized training facility covering all associated expenses with no additional cost to the Hospital. The training minimum contents are:
123 | 1.06 Training | I. Theory of Operation including circuit description and schematics.
124 | 1.06 Training | II. Hardware installation, assembly, and disassembly.
125 | 1.06 Training | III. Software installation and upgrade procedures.
126 | 1.06 Training | IV. Troubleshooting procedures.
127 | 1.06 Training | V. Preventive Maintenance and Calibration procedures.
128 | 1.06 Training | VI. Overview of consumables, supplies, and accessories.
129 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | 314867101 | 1193204 | UNIVERSAL CHARGER OTO/OPHT SCOPE COMPLETE W/TWO 3.5V RECHARGEABLE HANDLES W/O HEADS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Ophthalmoscope:
2 | Lamp | LED
3 | Type | Coaxial, shadow-free spot
4 | Number of aperture/filter combinations | 18
5 | Number of focusing lenses | 28
6 | Range of diopters | -25 to +40
7 | (Small, Medium, Large Spot) Red-free, Slit and Cobalt blue, Corneal Magnifying Lens | Required
8 | Otoscope:
9 | Lamp | LED
10 | Type | Advance, high durable
11 | Magnification | 4.2X
12 | Viewing area | Nearly complete view of the tympanic membrane
13 | Adjustable focus wheel | Requried
14 | Ear Speculum Type | Disposable compitable with current avalible in hospital
15 | Ejectable Tips | Required
16 | Insuffl ation port for pneumatic otoscopy | Required
17 | FDA Clearance | Required, Provide Certificate number
18 | SFDA Clearance | Required, Provide Certificate number
19 | CE Mark | Required, Provide Certificate number
20 | warranty | The vendor is responsible to provide a minimum of two-years full warranty commencing from the date of Technical acceptance.
21 | warranty | 1-The vendor is responsible to provide a minimum of two-years full warranty commencing from the date of Technical acceptance.
22 | warranty | 2-The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
23 | warranty | 3-The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
24 | warranty | 4-The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
25 | warranty | 5-Voltage: 220V , 60 Hz
26 | warranty | 6-The vendor to provide installation, set up, calibration, performance verification, preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
27 | warranty | 7-Site preparations, and pre-installtion works (civil, electrical, mechanical, networking, etc), are the responsibility of the HOSPITAL, unless, otherwise agreed upon/quoted by the SUPPLIER.
28 | warranty | 8-One (1) complete operations manual, to be provided in both hard and soft copies if available.
29 | warranty | 9-One (1) complete service manual (in English language) including schematics and spare parts list (if applicable). The manual is to be provided in both hard and soft copies if available.
30 | warranty | 10-The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
31 | warranty | 11-The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA).
32 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | 112471 | 1193204 | CHAIR DIALYSIS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | The proposed dialysis chair shall have the following features as minimum:
2 | It shall be a multi-positioning fully motorized (At least four motors) chair | yes
3 | It shall have large stable base | yes
4 | It shall have height adjustment | yes
5 | The chair back shall fully recline to the horizontal position | yes
6 | The chair foot area shall raise to the Trendelenburg position | yes
7 | microprocessor controlled handset
8 | It shall have hand-held control unit for easy adjustment of the chair/recliner | yes
9 | Tilt-up and Tilt-down armrest for easy of sit down and get out of the chair for more comfortability and safety | yes
10 | Increasing distance between the armrest | yes
11 | layered cushion against the liquid for the armrest | yes
12 | It shall have two large manually adjustable armrests: | yes
13 | Adjustable foot rest | yes
14 | Swivel for easy blood donor ingress and outgress | yes
15 | Large stable venipuncture platform | yes
16 | It shall have four large casters | yes
17 | Central Locking break | yes
18 | Allow full mobility of the chair, all caster to rotate 360 degrees | yes
19 | Lock for safety and security during blood drawing | yes
20 | All surfaces shall be easily cleaned and sanitized: | yes
21 | Surface material to be able to be removed, washed and reinstalled on the same chair | yes
22 | saving option for saving weight data | yes
23 | Exchangeable upholstery with Blood resistant artificial leather | yes
24 | Additional accessories like iPad holder, iv pole | yes
25 | Availability of all spare parts | yes
26 | Powder-coated frame for durable, high quality, finish | yes
27 | Premium grade upholstery | yes
28 | continuous measuring scale while the patient is undergoing dialysis | Must
29 | Replaceable clear protective cover over foot area | yes
30 | Option to choose the color of the material covering the chair | yes
31 | Dimensions
32 | Dimensions in seat position | L 145 x W 92 cm
33 | Size of upholstery | L 215 x W 60 x H 9,5 cm
34 | Armrest size | L 57 x W 17 cm
35
36 | Maximum patient load | 200 kg - 250 kg
37 | Entry or Reclining from any position | yes
38 | warranty | The vendor is responsible to provide a minimum of three full warranty commencing from the date of Technical acceptance including the motors, shafts, breaks, wheels and frame.
39 | warranty | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
40 | warranty | The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment.
41 | warranty | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
42 | warranty | Voltage: 100 to 240V , 60 Hz
43 | warranty | The vendor to provide installation, set up, calibration, performance verification, preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
44 | warranty | One (1) complete operations manual, to be provided in both hard and soft copies if available.
45 | warranty | One (1) complete service manual (in English language) including schematics and spare parts list (if applicable). The manual is to be provided in both hard and soft copies if available.
46 | warranty | Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
47 | warranty | Full local technical service training to be provided, by a certified trainer, at the time of installation (or at a mutually agreed upon time free of charge) covering the following areas as applicable: a) Theory of Operation including circuit description and schematics b) Hardware installation, assembly, and disassembly c) Software installation and upgrade procedures d) Troubleshooting procedures e) Preventive Maintenance and Calibration procedures f) Overview of consumables, supplies, and accessories
48 | warranty | If the above service training, manuals, and/or software are not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled.
49 | warranty | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
50 | warranty | equipment Downtime: maximum allowable time down is 3% of the total operational time, per annum (calculated based on 44 hrs./week X 52 weeks per annum). In case downtime exceeds this percentage, then the warranty will be extended automatically to twice of the actual down time.
51 | warranty | FDA , SFDA, CE mark approved | required
52 | warranty | the hospital has the right to pick the color combination when the PO is issued
53 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | 112466 | 1193204 | BED PATIENT REGULAR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Trendelenburg and Reverse Trendelenburg function. Please specify
2 | Strong and reliable side rail assembly
3 | 2 Permanent IV poles, Please specify
4 | Max. safe working load 200 Kgs or more, Please specify
5 | Height in low position: 15” (from bed deck), Please specify
6 | Height in high position: 32” (from bed deck), Please specify
7 | Bed head elevation: 60-70 degrees, Please specify
8 | Bed width (sleep deck): 36”,
9 | Bed length: 80” (extendable foot part preferable), Please specify
10 | Night light function
11 | Angle indicators
12 | Close side rail gap for patient safety
13 | Nurse control panel
14 | Patient lockout controls
15 | Wall bumpers
16 | CPR function
17 | Power requirements: 220V
18 | warranty | The vendor is responsible to provide a minimum of ONE (3) year full warranty commencing from the date of Technical acceptance
19 | 1. The vendor is responsible to provide a minimum of ONE (3) year full warranty commencing from the date of Technical acceptance.
20 | 2. The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
21 | 3. a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
22 | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
23 | 4. The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
24 | 5. Voltage: Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted)
25 | Frequency: 60Hz Power plug should be British style medical grade
26 | 6. The vendor to provide: Installation Performance
27 | 7. One (1) complete operations manual, to be provided preferably in both hard and soft copies.
28 | 8. One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
29 | 9. If the vendor fails to provide manuals, as stated above, then equipment acceptance will be put on hold until requirements provided.
30 | 10. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
31 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | 108263 | 1193204 | SPHYGMOMANOMETER ANEROID MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Laser-engraved dial face provides unmatched accuracy | yes
2 | Accurate to ±3 mm Hg | yes
3 | Recessed dial face and high contrast red pointer to increase visibility and reduce parallax errors | yes
4 | Jeweled movement for a longer life | yes
5 | 5-leg stand provides stability and maneuverability | yes
6 | Built-in storage basket desktop model | yes
7 | Latex-free | yes
8 | Mercury-free | yes
9 | Desktop Mobile | yes
10 | All 5 wheels rotate | yes
11 | Included cuffs, adult, pediatric and neonate | yes
12 | warranty | The vendor is responsible to provide a minimum of ONE (1) year full warranty commencing from the date of Technical acceptance.
13 | 1. The vendor is responsible to provide a minimum of ONE (_1_) year full warranty commencing from the date of Technical acceptance.
14 | 2. Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
15 | 3. The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
16 | 4. a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment.
17 | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
18 | 5. The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert.
19 | 6. The vendor to provide:
20 | a) Installation.
21 | b) Set up.
22 | c) Calibration.
23 | d) Performance verification.
24 | e) Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
25 | 7. One (1) complete operations manual, to be provided preferably in both hard and soft copies.
26 | 8. One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
27 | 9. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
28 | 10. The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA).
29 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | 33915 | 1193204 | THERMOMETER DIGITAL WALL MOUNT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Battery Operated | Yes
2 | Probe Cover Storage | Yes
3 | F/C Conversion | Yes
4 | Last Temperature Recall | Yes
5 | Monitor Mode | Yes
6 | Detachable Probe | Yes
7 | Oral / Axillary / Rectal Modes on Each Thermometer | Yes
8 | Color-Coded Removable Probe Well | Yes
9 | Wall-Holder Standard | Yes
10 | Backlight | Yes
11 | Security (Electronic) | Yes
12 | Configurable On-Screen Labeling for Identification of Instrument | Yes
13 | Walking Segments | Yes
14 | Monitor Mode | Yes
15 | Patient Temperature Exceeds Upper Measurement Range Limit of 110.0 °F (43.3 °C) | Yes
16 | Patient Temperature Is Lower than Measurement Range Limit of 80.0 °F (26.7 °C) | Yes
17 | Pulse Timer | Yes
18 | Low Battery | Yes
19 | Security Mode | Yes
20 | warranty | The vendor is responsible to provide a minimum of 2-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel(extended warranty periods if consumable deal is offered).
21 | The vendor is responsible to provide a minimum of 2-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel(extended warranty periods if consumable deal is offered).
22 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
23 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period).
24 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
25 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital.
25 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | 102142 | 1193204 | LIGHT EXAM MOBILE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 28 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | Examination light
2 | CONFIGURATION | Portable
3 | APPLICATIONS | ER, trauma, procedure suites, clinic
4 | Light source | LED,
5 | Output source | 12 VDC
6 | Safety Standard | EN60601, IEC 60601
7 | Device Classification | Clas II Equipment, Type BF
8 | Power Requirement | 100 – 240, 50/60Hz, otherwise 220V
9 | Power Standard | American US standard power Plugs
10 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of ONE (1) year full warranty commencing from the date of first case.
11 | 1.01 Warranty and Technical Support | b) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
12 | 1.01 Warranty and Technical Support | c) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
13 | 1.01 Warranty and Technical Support | d) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
14 | 1.02 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
15 | 1.02 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
16 | 1.02 Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
17 | 1.02 Certification and equipment recalls | d) Vendor to provide Saudi FDA certificate.
18 | 1.03 Power Requirement | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted.
19 | 1.03 Power Requirement | b) The Frequency must be 60Hz.
20 | 1.03 Power Requirement | c) Power plug should be UK style medical grade.
21 | 1.04 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
22 | 1.04 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
23 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | 92029 | 1193204, 1191576 | SCALE PATIENT WITH HEIGHT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 29 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA Clearance
2 | SFDA Clearance
3 | CE Mark
4 | Maximum weight
5 | Measurement Accuracy
6 | Display
7 | Automatic Zeroing
8 | Height Measuring Rod For Patient Scale
9 | Electrical Specification
10 | Power source – auto powered with adaptor with hospital standard plug.
11 | MISCELLANEOUS INFORMATION
12 | Weighting Mechanism | Specify (Electronic load cell is preferred)
13 | Dimensions (H x W x D) | Specify
14 | Platform (W x D) | Specify
15 | Weight | Specify
16 | warranty | The vendor is responsible to provide a minimum of (1) year full warranty commencing from the date of first case
17 | 1.02 Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
18 | 1.02 Software and upgrades | b) original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
19 | 1.03 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
20 | 1.03 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
21 | 1.03 Certification and equipment recalls | c) The Equipment’s should have SFDA certification
22 | 1.03 Certification and equipment recalls | d) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
23 | 1.04 Power Requirement | a) Battery operated.
24 | 1.05 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
25 | 1.05 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
26 | 1.05 Manuals | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
27 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | 347866100 | 1193598 | BLANKET WARMING SYSTEM AIR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CONFIGURATION | Mobile, rolling cart, pole-mounted
2 | APPLICATIONS | Preoperative, OR, PACU, ER, trauma, procedure suites
3 | DISPLAY TYPE | LED, LCD
4 | Hour meter Indicator | Yes
5 | High/low temperature Indicator | Yes
6 | Temperature Indicator | Yes
7 | Over temperature Indicator | Yes
8 | Disposable Blankets | Yes
9 | TEMP SETTINGS, °C | ≤46
10 | SAFETY THERMOSTAT High limit, °C | 53
11 | ALARM CONDITIONS | Overtemperature required,others (please specify)
12 | Alarm Type of indicator | Audible/visual
13 | NOISE LEVEL, dB | Less than 50
14 | WARM-UP TIME, min | 5
15 | Fan Failure Alarm | Yes
16 | FILTRATION SYSTEM | HEPA or better
17 | Unit Supplied Against Consumables | Highly Recommended (please specify in details)
18 | Safety Standard | EN60601, IEC 60601
19 | Device Classification | Clas II Equipment, Type BF
20 | Power Requirement | 100 – 240, 50/60Hz, otherwise 110V
21 | Power Standard | American US standard power Plugs
22 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of ONE (1) year full warranty commencing from the date of first case.
23 | 1.01 Warranty and Technical Support | b) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
24 | 1.01 Warranty and Technical Support | c) The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
25 | 1.01 Warranty and Technical Support | d) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
26 | 1.01 Warranty and Technical Support | e) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
27 | 1.02 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
28 | 1.02 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
29 | 1.02 Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
30 | 1.02 Certification and equipment recalls | d) Vendor to provide Saudi FDA certificate.
31 | 1.03 Power Requirement | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted.
32 | 1.03 Power Requirement | b) The Frequency must be 60Hz.
33 |  | c) Power plug should be UK style medical grade.
34 | 1.04 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
35 | 1.04 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
36 | 1.05 Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
37 | 1.05 Training | b) The vendor shall provide Local service training for ONE service engineers. The training minimum contents are:
38 | 1.05 Training | I. Theory of Operation including circuit description and schematics.
39 | 1.05 Training | II. Hardware installation, assembly, and disassembly.
40 | 1.05 Training | III. Software installation and upgrade procedures.
41 | 1.05 Training | IV. Troubleshooting procedures.
42 | 1.05 Training | V. Preventive Maintenance and Calibration procedures.
43 | 1.05 Training | VI. Overview of consumables, supplies, and accessories.
44 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | 411530210 | 1193758 | CARDIAC OUTPUT NON & MINIMALLY INVASIVE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 31 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | high end units:
2 | FDA CLEARANCE | Yes
3 | CE MARK | Yes
4 | SFDA (MDMA) certificate | Required
5 | Manufacturer | Please specify
6 | Model | Please specify
7 | Standalone Unit | Yes
8 | Battery Backup | Required minimum 30 min
9 | Easy Swappable Batteries | Yes
10 | Modular | Required, interchangeable
11 | Measurement | one-time, periodic and continuous cardiac output monitoring
12 | Data Displayed | CCO and associated parameters, indexed and non-indexed (SV, SVR, RVEF, EDV) oxygen saturation (mixed venous SvO2 or central venous ScvO2), iCO and associated parameters, both indexed and non-indexed (SV, SVR). Required
13 | Alarm | Audible and Visual
14 | Trends | SvO2, ScvO2, continuous CO, CCI, COavg, Clavg, EDV, EDVI, SVR, SVRI, RV EF, SV and SVI
15 | User Interface | HD capacitive touch screen no less than 11”
16 | Integration With Cerner | Required
17 | Video output | HDMI Or DVI required
18 | USB log | Required
19 | low end units:
20 | FDA CLEARANCE | Yes
21 | CE MARK | Yes
22 | SFDA (MDMA) certificate | Required
23 | Manufacturer | Please specify
24 | Model | Please specify
25 | Standalone Unit | Yes
26 | Battery Backup | Required minimum 30 min
27 | Number of channels | Minimum 4
28 | Measurement | one-time, periodic and continuous cardiac output monitoring
29 | Data Displayed | CCO and associated parameters, indexed and non-indexed (SV, SVR, RVEF, EDV) oxygen saturation (mixed venous SvO2 or central venous ScvO2
30 | Alarm | Audible and Visual
31 | Trends | SvO2, ScvO2, continuous CO, CCI, COavg and Clavg
32 | User Interface | HD capacitive touch screen no less than 8”
33 | Integration With Cerner | Required
34 | USB log | Required
35 | General | a) The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
36 | General | b) The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order.
37 | General | c) The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models.
38 | Warranty and Technical Support | The vendor is responsible to provide a minimum of three (3) year full warranty commencing from the date of preliminary acceptance.
39 | Warranty and Technical Support | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
40 | Warranty and Technical Support | f) Response time during warranty should not exceed 8 hour, for working days, or 1 day, for official holidays and weekends, from when notification of failure or problem is received by vendor.
41 | Warranty and Technical Support | g) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
42 | Warranty and Technical Support | h) The vendor to provide: Installation, set up, Calibration, Performance verification and Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
43 | Warranty and Technical Support | i) Service response time for routine calls should be one (2) hour.
44 | Warranty and Technical Support | j) Service engineer must provide a service report each time they evaluate and/or work on the equipment to Clinical Engineering Department.
45 | Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
46 | Software and upgrades | b) If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
47 | Software and upgrades | c) Vendors shall guarantee system upgradeability of all equipment supplied and must state the period of guarantee. Any software updates within the life period are to be incorporated into the installed equipment at no additional charge to the hospital.
48 | Certification and equipment recall | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent).
49 | Certification and equipment recall | b) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
50 | Power Requirement | a) Equipment must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise, 220 V is accepted.
51 | Power Requirement | b) The Frequency must be 60Hz.
52 | Power Requirement | c) Power plug should be UK style medical grade.
53 | Power Requirement | d) ) the above applies to all equipment unless stated in its specification table.
54 | manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
55 | manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
56 | Training | Included with each vendor offer shall include all details of the training scope, location, details provided to department personnel.
57 | Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
58 | Training | b) The vendor shall provide a factory-level technical service training for (1) clinical engineer at the factory or authorized training facility covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital covering, but not limited to, the following Areas as applicable:
59 | Training | I. Theory of Operation including circuit description and schematics
60 | Training | II. Hardware installation, assembly, and disassembly
61 | Training | III. Software installation and upgrade procedures
62 | Training | IV. Troubleshooting procedures
63 | Training | V. Preventive Maintenance and Calibration procedures
64 | Training | VI. Overview of consumables, supplies, and accessories
65 | TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | a) In addition to the requested equipment, the vendor shall provide the any specialized test tools, manuals and/or if applicable service software that used in equipment maintenance and to be delivered at the time of equipment installation.
66 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | 410434502 | 1194260 | Advanced and Portable ECMO Machine |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 32 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA (MDMA)
2 | FDA / CE
3 | Name of Manufacturer | Please specify
4 | Model of the equipment | Please specify
5 | No. of needed equipment | 1
6 | Instrumentation:
7 | A) PUMP TYPE | Centrifugal
8 | B) Integrated Oxygenator and Centrifugal Pump
9 | C) Max rpm | 5000
10 | D) Drive system | Magnet
11 | E) Flow rate, L/min | 0-10L/Min
12 | F) Flow sensor
13 | G) External Pressure sensor
14 | H) Internal Pressure sensor
15 | I) Internal Temperature Sensor
16 | J) Venous oxygen saturation sensor
17 | K) Hemoglobin Sensor
18 | L) Hematocrit Sensor
19 | M) Bubble Sensor
20 | N) Manual crank
21 | O) Interface touch and colored screen
22 | P) Heater unit
23 | Q) Cart +IV pole + Gas cylinder holder
24 | R) Gas blender (Air/O2)
25 | Electrical Requirements:
26 | Change Frequency | 50-60Hz
27 | Grounding | Correctly grounded electrical outlet
28 | Line Voltage | Dual :110-220 VAC OR 220V Accepted
29 | Isolation Transformer (medical Grade)
30 | Battery run time | Minimum 2 hrs.
31 | Included:
32 | Mechanical Circulatory Support Device | Main Component
33 | Pole Infusion Metal | accessory
34 | Heater Unit 230v | Component
35 | Blender Gas | Component
36 | Cart Sprinter | accessory
37 | Shell Top Kit for Main Device | accessory
38 | Tube Water Supply 105 M for Heater Unit | accessory
39 | Holder Gas Bottle | accessory
40 | Holder Water Tube for Heater Unit | accessory
41 | Emergency Drive Holder for Mast System | Component
42 | Holder Mobile Transport HKH 8860 | Component
43 | ARTICLE 1: GENERAL SUMMARY | 1.1 The request for proposal contains the ECMO machine
44 | ARTICLE 1: GENERAL SUMMARY | 1.2 The Bid must include all consumables and servicing costs; both regularly scheduled preventive maintenance as well as emergency maintenance as per manufacturer’s requirement, available on a 24 hour basis. The vendor will cover the cost of all replacement parts and factory upgrades.
45 | ARTICLE 1: GENERAL SUMMARY | 1.3 Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and that:
46 | ARTICLE 1: GENERAL SUMMARY | · the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and
47 | ARTICLE 1: GENERAL SUMMARY | · the equipment meet the standard of electromagnetic compatibility
48 | ARTICLE 1: GENERAL SUMMARY | 1.4 The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as
49 | ARTICLE 1: GENERAL SUMMARY | 1.5 SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
50 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.1 The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
51 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.2 The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order.
52 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.3 The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models.
53 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.4 Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released with same price. To be discussed with contract dept.
54 | ARTICLE 3: TRAINING | 3.1 Included with each vendor offer shall include all details of the training scope, location, details provided to department personnel.
55 | ARTICLE 3: TRAINING | 3.2 The hospital reserves the right not to accept the training and the vendor must provide a replacement at no additional cost to hospital.
56 | ARTICLE 3: TRAINING | 3.3 The vendor must provide two application training sessions; the first visit just after equipment preliminary acceptance and the second visit within six (6) months.
57 | ARTICLE 3: TRAINING | 3.4 The vendor will specify, with advance notification, required for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
58 | ARTICLE 3: TRAINING | 3.5 The vendor shall provide a factory-level technical service training for (1) clinical engineer at the factory covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital covering, but not limited to, the following Areas as applicable:
59 | ARTICLE 3: TRAINING | · Theory of Operation including circuit description and schematics
60 | ARTICLE 3: TRAINING | · Hardware installation, assembly, and disassembly
61 | ARTICLE 3: TRAINING | · Software installation and upgrade procedures
62 | ARTICLE 3: TRAINING | · Troubleshooting procedures
63 | ARTICLE 3: TRAINING | · Preventive Maintenance and Calibration procedures
64 | ARTICLE 3: TRAINING | · Overview of consumables, supplies, and accessories
65 | ARTICLE 4: REFERENCES | 4.1 Vendors must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operates the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities.
66 | ARTICLE 4: REFERENCES | 4.2 KFHS&RC has a right to request a site visit to the reference facilities for further technical evaluation and the vendor has to arrange and cover all associated expenses including visa fees, travel insurance, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital.
67 | ARTICLE 5: SYSTEM’s UPGRADES | 5.1 Vendors shall guarantee system upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the hospital.
68 | ARTICLE 6: TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | 6.1 In addition to the requested equipment, the vendor shall provide the test tools, manuals and/or software should be delivered at the time of equipment installation.
69 | ARTICLE 7: EQUIPMENT WARRANTY | 7.1 The vendor is responsible to provide (3) years full warranty commencing from the date of preliminary acceptance certificate.
70 | ARTICLE 8: EXAMINATION OF SPACE AND SERVICES (when applicable) | 8.1 Vendor personnel shall familiarize themselves with the features and characteristics of the building which will affect the delivery and/or installation of their equipment. The vendor shall check space, height, accessibility to all applicable areas to ensure ease of installation. Vendors must advise if the proposed equipment will NOT fit through a doorway.
71 | ARTICLE 8: EXAMINATION OF SPACE AND SERVICES (when applicable) | 8.2 The vendor shall report inadequacies to the owner immediately to facilitate remedial measures prior to scheduling installation.
72 | ARTICLE 8: EXAMINATION OF SPACE AND SERVICES (when applicable) | 8.3 The vendor shall accept responsibility for error or neglect to familiarize with access, space and service conditions. There will be no storage of furniture or equipment on the job site.
73 | ARTICLE 9: SITE PREPARATION / SCOPE OF WORK | 9.1 The vendor will be responsible for site preparation and pre-installation requirements with coordination of Hospital Project Management. The vendor is responsible for providing any specialized materials to complete pre-installation requirements of the medical equipment.
74 | ARTICLE 10: DELIVERY AND INSTALLATION | 10.1 Delivery and installation shall be coordinated with hospital in accordance with contractual completion period
75 | ARTICLE 10: DELIVERY AND INSTALLATION | 10.2 The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
76 | ARTICLE 10: DELIVERY AND INSTALLATION | 10.3 The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc) for each employee they will have on site so that appropriate hospital gate passes can be issued.
77 | ARTICLE 11: CLEANING UP AND MAKING GOOD DAMAGES | 11.1 The vendor shall keep the premises, including assigned rooms, all passages and auxiliary spaces free from accumulation of waste material or rubbish caused by him at the end of each day.
78 | ARTICLE 11: CLEANING UP AND MAKING GOOD DAMAGES | 11.2 The vendor shall assume full responsibility for all damages to the building plant and previously placed equipment, including gouges and smudges to the owner’s property made by their own personnel and/or sub-contractors.
79 | ARTICLE 11: CLEANING UP AND MAKING GOOD DAMAGES | 11.3 Any resulting repairs by the vendor shall match the adjacent existing finishes and materials. The repairs are to be done by appropriate trade contractors who did the work originally and paid by the vendor.
80 | ARTICLE 11: CLEANING UP AND MAKING GOOD DAMAGES | 11.4 The vendor shall comply with the required pre-infection control procedures.
81 | ARTICLE 12: DOCUMENTATION | 12.1 Two (2) complete operational manuals for hardware and software.
82 | ARTICLE 12: DOCUMENTATION | 12.2 Two (2) service manuals for hardware and software.
83 | ARTICLE 12: DOCUMENTATION | 12.3 Two (2) spare part manuals (if applicable).
84 | ARTICLE 12: DOCUMENTATION | 12.4 Completed schematic diagrams with documentation of onsite modifications (if required).
85 | ARTICLE 12: DOCUMENTATION | 12.5 All above documentation must be provided in English.
86 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | 13.1 The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
87 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | • Sales Literature;
88 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin;
89 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
90 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | • NEMA specifications and/or other international regulatory agencies.
91 | ARTICLE 13: EQUIPMENT ACCEPTANCE TESTING | 13.2 Any deficiencies identified upon preliminary acceptance testing must be rectified by the vendor to the hospital’s satisfaction prior to commencing preliminary acceptance.
92 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | Vendor must provide the following information as it applies to their service organization:
93 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.1 Local service is imperative.
94 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.2 Local service location.
95 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.3 Service response time for routine and after hour calls. The vendor shall define the term ‘Response Time’ and provide their minimum and maximum time limits.
96 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.4 For each piece of equipment specified within this document, list percentage of parts that are critical for normal operation of the equipment that are not available from the local service centre, and would take more than one (1) day to receive from another area.
97 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.5 Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment.
98 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.6 The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced.
99 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.7 Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods stated above will be extended to cover 3 times the actual downtime.
100 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.8 Service response time for routine calls should be one (1) hour.
101 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.9 Service engineer must provide a service report each time they evaluate and/or work on the equipment to Clinical Engineering Department.
102 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.10 All empty boxes, tools, spare parts, defective parts, replacement parts and the like, will be removed from the premises immediately. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc
103 | ARTICLE 14: MANUFACTURER’S SERVICE COMMITMENT | 14.11 Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized and clutter-free).
104 | ARTICLE 15: GENERAL COMMENTS | 15.1 The vendor is requested to fill the technical specification sheets -. Not providing data requested can delay the vendor selection process wherein the hospital does reserve the right to ask a vendor to recompile and re submit the data.
105 | ARTICLE 15: GENERAL COMMENTS | 15.2 If data is not available for a certain requested data field the vendor should not leave the field blank but instead use the term ‘No Data’ in the field. If this term is not used, it will be assumed that the data sheet is incomplete and a resubmission may be requested. The vendor does have the right to reference the field to an attachment to better clarify the entry.
106 | ARTICLE 15: GENERAL COMMENTS | 15.3 The vendor must provide sufficient manufacture documentation to support the data quoted.
107 | ARTICLE 15: GENERAL COMMENTS | 15.4 The vendor should provide additional information that may not have been requested on the data specification sheets that can elaborate on the quoted technology capabilities and clinical applications.
108 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | 310434534 | 1194271 | ECMO MACHINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 33 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA (MDMA)
2 | FDA / CE
3 | Name of Manufacturer | Please specify
4 | Model of the equipment | Please specify
5 | No. of needed equipment | 6
6 | Instrumentation:
7 | A) PUMP TYPE | Centrifugal
8 | B) PUMP Approved usage period | Minimum 14 days, provide certificate for approval.
9 | C) Max rpm | 5000
10 | D) Drive system | Magnet
11 | E) Flow rate, L/min | 0-10L/Min
12 | F) flow sensor
13 | G) Bubble Sensor
14 | H) Manual crank
15 | I) Heater unit
16 | J) Cart +IV pole + Gas cylinder holder
17 | K) Gas blender (Air/O2)
18 | L) Pump Minimal Priming Volume | Less than 35 ml
19 | Electrical Requirements:
20 | Change Frequency | 60Hz
21 | Grounding | UK hospital grade, Correctly grounded electrical outlet
22 | Line Voltage | 220 VAC
23 | Isolation Transformer (medical Grade)
24 | Battery run time | Minimum 2 hrs.
25 | Included:
26 | CONSOLE WITH ICU KIT RFC 230 V / 50 HZ | Main Component
27 | DRIVE UNIT, BLUE | Component
28 | EMERGENCY DRIVE | Component
29 | HEATER UNIT 230V | Component
30 | COUNTRY KIT HU 35 | accessory
31 | Fixation System For Heater Unit (Sprinter Cart) | accessory
32 | sprinter cart | accessory
33 | Top Shelf For Main Console | accessory
34 | Second Handle For Sprinter Cart | accessory
35 | Medical Isolating Transformer Kit 230 V, 50/60 HZ, supply /output | Component
36 | Gas Blender | Component
37 | holder for PLS system | Accessory
38 | CABLE CONNECTING MEDICAL ISOLATING TRANSFORMER KIT | Accessory
39 | TUBING HU35 1.5M X 1 WATER SUPPLY FOR HEATER UNIT | Accessory
40 | warranty | The vendor is responsible to provide (3) years full warranty commencing from the date of preliminary acceptance certificate
41 | ARTICLE 1: GENERAL SUMMARY | 1.1 The request for proposal contains the ECMO machine for Department of Heart center
42 | ARTICLE 1: GENERAL SUMMARY | 1.2 The Bid must include all consumables and servicing costs; both regularly scheduled preventive maintenance as well as emergency maintenance as per manufacturer’s requirement, available on a 24-hour basis. The vendor will cover the cost of all replacement parts and factory upgrades.
43 | ARTICLE 1: GENERAL SUMMARY | 1.3 Each piece of equipment will be described fully and shall be accompanied by manufacturer’s literature, and that:
44 | ARTICLE 1: GENERAL SUMMARY | · the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and
45 | ARTICLE 1: GENERAL SUMMARY | · the equipment meets the standard of electromagnetic compatibility
46 | ARTICLE 1: GENERAL SUMMARY | 1.4 The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as
47 | ARTICLE 1: GENERAL SUMMARY | 1.5 SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall.
48 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.1 The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation.
49 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.2 The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order.
50 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.3 The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models.
51 | ARTICLE 2: TECHNOLOGY ENHANCEMENTS | 2.4 Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released with same price. To be discussed with Contract Dept.
52 | ARTICLE 3: TRAINING | 3.1 Included with each vendor offer shall include all details of the training scope, location, details provided to department personnel.
53 | ARTICLE 3: TRAINING | 3.2 The hospital reserves the right not to accept the training and the vendor must provide a replacement at no additional cost to hospital.
54 | ARTICLE 3: TRAINING | 3.3 The vendor must provide two application training sessions; the first visit just after equipment preliminary acceptance and the second visit within six (6) months.
55 | ARTICLE 3: TRAINING | 3.4 The vendor will specify, with advance notification, required for booking their application training sessions. The vendor will adhere to the predetermined and agreed upon schedule.
56 | ARTICLE 3: TRAINING | 3.5 The vendor shall provide a factory-level technical service training for (1) clinical engineer at the factory covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital covering, but not limited to, the following areas as applicable:
57 | ARTICLE 3: TRAINING | · Theory of Operation including circuit description and schematics
58 | ARTICLE 3: TRAINING | · Hardware installation, assembly, and disassembly
59 | ARTICLE 3: TRAINING | · Software installation and upgrade procedures
60 | ARTICLE 3: TRAINING | · Troubleshooting procedures
61 | ARTICLE 3: TRAINING | · Preventive Maintenance and Calibration procedures
62 | ARTICLE 3: TRAINING | · Overview of consumables, supplies, and accessories
63 | ARTICLE 4: REFERENCES | 4.1 Vendors must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operates the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities.
64 | ARTICLE 4: REFERENCES | 4.2 KFHS&RC has a right to request a site visit to the reference facilities for further technical evaluation and the vendor has to arrange and cover all associated expenses including visa fees, travel insurance, round-trip tickets, transportation, meals, and accommodation – with no additional cost to the Hospital.
65 | ARTICLE 5: SYSTEM’s UPGRADES | 5.1 Vendors shall guarantee system upgradeability of all equipment supplied and must state the period of guarantee. Any system upgrades within the warranty period are to be incorporated into the installed equipment at no additional charge to the hospital.
66 | ARTICLE 6: TEST TOOLS AND/OR ADDITIONAL SOFTWARE PURCHASE | 6.1 In addition to the requested equipment, the vendor shall provide the test tools, manuals and/or software should be delivered at the time of equipment installation.
67 | ARTICLE 7: EQUIPMENT WARRANTY | 7.1 The vendor is responsible to provide (3) years full warranty commencing from the date of preliminary acceptance certificate.
68 | ARTICLE 8: DELIVERY AND INSTALLATION | 8.1 Delivery and installation shall be coordinated with hospital in accordance with contractual completion period
69 | ARTICLE 8: DELIVERY AND INSTALLATION | 8.2 The vendor’s personnel shall use authorized passages only and shall assume all liability if the vendor’s personnel fail to exercise proper care for safety, considering that the place of delivery and installation may be under construction.
70 | ARTICLE 8: DELIVERY AND INSTALLATION | 8.3 The vendor shall provide a copy of personal documentation (name, title, company, iqama number, etc. for each employee they will have on site so that appropriate hospital gate passes can be issued.
71 | ARTICLE 9: DOCUMENTATION | 9.1 Two (2) complete operational manuals for hardware and software.
72 | ARTICLE 9: DOCUMENTATION | 9.2 Two (2) service manuals for hardware and software.
73 | ARTICLE 9: DOCUMENTATION | 9.3 Two (2) spare part manuals (if applicable).
74 | ARTICLE 9: DOCUMENTATION | 9.4 Completed schematic diagrams with documentation of onsite modifications (if required).
75 | ARTICLE 9: DOCUMENTATION | 9.5 All above documentation must be provided in English.
76 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | 10.1 The purpose of this testing is to verify that the performance of the installed equipment is within the parameters of the specifications provided by the manufacturer in the form of:
77 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | • Sales Literature;
78 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | • Technical Bulletin;
79 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | • Product Data Sheets supplied by the manufacturer; and
80 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | • NEMA specifications and/or other international regulatory agencies.
81 | ARTICLE 10: EQUIPMENT ACCEPTANCE TESTING | 10.2 Any deficiencies identified upon preliminary acceptance testing must be rectified by the vendor to the hospital’s satisfaction prior to commencing preliminary acceptance.
82 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | Vendor must provide the following information as it applies to their service organization:
83 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.1 Local service is imperative.
84 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.2 Local service location.
85 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.3 Service response time for routine and after hour calls. The vendor shall define the term ‘Response Time’ and provide their minimum and maximum time limits.
86 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.4 For each piece of equipment specified within this document, list percentage of parts that are critical for normal operation of the equipment that are not available from the local service center, and would take more than one (1) day to receive from another area.
87 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.5 Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment.
88 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.6 The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced.
89 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.7 Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods stated above will be extended to cover 3 times the actual downtime.
90 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.8 Service response time for routine calls should be one (1) hour.
91 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.9 Service engineer must provide a service report each time they evaluate and/or work on the equipment to Clinical Engineering Department.
92 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.10 All empty boxes, tools, spare parts, defective parts, replacement parts and the like, will be removed from the premises immediately. It is not permissible to have such items littering equipment technical rooms, hallways, and/or scanning rooms, etc.
93 | ARTICLE 11: MANUFACTURER’S SERVICE COMMITMENT | 11.11 Service engineers are expected to properly maintain all equipment technical rooms (i.e. clean, well-organized and clutter-free).
94 | ARTICLE 12: GENERAL COMMENTS | 12.1 The vendor is requested to fill the technical specification sheets. Not providing data requested can delay the vendor selection process wherein the hospital does reserve the right to ask a vendor to recompile and re submit the data.
95 | ARTICLE 12: GENERAL COMMENTS | 12.2 If data is not available for a certain requested data field the vendor should not leave the field blank but instead use the term ‘No Data’ in the field. If this term is not used, it will be assumed that the data sheet is incomplete and a resubmission may be requested. The vendor does have the right to reference the field to an attachment to better clarify the entry.
96 | ARTICLE 12: GENERAL COMMENTS | 12.3 The vendor must provide sufficient manufacture documentation to support the data quoted.
97 | ARTICLE 12: GENERAL COMMENTS | The vendor should provide additional information that may not have been requested on the data specification sheets that can elaborate on the quoted technology capabilities and clinical applications.
98 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | 414268810 | 1194570 | MACHINE ANESTHESIA ADVANCED VENTILATION WITH MONITORING SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 34 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Anesthesia machine:
2 | Pipeline gas inlets: | 3 (O2, N2O, air)
3 | Gas cylinder yokes: | O2, N2O
4 | Number of yokes | The machine to accommodate 1 oxygen and 1 nitrous oxide pin index yokes.
5 | Vaporizers/Agents, supported: | Isoflurane, desflurane, sevoflurane and any other new agent that might be introduced in future.
6 | Gas hoses color coding, hoses probes and gas labeling standard | US standard for color coding of hoses and labeling with Ohio quick connect adapter.
7 | Gases hoses length | 3 meters
8 | Number and Type of Vaporizer on each Unit: | 2, Sevoflurane and Desoflurane
9 | Interlock Mechanism: | Yes
10 | SUCTION SYSTEM: | Yes (operate separately through vacuum hose not by Venturi mechanism)
11 | O2 FAIL-SAFE: | Audible, visual, N2O shutoff
12 | HYPOXIC MIXTURE FAIL-SAF: | Yes (methods vary, please specify)
13 | Auxiliary O2 outlet: | Yes, with flow meter
14 | Auxiliary common gas outlet | Yes
15 | Ventilator Type: | High accurate ICU grade mechanical ventilator suitable for Adult, Pediatric & Neonatal
16 | Ventilation modes: | Manual, spontaneous, VCV, PCV, SPN-CPAP/PS, PC-BIPAP/PS, VC-SIMV/PS, PC-APRV.
17 | Tidal volume Range | 20-2000ml on VCV, 5-2000ml on PCV
18 | Minute volume range: | 0.05-26 l/min
19 | Frequency, bpm | 4 to 85
20 | Inspiratory flow, L/min | 1-140 L/min or higher
21 | IE ratio | 2:1 to 4:1,
22 | Inspiratory pause: | Yes 0-50% minimum
23 | Pressure limit, cm H2O | 5 to 80 (adjustable)
24 | PEEP, cm H2O | 0 to 20 (adjustable)
25 | Automatic System checks | Fully Automated self-test on all the system component (electronics and pneumatic) with a full mechanical and manual ventilator leak test and compliance calculation.
26 | Auto Compliance Compensation | Yes
27 | Scavenging System | Open reservoir, passive
28 | Auto Record Keeper | Yes
29 | Airway pressure monitoring | Yes, Insp & Expiratory
30 | Prioritized alarms system: | Advisory, caution, alarm
31 | High-pressure alarm: | Yes
32 | Sub atmospheric pressure alarm: | Yes
33 | Continuing pressure alarm: | Yes
34 | Low pressure/apnea: | Yes
35 | Expiratory volume/flow: | Yes
36 | Type of flow sensor: | U.S. or Heated Wire (if other please specify)
37 | Apnea alarm: | Yes
38 | Reverse-flow alarm: | Yes
39 | High/low minute volume alarm: | Yes
40 | O2 concentration: | Yes, Inspiratory & Expiratory
41 | Type of O2 sensor: | Paramagnetic or higher technology (please specify)
42 | Response time | less than 5 sec.,
43 | CO2 concentration: | Mainstream is recommended
44 | Apnea alarm from CO2: | Yes
45 | CO2 water trap | yes
46 | N2O alarm: | Yes
47 | All Agent monitors: | Yes
48 | Type of agents supported: | Desflurane, Sevoflurane, and any future agent can be added with software upgrade.
49 | Auto Agent ID: | Yes, more than one agent can be identified at the same time
50 | Agent conc. alarm: | Low/hi Insp, and hi Exp,
51 | Physical design: | minimum of two drawers, and writing shelf, mobile unit with lockable casters
52 | Auxiliary power outlets: | At Least 4 with dedicated circuit breaker (American standard) 110V NEMA 5-15 grounded (Type B)
53 | Cable Management Arm/system: | Yes with Swivel side arm.
54 | Back up battery: | Yes, is a must
55 | On Battery use per charge: | Minimum 30 Minutes
56 | Gas Concentration Range, Volume %:
57 | Desflurane | 0-20
58 | Sevoflurane | 0-8
59 | N2O | 0-80
60 | CO2 | 0-10
61 | O2 | 0-100
62 | RISE TIME, msec:
63 | Agent | <500
64 | N2O | <500
65 | CO2 | <400
66 | O2 | <400
67 | ACCURACY:
68 | Agent, volume % | 0.25
69 | N2O, volume % | 5
70 | CO2, mm Hg | 0.4
71 | O2, volume % | 2
72 | CALIBRATION Requirement | Once per Year or more
73 | WATER TRAP VOL, mL | >8
74 | Screen | Moveable, Not less than 15" active-color touch TFT or bette with a knob to confirm selected option.
75 | ALARM LIMITS:
76 | HIGH/LOW, % Agent | 0.1-20, off/0-10, off
77 | HIGH/LOW, % N2O | 20-100, off/1-50, off
78 | HIGH/LOW, % ETCO2 | 1-100 mm Hg, off/0-100 mm Hg, off
79 | HIGH/LOW, % Inspired CO2 | 0-100 mm Hg, off/0-100 mm Hg, off
80 | HIGH/LOW, % Inspired O2 | 18-100, off/18-100
81 | Data Export | Compatible with Cerner, and hospital information system.
82 | Sterilization | Yes, Ventilator to be simple to remove as one part with minimum components to remove.
83 | Patient Monitoring:
84 | Monitor configuration | Must be modular to allow swapping parameters module when needed.
85 | ECG | Should support 3, 5, and 12-lead
86 | Heart Rate | 30-250
87 | St Segment | yes
88 | Memory trending of physiologic parameters | Minimum of 2 hours (Please specify if other)
89 | Noninvasive BP | Reading range: 15 mm hg to ≥ 250 mm hg with large adult, adult, pediatric and neonate cuff sizes
90 | Invasive BP | Not less than 4
91 | Temperature | Not less than 3 (two for skin and one for rectal )
92 | Pulse Oximeter | Not less than two probes (one on hand and other on foot). Nellcor standard. With adult, pediatric and neonate probes.
93 | Other parameters | Must include BIS (anesthesia depth monitor), nerve stimulator. Also preferred to have cerebral oximeter monitor.
94 | Cardiac output | Yes, thermo dilution preferred.
95 | Number of waveforms displayed in the screen. | Monitor must Display Minimum of six physiologic/multiage-analysis waveforms on each display.
96 | Display Type | Colored TFT or LCD with wide viewing angle not less than 19”
97 | Integrated | optional
98 | Interface with Others | Yes
99 | User Interface | Touch screen and with rotary knop
100 | Data Export | Compatible with Cerner, and heart net hospital information system.
101 | Prioritized Alarms | 3 (caution, advisory, alarm)
102 | Physical Features | Must have the compatible mount to fit on the anesthesia machine.
103 | warranty | The vendor is responsible to provide a minimum of three (3) year full warranty commencing from the date of preliminary acceptance.
104 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | 452169101 | 1194845 | MONITOR BODY COMPOSITION (BCM) |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 35 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Usage | Overhydration is common among patients with renal failure and is closely linked to cardiovascular disease, the leading cause of death in dialysis patients. Achieving optimal fluid balance is therefore one of the central challenges in routine dialysis practice.
2 | Overhydration (OH) (pre-/postdialytic) [L] | YES
3 | Lean tissue index (LTI) [kg/m2] | YES
4 | Fat tissue index (FTI) [kg/m2] | YES
5 | Total body water (TBW) (Urea distribution volume V) [L] | YES
6 | Extracellular water (ECW) [L] | YES
7 | Intracellular water (ICW) [L] | YES
8 | Intracellular water (ICW) [L] | YES
9 | Lean tissue mass [kg] and [%] | YES
10 | Fat mass [kg] | YES
11 | Adipose tissue mass [kg] and [%] | YES
12 | Body Cell Mass [kg] | YES
13 | Measurement time approx. 2 min | YES
14 | Data output LC-Display; integrated SmartCard writer | YES
15 | Measuring frequency range 50 discrete frequencies in the range from 5 – 1000 kHz | YES
16 | Battery Lithium-Ion battery, capacity 5 hours | YES
17 | Clinical information screens | YES
18 | AC adapter 100 – 240 V AC; 50 – 60 H, UK plug hospital grade | YES
19 | 1-The vendor is responsible to provide a minimum of Three-years full warranty commencing from the date of Technical acceptance.
20 | 2-The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
21 | 3-The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment.
22 | 4-The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as,SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
23 | 5-The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
24 | 6-Equipment Downtime: maximum allowable time down is 3% of the total operational time, per annum (calculated based on 44 hrs./week X 52 weeks per annum). In case downtime exceeds this percentage, then the warranty will be extended automatically to twice of the actual down time.
25 | 7-The equipment should have SFDA certificate.
26 | 8-The vendor to provide: Installation, Set up Calibration Performance verification Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge.
27 | 9-Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
28 | 10-Full Technical Service training to be provided by qualified Engineer
29 | 11- One (1) complete operations manual, to be provided preferably in both hard and soft copies.
30 | 12- One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
31 | 13- If the vendor fails to provide user training, technical Service training, manuals, as stated above, then equipment acceptance will be put on hold until requirements provided.
32 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | 100611 | 1200563 | LAMP HEATING |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 36 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Mobile single head light on a stable base | YES
2 | Adjustable height and light hood | YES
3 | Radiant heat and light | YES
4 | Radiates invisible infrared heat | YES
5 | Adjustable light hoods | YES
6 | Distance measurer | YES
7 | A 15 minute on/off safety timer and an adjustable intensity control are standard | YES
8 | Fully maneuverable | YES
9 | Warm-Lamp shall have the following technical data: | YES
10 | Incandescent light. 250W | YES
11 | operating volt 220v Ac | YES
12 | warranty | The vendor is responsible to provide a minimum of two-years full warranty commencing from the date of Technical acceptance.
13 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model.
14 | The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment.
15 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall.
16 | Voltage: 220V , 60 Hz
17 | The vendor to provide installation, set up.
18 | One (1) complete operations manual, to be provided in both hard and soft copies if available.
19 | One (1) complete service manual (in English language) including schematics and spare parts list.
20 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | 442271600 | 1210914 | AIRWAY CLEARANCE LUNG EXPANSION SYSTEM THREE IN ONE THERAPY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 37 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MODEL: | Please specify
2 | PARAMETER ACCURACY
3 | Dimensions | Height: 10” (254 mm), Width: 15” (381 mm), Depth: 4” (102 mm)
4 | Weight | ≤ 3 kg, including stand
5 | Mode (exact or similar) and features | Continuous Positive Expiratory Pressure (CPEP) Chest High Frequency Oscillation (CHFO) Aerosol Delivery
6 | Pressures (cm H2O) | 1 to 28
7 | Frequency for CHFO mode range | 0.5 to 5 Hz.
8 | Patient Interface | Mouthpiece, ventilator circuit
9 | Controller (Driving Gas) | Pneumatic and air or oxygen
10 | Controller settings | 1. On/off 2. Mode selection 3. Frequency selection for 4. CHFO mode 5. Pressure adjustment for 6. CPEP mode 7. Pressure manometer 8. Connection for patient circuit
11 | Ventilator connection | Placed in-line in the inspiratory limb of the ventilator circuit with a standard “T” adapter. Only for acute care Environment
12 | Patient Circuit | Disposable circuit includes: connection for in-line nebulizer Works in hospital and home
13 | CHFO Continuous High Frequency Oscillations | Frequencies up to 300 beats per minute and peak positive pressures < 30 cmH2O
14 | CPEP Continuous Positive Expiratory Pressure | < 30 cm H2O
15 | Mode of Operation | Pneumatic system.
16 | Patient Population | 5 years old and above whom
17 | Environments of Use | Hospital, sub-acute facilities, Nursing care, Homecare
18 | Included:
19 | Airway Clearance Lung Expansion System Three In One Therapy | Main component QTY 3
20 | Patient Circuit With Nebulizer For Lung Expansion System Disposable 5 Circuits/ Kit | Accessory QTY 3
21 | Warranty | Warranty (Two Years) including Labor, spare parts and PPM kits
22 | - Operating voltage should be auto 110 or 220 VAC. | Required
23 | - Power Plug type - set of American and British | Required
24 | - Vendor to provide service manual | Required
25 | - The vendor shall provide local service training clinical engineering covering the following: | Required
26 | a) Theory of Operation including circuit description and schematics
27 | b) Hardware installation, assembly, and disassembly
28 | c) Software installation and upgrade procedures
29 | d) Troubleshooting procedures
30 | e) Preventive Maintenance and Calibration procedures
31 | f) Overview of consumables, supplies, and accessories
32 | - The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production | Required
33 | - The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent). | Required
34 | - The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | Required
35 | - Authority Clarification | SFDA Required
36 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | 1157162 | 1157162 | Endoscopic Video System For Difficult Intubation |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 38 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MONITOR SPECIFICATION:
2 | Resolution, At least 1280 X 800 Pixel or more | Required
3 | Color, 256 or more | Required
4 | Size, 7" and above | Required
5 | Type: OLED Or TFT | Required
6 | Inputs, HDMI, USB | Required
7 | Two Camera Inputs | Required
8 | Media Storage for real time capturing | Required
9 | IP4 Splash proofing | Required
10 | Dual Voltage 110 – 230 V | Required
11 | Frequency : Auto 50Hz/60Hz | Required
12 | Internal Rechargeable Battery | Required
13 | Battery Type: Lithium Ion | Required
14 | IMAGER SPECIFICATION:
15 | Immersion Protection IPX8 | Required
16 | Camera Technology CMOS | Required
17 | Resolution: 640 X 480 Pixel or More | Required
18 | Illumination, White | Required
19 | Disinfection temp. up to 65 °C | Required
20 | Compatible with Offered Monitor | Required
21 | Compatible with offered Blades | Required
22 | FLEXIBLE SCOPE SPECIFICATION:
23 | Camera Technology CMOS | Required
24 | Suction Valve | Required
25 | Compatible with offered Monitor | Required
26 | Deflection up/down : 140° / 140° | Required
27 | Viewing Angle 100° | Required
28 | Working Length between 60 – 65 cm | Required
29 | Total Length: between 90 – 100 cm | Required
30 | Working channel inner diameter 1.5 mm | Required
31 | Working channel Distal tip outer dia. 4mm | Required
32 | Integrated Light Source | Required
33 | Light type : LED | Required
34 | Light Channels: TWO | Required
35 | TROLLY SPECIFICATION:
36 | Wheels : 4 antistatic | Required
37 | Locking brakes : 2 Wheels equipped | Required
38 | Equipment Rail | Required
39 | Tube 100cm adjustable | Required
40 | Material: Stainless steel | Required
41 | Drawers | Required
42 | Stainless steel round pip | Required
43 | RIGID SCOPE SPECIFICATION:
44 | Camera Technology CMOS | Required
45 | Suction Valve | Required
46 | Compatible with offered Monitor | Required
47 | Deflectable Tip | Required
48 | Video Documentation Button | Required
49 | Disinfect able | Required
50 | Viewing angle upto 65° | Required
51 | IPX8 | Required
52 | Angulation tip upto 60° | Required
53 | 1.01 Warranty and Technical Support | a) The vendor is responsible to provide a minimum of ONE (1) year full warranty commencing from the date of first case.
54 | 1.01 Warranty and Technical Support | b) Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime.
55 | 1.01 Warranty and Technical Support | c) Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor.
56 | 1.01 Warranty and Technical Support | d) The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced.
57 | 1.01 Warranty and Technical Support | e) The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge.
58 | 1.01 Warranty and Technical Support | f) The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period.
59 | 1.02 Software and upgrades | a) The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge.
60 | 1.02 Software and upgrades | b) If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge.
61 | 1.03 Certification and equipment recalls | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent).
62 | 1.03 Certification and equipment recalls | b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility.
63 | 1.03 Certification and equipment recalls | c) The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall.
64 | 1.04 Power Requirement | a) Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V)
65 | 1.04 Power Requirement | b) The Frequency must be 60Hz.
66 | 1.04 Power Requirement | c) Power plug should be British style medical grade.
67 | 1.05 Manuals | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies.
68 | 1.05 Manuals | b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies.
69 | 1.06 Training | a) Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge.
70 | 1.06 Training | b) The vendor shall provide factory-level service training for ONE service engineers, at the factory or authorized training facility covering all associated expenses including visa fees, travel insurance, tuition, round-trip tickets, transportation, meals, and accommodation, with no additional cost to the Hospital. The training minimum contents are:
71 | 1.06 Training | I. Theory of Operation including circuit description and schematics.
72 | 1.06 Training | II. Hardware installation, assembly, and disassembly.
73 | 1.06 Training | III. Software installation and upgrade procedures.
74 | 1.06 Training | IV. Troubleshooting procedures.
75 | 1.06 Training | V. Preventive Maintenance and Calibration procedures.
76 | 1.06 Training | Overview of consumables, supplies, and accessories.
77 | NOTE: | All training sessions, visits and factory level training fees must be quoted (itemized) and not included with the main device.
Company Stamp |  |  |  | Signature