## Index
INDEX
Summary | This sheet contains the summary of all offers from the supplier/bidder.
ITEM 1 | This sheet contains the details of the bid with each technical specification for | ITEM 1 | MONITOR TRANSCUTANEOUS OXYGEN-CARBON DIOXIDE
ITEM 2 | This sheet contains the details of the bid with each technical specification for | ITEM 2 | CRITICAL CARE VENTILATOR NIV FOR ADULT/PEDIATRIC PATIENT MODEL V680
ITEM 3 | This sheet contains the details of the bid with each technical specification for | ITEM 3 | TELESCOPE 0 DEGREE HOPKINS II STRAIGHT FORWARD ENLARGED VIEW DIAMETER 10 MM LENGTH 31 CM AUTOCLAVABLE FOR INDOCYANIN GREEN ICG AND AF FIBER OPTIC LIGHT TRANSMISSION
ITEM 4 | This sheet contains the details of the bid with each technical specification for | ITEM 4 | TELESCOPE HOPKINS II 30 DEGREES 10 MM ENLARGED VIEW 31CM, AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED # 26003BA
ITEM 5 | This sheet contains the details of the bid with each technical specification for | ITEM 5 | TELESCOPE 30 DEGREE 5MM LENGTH 29CM AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE RED 26046BA
ITEM 6 | This sheet contains the details of the bid with each technical specification for | ITEM 6 | TELESCOPE HOPKINS II 0 DEGREE STRAIGHT FORWARD ENLARGED VIEW DIAMETER 3.3MM LENGTH 25CM AUTOCLAVABLE FIBRE OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE: GREEN
ITEM 7 | This sheet contains the details of the bid with each technical specification for | ITEM 7 | TELESCOPE HOPKINS II 3.3MM STRAIGHT FORWARD 30 DEGREE LENGTH 25CM COLOR CODE RED AUTOCLAVABLE
ITEM 8 | This sheet contains the details of the bid with each technical specification for | ITEM 8 | VIDEO GASTROSCOPE SLIM 90K WORKING LENGTH 1050MM OUTER DIAMETER 8.0MM FIELD OF VIEW 140DEG
ITEM 9 | This sheet contains the details of the bid with each technical specification for | ITEM 9 | VIDEO COLONOSCOPE SLIM HD+ WORKING LENGTH 1700MM OUTER DIAMETER 11.6MM FIELD OF VIEW 140DEG
ITEM 10 | This sheet contains the details of the bid with each technical specification for | ITEM 10 | VIDEO GASTROSCOPE DOUBLE CHANNEL WORKING LENGTH 1050MM OUTER DIAMETER 12.8MM FIELD OF VIEW 140DEG
ITEM 11 | This sheet contains the details of the bid with each technical specification for | ITEM 11 | VIDEO COLONOSCOPE DOUBLE CHANNEL WORKING LENGTH 1700MM OUTER DIAMETER 13.2MM FIELD OF VIEW 140DEG
ITEM 12 | This sheet contains the details of the bid with each technical specification for | ITEM 12 | VIDEO COLONOSCOPE HGH DEFINITION DEC38- I10L SERIES 13.2MM/3.8MM
ITEM 13 | This sheet contains the details of the bid with each technical specification for | ITEM 13 | GASTROSCOPE HIGH DEFINITION VIDEO DEG29- I10 SERIES 9.8MM /3.2MM
ITEM 14 | This sheet contains the details of the bid with each technical specification for | ITEM 14 | VIDEO PROCESSOR HIGH DEFINITION DEPK-17000 SERIES
ITEM 15 | This sheet contains the details of the bid with each technical specification for | ITEM 15 | VIDEO BRONCHOSCOPE 5.2MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.0MM FIELD OF VIEW 120DEG
ITEM 16 | This sheet contains the details of the bid with each technical specification for | ITEM 16 | VIDEO BRONCHOSCOPE 6.4MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.8MM FIELD OF VIEW 120DEG
ITEM 17 | This sheet contains the details of the bid with each technical specification for | ITEM 17 | INSUFFLATOR CO2 UNIT
ITEM 18 | This sheet contains the details of the bid with each technical specification for | ITEM 18 | ENDOSCOPIC IRRIGATION PUMP DIMENSIONS 121MM X 197MM X 349MM
ITEM 19 | This sheet contains the details of the bid with each technical specification for | ITEM 19 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 27-INCH
ITEM 20 | This sheet contains the details of the bid with each technical specification for | ITEM 20 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 32-INCH
ITEM 21 | This sheet contains the details of the bid with each technical specification for | ITEM 21 | VIDEO DUODENOSCOPE 11.6MM TUBE WORKING LENGTH 1250MM INSTRUMENT CHANNEL 4.2MM FIELD OF VIEW 100DEG
ITEM 22 | This sheet contains the details of the bid with each technical specification for | ITEM 22 | VIDEO DUODENOSCOPE HARD TIP 12.5MM SOFT TIP 10.8MM CHANNEL DIAMETER 3.2MM FIELD OF VIEW 100DEG
ITEM 23 | This sheet contains the details of the bid with each technical specification for | ITEM 23 | ELECTROSURGICAL AND ARGON COAGULATION UNIT
ITEM 24 | This sheet contains the details of the bid with each technical specification for | ITEM 24 | GAMMA DETECTION SYSTEM CONTROL UNIT WITH SOFTWARE
ITEM 25 | This sheet contains the details of the bid with each technical specification for | ITEM 25 | PROBE ELECTRONIC BLUETOOTH STRAIGHT 14MM GAMMA DETECTION SYSTEM CONTROL UNIT
ITEM 26 | This sheet contains the details of the bid with each technical specification for | ITEM 26 | DERMATOME ELECTRIC SLIMLINE MODEL S COMPLETE KIT UNIVERSAL POWER SUPPLY 110-230 VOLTS, # 3539700
ITEM 27 | This sheet contains the details of the bid with each technical specification for | ITEM 27 | OPHTHALMIC FUNDUS IMAGING PHOTO AND ANGIOGRAPHY
ITEM 28 | This sheet contains the details of the bid with each technical specification for | ITEM 28 | OPTICAL BIOMETRIC
ITEM 29 | This sheet contains the details of the bid with each technical specification for | ITEM 29 | SCANNER OPHTHALMIC VISUAL FIELD ANALAYZER 3
ITEM 30 | This sheet contains the details of the bid with each technical specification for | ITEM 30 | MICROSCOPE OPHTHALMOLOGY SURGICAL ON FLOORSTAND WITHOUT ILLUMINATION UNIT
ITEM 31 | This sheet contains the details of the bid with each technical specification for | ITEM 31 | AUTOREFRACTOR KERATOMETER UNIT
ITEM 32 | This sheet contains the details of the bid with each technical specification for | ITEM 32 | EXAMINATION CHAIR MOTORIZED RECLINE
ITEM 33 | This sheet contains the details of the bid with each technical specification for | ITEM 33 | ADVANCED SLIT LAMP SOPHISTICATED BIO MICROSCOPY AND LED POWER, INCLINED ADAPTER AND STEREO VARIATOR WITH COMPLETE SYSTEM ACCESSORIES SET
ITEM 34 | This sheet contains the details of the bid with each technical specification for | ITEM 34 | VISUAL ACUITY SYSTEM
ITEM 35 | This sheet contains the details of the bid with each technical specification for | ITEM 35 | DETECTOR AUTO-LENSMASTER LENS TYPE DTECTION UV MEASUREMENT 5.7-INCH COLOR LCD WITH TOUCH SCREEN
ITEM 36 | This sheet contains the details of the bid with each technical specification for | ITEM 36 | SCANNER OPHTHALMIC DIAGNOSTIC WITH INTEGRATED AXIAL LENGTH MEASURMENT
ITEM 37 | This sheet contains the details of the bid with each technical specification for | ITEM 37 | TONOMETER NON-CONTACT WITH ADJUSTABLE LCD SCREEN
ITEM 38 | This sheet contains the details of the bid with each technical specification for | ITEM 38 | PHACOEMULSIFICATION MACHINE CENTURION VISION SYSTEM
ITEM 39 | This sheet contains the details of the bid with each technical specification for | ITEM 39 | VITRECTOMY MACHINE VISION SYSTEM LXT INCLUDES EMBEDDED LASER
ITEM 40 | This sheet contains the details of the bid with each technical specification for | ITEM 40 | OPHTHALMIC MULTISPOT LASER SYSTEM FOR RETINAL DISORDERS
ITEM 41 | This sheet contains the details of the bid with each technical specification for | ITEM 41 | OPHTHALMIC COMBINED LASER MACHINE ND:YAG PLUS ARGON-LASER TABLETOP
ITEM 42 | This sheet contains the details of the bid with each technical specification for | ITEM 42 | ULTRASOUND OPHTHALMIC A/B SCAN
ITEM 43 | This sheet contains the details of the bid with each technical specification for | ITEM 43 | OPHTHALMIC CHAIR PROCEDURE
ITEM 44 | This sheet contains the details of the bid with each technical specification for | ITEM 44 | OPTICAL COHERENCE TOMOGRAPHY A-SCAN WITH FUNDUS CAMERA
ITEM 45 | This sheet contains the details of the bid with each technical specification for | ITEM 45 | TONOMETER BATTERY OPERATED MEASURES INTRAOCULAR PRESSURE IOP PORTABLE
ITEM 46 | This sheet contains the details of the bid with each technical specification for | ITEM 46 | OPHTHALMIC CLASSIC SLIT LAMP PORTABLE
ITEM 47 | This sheet contains the details of the bid with each technical specification for | ITEM 47 | SEAT BALANCE ADVANCE BRAKE WITH HEIGHT ADJUSTMENT VIA FOOT ACTIVATION AND A BRAKE
ITEM 48 | This sheet contains the details of the bid with each technical specification for | ITEM 48 | PATIENT OPHTHALMOLOGY SURGICAL OPERATING TABLE
ITEM 49 | This sheet contains the details of the bid with each technical specification for | ITEM 49 | ANALYZER WITH THE CAPACITY FOR ION-SELECTIVE ELECTRODE (ISE) DETERMINATION OF SODIUM POTASSIUM, AND CHLORIDE IN SERUM, PLASMA, AND URINE
ITEM 50 | This sheet contains the details of the bid with each technical specification for | ITEM 50 | ANALYZER WITH ELECTROCHEMILUMINESCENCE (ECL) TECHNOLOGY FOR IMMUNOASSAY ANALYSIS
ITEM 51 | This sheet contains the details of the bid with each technical specification for | ITEM 51 | ANALYZER EVOLUTION COAGULATION
ITEM 52 | This sheet contains the details of the bid with each technical specification for | ITEM 52 | ANALYZER DIFFERENTIAL FULLY INTEGRATED IPU INCLUDING LCD COLOUR TOUCHSCREEN
ITEM 53 | This sheet contains the details of the bid with each technical specification for | ITEM 53 | FLOW HOOD LAMINAR
ITEM 54 | This sheet contains the details of the bid with each technical specification for | ITEM 54 | TABLET COUNTER MACHINE
ITEM 55 | This sheet contains the details of the bid with each technical specification for | ITEM 55 | MACHINE FLUIDOSE SERIES 5 MPI UNIT DOSE PACKAGING SYSTEM
ITEM 56 | This sheet contains the details of the bid with each technical specification for | ITEM 56 | SUCTION UNIT PORTABLE
ITEM 57 | This sheet contains the details of the bid with each technical specification for | ITEM 57 | WARMER UNIT BLOOD FLUID
ITEM 58 | This sheet contains the details of the bid with each technical specification for | ITEM 58 | MEDICATION FRIDGE 5.5 CUBIC FEE
ITEM 59 | This sheet contains the details of the bid with each technical specification for | ITEM 59 | MEDICATION FRIDGE LARGE
ITEM 60 | This sheet contains the details of the bid with each technical specification for | ITEM 60 | MEDICATION FRIDGE 11.5 CUBIC FEET
ITEM 61 | This sheet contains the details of the bid with each technical specification for | ITEM 61 | REFRIGERATOR SMALL MEDICAL-GRADE
ITEM 62 | This sheet contains the details of the bid with each technical specification for | ITEM 62 | WHEELCHAIR MULTIFUNCTION
ITEM 63 | This sheet contains the details of the bid with each technical specification for | ITEM 63 | WHEELCHAIR BARIATRIC

## Summary
SN | CODE | ITEM DESCRIPTION | GROUP NUMBER | CATEGORY | QUANTITY | SUPPLIER | SUPPLIER CODE | MANUFACTURER | COUNTRY OF ORIGIN | MODEL | MANUFACTURER CATALOGUE NUMBER | Quantity Quoted | Unit Price (SR) | Unit Price In Writing (SR) | Unit Price (SR)( including vat if applicable ) | Total Price (SR) including Vat for quoted quantity | COMPANY COMMENTS/ REMARKS: | delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date
1 | 447061100 | MONITOR TRANSCUTANEOUS OXYGEN-CARBON DIOXIDE |  | Medical Eqpt | 2 | [10 empty] | 0
2 | 442272201 | CRITICAL CARE VENTILATOR NIV FOR ADULT/PEDIATRIC PATIENT MODEL V680 |  | Medical Eqpt | 2 | [10 empty] | 0
3 | 413286112 | TELESCOPE 0 DEGREE HOPKINS II STRAIGHT FORWARD ENLARGED VIEW DIAMETER 10 MM LENGTH 31 CM AUTOCLAVABLE FOR INDOCYANIN GREEN ICG AND AF FIBER OPTIC LIGHT TRANSMISSION |  | Medical Eqpt | 2 | [10 empty] | 0
4 | 499900515 | TELESCOPE HOPKINS II 30 DEGREES 10 MM ENLARGED VIEW 31CM, AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED # 26003BA |  | Medical Eqpt | 4 | [10 empty] | 0
5 | 413286117 | TELESCOPE 30 DEGREE 5MM LENGTH 29CM AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE RED 26046BA |  | Medical Eqpt | 6 | [10 empty] | 0
6 | 411637814 | TELESCOPE HOPKINS II 0 DEGREE STRAIGHT FORWARD ENLARGED VIEW DIAMETER 3.3MM LENGTH 25CM AUTOCLAVABLE FIBRE OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE: GREEN |  | Medical Eqpt | 2 | [10 empty] | 0
7 | 453298105 | TELESCOPE HOPKINS II 3.3MM STRAIGHT FORWARD 30 DEGREE LENGTH 25CM COLOR CODE RED AUTOCLAVABLE |  | Medical Eqpt | 2 | [10 empty] | 0
8 | 411633221 | VIDEO GASTROSCOPE SLIM 90K WORKING LENGTH 1050MM OUTER DIAMETER 8.0MM FIELD OF VIEW 140DEG |  | Medical Eqpt | 1 | [10 empty] | 0
9 | 411633222 | VIDEO COLONOSCOPE SLIM HD+ WORKING LENGTH 1700MM OUTER DIAMETER 11.6MM FIELD OF VIEW 140DEG |  | Medical Eqpt | 1 | [10 empty] | 0
10 | 411633223 | VIDEO GASTROSCOPE DOUBLE CHANNEL WORKING LENGTH 1050MM OUTER DIAMETER 12.8MM FIELD OF VIEW 140DEG |  | Medical Eqpt | 1 | [10 empty] | 0
11 | 411633224 | VIDEO COLONOSCOPE DOUBLE CHANNEL WORKING LENGTH 1700MM OUTER DIAMETER 13.2MM FIELD OF VIEW 140DEG |  | Medical Eqpt | 1 | [10 empty] | 0
12 | 413614702 | VIDEO COLONOSCOPE HGH DEFINITION DEC38- I10L SERIES 13.2MM/3.8MM |  | Medical Eqpt | 2 | [10 empty] | 0
13 | 413614701 | GASTROSCOPE HIGH DEFINITION VIDEO DEG29- I10 SERIES 9.8MM /3.2MM |  | Medical Eqpt | 2 | [10 empty] | 0
14 | 413614700 | VIDEO PROCESSOR HIGH DEFINITION DEPK-17000 SERIES |  | Medical Eqpt | 1 | [10 empty] | 0
15 | 414249702 | VIDEO BRONCHOSCOPE 5.2MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.0MM FIELD OF VIEW 120DEG |  | Medical Eqpt | 1 | [10 empty] | 0
16 | 414249703 | VIDEO BRONCHOSCOPE 6.4MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.8MM FIELD OF VIEW 120DEG |  | Medical Eqpt | 1 | [10 empty] | 0
17 | 411637910 | INSUFFLATOR CO2 UNIT |  | Medical Eqpt | 1 | [10 empty] | 0
18 | 411640201 | ENDOSCOPIC IRRIGATION PUMP DIMENSIONS 121MM X 197MM X 349MM |  | Medical Eqpt | 1 | [10 empty] | 0
19 | 411637911 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 27-INCH |  | Medical Eqpt | 1 | [10 empty] | 0
20 | 411637912 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 32-INCH |  | Medical Eqpt | 1 | [10 empty] | 0
21 | 411633225 | VIDEO DUODENOSCOPE 11.6MM TUBE WORKING LENGTH 1250MM INSTRUMENT CHANNEL 4.2MM FIELD OF VIEW 100DEG |  | Medical Eqpt | 1 | [10 empty] | 0
22 | 411633226 | VIDEO DUODENOSCOPE HARD TIP 12.5MM SOFT TIP 10.8MM CHANNEL DIAMETER 3.2MM FIELD OF VIEW 100DEG |  | Medical Eqpt | 1 | [10 empty] | 0
23 | 411638813 | ELECTROSURGICAL AND ARGON COAGULATION UNIT |  | Medical Eqpt | 1 | [10 empty] | 0
24 | 411555406 | GAMMA DETECTION SYSTEM CONTROL UNIT WITH SOFTWARE |  | Medical Eqpt | 1 | [10 empty] | 0
25 | 411555405 | PROBE ELECTRONIC BLUETOOTH STRAIGHT 14MM GAMMA DETECTION SYSTEM CONTROL UNIT |  | Medical Eqpt | 1 | [10 empty] | 0
26 | 442295207 | DERMATOME ELECTRIC SLIMLINE MODEL S COMPLETE KIT UNIVERSAL POWER SUPPLY 110-230 VOLTS, # 3539700 |  | Medical Eqpt | 2 | [10 empty] | 0
27 | 414256513 | OPHTHALMIC FUNDUS IMAGING PHOTO AND ANGIOGRAPHY |  | Medical Eqpt | 1 | [10 empty] | 0
28 | 442183003 | OPTICAL BIOMETRIC |  | Medical Eqpt | 1 | [10 empty] | 0
29 | 414256512 | SCANNER OPHTHALMIC VISUAL FIELD ANALAYZER 3 |  | Medical Eqpt | 1 | [10 empty] | 0
30 | 411640503 | MICROSCOPE OPHTHALMOLOGY SURGICAL ON FLOORSTAND WITHOUT ILLUMINATION UNIT |  | Medical Eqpt | 1 | [10 empty] | 0
31 | 414253300 | AUTOREFRACTOR KERATOMETER UNIT |  | Medical Eqpt | 2 | [10 empty] | 0
32 | 442183007 | EXAMINATION CHAIR MOTORIZED RECLINE |  | Medical Eqpt | 4 | [10 empty] | 0
33 | 442183005 | ADVANCED SLIT LAMP SOPHISTICATED BIO MICROSCOPY AND LED POWER, INCLINED ADAPTER AND STEREO VARIATOR WITH COMPLETE SYSTEM ACCESSORIES SET |  | Medical Eqpt | 4 | [10 empty] | 0
34 | 414283905 | VISUAL ACUITY SYSTEM |  | Medical Eqpt | 6 | [10 empty] | 0
35 | 411541200 | DETECTOR AUTO-LENSMASTER LENS TYPE DTECTION UV MEASUREMENT 5.7-INCH COLOR LCD WITH TOUCH SCREEN |  | Medical Eqpt | 2 | [10 empty] | 0
36 | 414256510 | SCANNER OPHTHALMIC DIAGNOSTIC WITH INTEGRATED AXIAL LENGTH MEASURMENT |  | Medical Eqpt | 1 | [10 empty] | 0
37 | 410442902 | TONOMETER NON-CONTACT WITH ADJUSTABLE LCD SCREEN |  | Medical Eqpt | 2 | [10 empty] | 0
38 | 411639800 | PHACOEMULSIFICATION MACHINE CENTURION VISION SYSTEM |  | Medical Eqpt | 1 | [10 empty] | 0
39 | 411639801 | VITRECTOMY MACHINE VISION SYSTEM LXT INCLUDES EMBEDDED LASER |  | Medical Eqpt | 1 | [10 empty] | 0
40 | 414256509 | OPHTHALMIC MULTISPOT LASER SYSTEM FOR RETINAL DISORDERS |  | Medical Eqpt | 1 | [10 empty] | 0
41 | 414256514 | OPHTHALMIC COMBINED LASER MACHINE ND:YAG PLUS ARGON-LASER TABLETOP |  | Medical Eqpt | 1 | [10 empty] | 0
42 | 414256508 | ULTRASOUND OPHTHALMIC A/B SCAN |  | Medical Eqpt | 1 | [10 empty] | 0
43 | 411548111 | OPHTHALMIC CHAIR PROCEDURE |  | Medical Eqpt | 1 | [10 empty] | 0
44 | 414256511 | OPTICAL COHERENCE TOMOGRAPHY A-SCAN WITH FUNDUS CAMERA |  | Medical Eqpt | 1 | [10 empty] | 0
45 | 410442903 | TONOMETER BATTERY OPERATED MEASURES INTRAOCULAR PRESSURE IOP PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
46 | 410442702 | OPHTHALMIC CLASSIC SLIT LAMP PORTABLE |  | Medical Eqpt | 2 | [10 empty] | 0
47 | 411548112 | SEAT BALANCE ADVANCE BRAKE WITH HEIGHT ADJUSTMENT VIA FOOT ACTIVATION AND A BRAKE |  | Non-medical Eqpt | 7 | [10 empty] | 0
48 | 411548113 | PATIENT OPHTHALMOLOGY SURGICAL OPERATING TABLE |  | Medical Eqpt | 2 | [10 empty] | 0
49 | 411689102 | ANALYZER WITH THE CAPACITY FOR ION-SELECTIVE ELECTRODE (ISE) DETERMINATION OF SODIUM POTASSIUM, AND CHLORIDE IN SERUM, PLASMA, AND URINE |  | Medical Eqpt | 1 | [10 empty] | 0
50 | 411689103 | ANALYZER WITH ELECTROCHEMILUMINESCENCE (ECL) TECHNOLOGY FOR IMMUNOASSAY ANALYSIS |  | Medical Eqpt | 1 | [10 empty] | 0
51 | 411689104 | ANALYZER EVOLUTION COAGULATION |  | Medical Eqpt | 1 | [10 empty] | 0
52 | 411689105 | ANALYZER DIFFERENTIAL FULLY INTEGRATED IPU INCLUDING LCD COLOUR TOUCHSCREEN |  | Medical Eqpt | 1 | [10 empty] | 0
53 | 438793100 | FLOW HOOD LAMINAR |  | Medical Eqpt | 4 | [10 empty] | 0
54 | 414255002 | TABLET COUNTER MACHINE |  | Medical Eqpt | 4 | [10 empty] | 0
55 | 441827101 | MACHINE FLUIDOSE SERIES 5 MPI UNIT DOSE PACKAGING SYSTEM |  | Medical Eqpt | 2 | [10 empty] | 0
56 | 447190100 | SUCTION UNIT PORTABLE |  | Medical Eqpt | 30 | [10 empty] | 0
57 | 417098109 | WARMER UNIT BLOOD FLUID |  | Medical Eqpt | 4 | [10 empty] | 0
58 | 411661003 | MEDICATION FRIDGE 5.5 CUBIC FEE |  | Non-medical Eqpt | 4 | [10 empty] | 0
59 | 417843000 | MEDICATION FRIDGE LARGE |  | Non-medical Eqpt | 6 | [10 empty] | 0
60 | 93196 | MEDICATION FRIDGE 11.5 CUBIC FEET |  | Non-medical Eqpt | 4 | [10 empty] | 0
61 | 411661007 | REFRIGERATOR SMALL MEDICAL-GRADE |  | Medical Eqpt | 4 | [10 empty] | 0
62 | 432359100 | WHEELCHAIR MULTIFUNCTION |  | Medical Eqpt | 50 | [10 empty] | 0
63 | 410470200 | WHEELCHAIR BARIATRIC |  | Medical Eqpt | 10 | [10 empty] | 0
|  | CATEGORY | NO. OF Items Offered | Total Amount of Offers | General Terms and Conditions/ Remarks from Supplier |  |  |  |  | IMPORTANT NOTE TO BIDDER
|  | Medical Equipment | 0 | 0 | [5 empty] | The model, manufacturer and other details for each individual component should be specified in the offer. Please refer to the tender terms and conditions for more details.
|  | Non-medical equipment | 0 | 0 | [5 empty] | If you have any clarification or need assistance to fill up this tender file, please contact us at 920018184 extn 1064 or send email to med.equipment@nupco.com .
|  |  | 0 | 0 | [5 empty] | مطلوب تعبئة الحقول أعلاه جميعها وسيتم صرف النظر عن أي عرض لا يلتزم بذلك

## ITEM 1
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
1 | 447061100 | 0 | MONITOR TRANSCUTANEOUS OXYGEN-CARBON DIOXIDE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA/SFDA CLEARANCE | YES
2 | Optional CE MARK (MDD) | YES
3 | CONFIGURATION: Heated humidifier with integrated flow generator; optional O2 inlet kit | YES
4 | CLINICAL APPLICATIONS: High-flow nasal cannula therapy | YES
5 | PATIENT TYPE: Adult, pediatric, neonatal | YES
6 | FLOW RANGE, LPM: Minimum of 2-40 for adults, 2-20 for infants, 2-25 pediatric | YES
7 | FiO2 CONTROL RANGE: 21-100% | YES
8 | Control method: Internal, electronic | YES
9 | PATIENT INTERFACE OPTIONS: high- and low-flow nasal cannulas | YES
10 | TEMPERATURE ADJUSTMENT RANGE, °C: Minimum of 33-40 | YES
11 | HEATED CIRCUIT CAPABILITY: Required | YES
12 | WATER LEVEL DETECTION: Required | YES
13 | DISPLAY: Digital LED | YES
14 | AUDIBLE/VISUAL TEMPERATURE ALARMS | YES
15 | AUDIBLE/VISUAL WATER LEVEL ALARMS | YES
16 | AUDIBLE/VISUAL OXYGEN LEVEL ALARMS | YES
17 | OTHER ALARMS: Power-out, leak, blockage, breathing tube fault, flow error, temperature error | YES
18 | NURSE CALL CONNECTIVITY: Optional | YES
19 | PATIENT CIRCUIT REQUIREMENTS: Proprietary disposable water path, cannulas, tubing sets, and heated wire circuits | YES
20 | MOUNTING: Rear-mounted clamp for approved roll stands | YES
21 | POWER REQUIREMENTS: 240 VAC, 50/60 Hz | YES
22 | BATTERY BACKUP: Optional | YES
23 | GAS SUPPLY REQUIREMENTS: Medical air and oxygen at inlet pressures between 4 and 85 psi (28-586 KPa)/Heliox Optional | YES
24 | Warranty: Min 3 Years, Extended Warranty will be prefered | YES
25 | Included Accessories
26 | Accessoriey 1
27 | Accessoriey 2
28 | Accessoriey 3
29 | Accessoriey 4
30 | Accessoriey 5
31 | Accessoriey 6
32 | Accessoriey 7
33 | Accessoriey 8
34 | Optional Accessories
35 | Accessoriey 1 | Qty : 10
36 | Accessoriey 2 | Qty : 10
37 | Accessoriey 3 | Qty : 4
38 | Accessoriey 4 | Qty : 8
39 | Accessoriey 5 | Qty : 6
40 | Accessoriey 6 | Qty : 25
41 | Accessoriey 7 | Qty : 12
42 | Accessoriey 8 | Qty : 17
43 | Other Specification
44 | Vendor will do upgrade during Warranty FOC | YES
45 | Pre-Installation & Commissioning | Included
46 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
47 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
48 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
49 | The vendor to provide: a) Installation. b) Set up. c) Performance verification. d) Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
50 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
51 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
52 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
53 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
54 | All needed Functional Verifications and Calibrations will be included at no additional charges | YES
Company Stamp |  |  |  | Signature

## ITEM 2
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
2 | 442272201 | 0 | CRITICAL CARE VENTILATOR NIV FOR ADULT/PEDIATRIC PATIENT MODEL V680 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Equipment needed for Following Patient Type: Adult to Infant (2.5 Kg – 250 Kg) | YES
2 | CONTROL PARAMETERS AS FOLLOWS:
3 | Inspiration Time, Sec: 0.5 - 5 | YES
4 | Inspiratory Pressure, cmH2O: 1 – 30 | YES
5 | Expiration Time, Sec: 0.5 - 5 | YES
6 | Expiration Pressure, cmH2O: 1 - 30 | YES
7 | Ramp Feature | YES
8 | Trigger Sensitivity: Digital Auto Trak | YES
9 | Breath Rate: 0 – 40 (AVAPS / Auto) mode only | YES
10 | Panel Lock | YES
11 | Rise Profile: 1 - 5 | YES
12 | APNEA Adjustment | YES
13 | EQUIPMENT ALARMS AS FOLLOWS:
14 | Mask Disconnection | YES
15 | Power Failure | YES
16 | Vent Inoperative | YES
17 | Low Battery | YES
18 | Self-diagnostics and reference error code | YES
19 | Overheat | YES
20 | MISCELLANEOUS INFORMATION AS FOLLOWS:
21 | send reports to printer | YES
22 | Save data to external media | YES
23 | Send data via network | YES
24 | MONITORING PARAMETER AS FOLLOWS:
25 | Inspiration Time | YES
26 | Expiration Time | YES
27 | Inspiration Pressure | YES
28 | Expiration Pressure | YES
29 | Breath Rate | YES
30 | I:E Ratio | YES
31 | Mean Airway Pressure MAP | YES
32 | Minute Volume | YES
33 | Tidal Volume | YES
34 | Leak Percentage | YES
35 | Flow | YES
36 | PATIENT ALARMS AS FOLLOWS:
37 | High Respiration Rate | YES
38 | APNEA | YES
39 | High Pressure Alarm | YES
40 | Low Minute Volume (Auto mode only) | YES
41 | High/Low Tidal Volume (Auto mode only) | YES
42 | Rebreathing | YES
43 | PHYSICAL PARAMETERS AND ENVIRONMENT SPECIFICATION AS FOLLOWS:
44 | Acceptable Noise ≤50dB | YES
45 | Altitude : - 300m to 3000m from the sea level | YES
46 | Temperature : 15 to 45 C | YES
47 | Dimension | Pls Specify
48 | ELECTRICAL SPECIFICATION AS FOLLOWS:
49 | BIPAP must have dual voltage operational compatibility at automatic selection of frequency. | YES
50 | Voltage Range : 100 ~ 240 Volts otherwise 240v is acceptable | YES
51 | Frequency : 50Hz / 60Hz | YES
52 | Maximum Current : ≤ 3 Amp | YES
53 | Battery backup time must be 6 hrs. at-least | YES
54 | Battery type should be Lithium Ion Or Non Spill-able (SLA will not accepted) | YES
55 | VENTILATION MODES AS FOLLOWS:
56 | CPAP | YES
57 | Synchronized S Mode | YES
58 | S/T Mode | YES
59 | T Mode | YES
60 | PC Mode | YES
61 | AVAP Mode | YES
62 | PCV Mode | YES
63 | BiLEVEL | YES
64 | The system must compliance with the following safety standards:
65 | - ISO 80601-1-12 | YES
66 | - ISO 80601-1-55 | YES
67 | - IEC 60801-1 Type B – for ventilator base unit or any part in direct contact with patient | YES
68 | - IPX1 | YES
69 | Included Accessories
70 | Accessoriey 1
71 | Accessoriey 2
72 | Accessoriey 3
73 | Accessoriey 4
74 | Accessoriey 5
75 | Accessoriey 6
76 | Accessoriey 7
77 | Accessoriey 8
78 | Optional Accessories
79 | Accessoriey 1
80 | Accessoriey 2
81 | Accessoriey 3
82 | Accessoriey 4
83 | Accessoriey 5
84 | Accessoriey 6
85 | Accessoriey 7
86 | Accessoriey 8
87 | Other Specification
88 | Vendor will do upgrade during Warranty FOC | YES
89 | Pre-Installation & Commissioning | Included
90 | Hospital Engineer can update / upgrade/download software into the ventilator in-house, if required any equipment/hardware/ password/software must be provided to hospital FOC. | YES
91 | The equipment or any of its accessories must not be presently listed in any recall | YES
92 | Vendor Should provide any cable for interfacing to laptop for calibration and maintenance FOC | YES
93 | Ventilator should have service mode and it must be password protected. | YES
94 | Audible Alarm volume can be adjusted : range 35dB to 90dB | YES
95 | Warning message to replace battery | YES
96 | Dual Limb active patient circuit support | YES
97 | The vendor is responsible to provide a minimum of THREE (3) year full warranty commencing from the date of first case. | YES
98 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
99 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
100 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
101 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
102 | The contractor will provide all the essential tools and test equipment required to maintain or calibrate the equipment free of charge. | YES
103 | The vendor should shoulder the periodic maintenance with pm kits, Batteries and other parts cost during warranty period. | YES
104 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
105 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
106 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent). The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
107 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
108 | Vendor to provide approved SFDA Certificate. | YES
109 | Medical equipment to be used in the hospital will comply with hoses color code and gas outlets for medical gases system as currently standardized and used at KFSH Madinah: | YES
110 | Ventilator should be capable to work with gas Pressure: 30 to 70 PSI | YES
111 | Power plug should be British style medical hospital grade. | YES
112 | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies. b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 3
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
3 | 413286112 | 0 | TELESCOPE 0 DEGREE HOPKINS II STRAIGHT FORWARD ENLARGED VIEW DIAMETER 10 MM LENGTH 31 CM AUTOCLAVABLE FOR INDOCYANIN GREEN ICG AND AF FIBER OPTIC LIGHT TRANSMISSION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Technical Parameters | YES
2 | Scope Type: Diagnostic/operating | YES
3 | Optical design system: rod-lens system and the single-image fiber system. | YES
4 | Diameter, mm : 10 MM | YES
5 | working sleeve or sheath Provided | YES
6 | Insertion length with sleeve, mm 80 to 200 | YES
7 | field of view between 70° and 100°. | YES
8 | Directions of view 0, 30, or 70 (30 is the most preferred, and pls refer to item description for details) | YES
9 | IRRIGATION FLOW, mL/min 200 to 700 | YES
10 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
11 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
12 | Video mounts, Preferred C-mount | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
35 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
36 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
37 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be inf | YES
38 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of productio | YES
40 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
42 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
43 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of th | YES
Company Stamp |  |  |  | Signature

## ITEM 4
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
4 | 499900515 | 0 | TELESCOPE HOPKINS II 30 DEGREES 10 MM ENLARGED VIEW 31CM, AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED # 26003BA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Optical design system: rod-lens system and the single-image fiber system. | YES
3 | Diameter, mm : 10 MM | YES
4 | working sleeve or sheath Provided | YES
5 | Insertion length with sleeve, mm 80 to 200 | YES
6 | field of view between 70° and 100°. | YES
7 | Directions of view 0, 30, or 70 (30 is the most preferred, and pls refer to item description for details) | YES
8 | IRRIGATION FLOW, mL/min 200 to 700 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Length cm: 31 | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
35 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
36 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
37 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
38 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
40 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
42 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
43 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 5
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
5 | 413286117 | 0 | TELESCOPE 30 DEGREE 5MM LENGTH 29CM AUTOCLAVABLE FIBER OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE RED 26046BA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Optical design system: rod-lens system and the single-image fiber system. | YES
3 | Diameter, mm : 5 MM | YES
4 | working sleeve or sheath Provided | YES
5 | Insertion length with sleeve, mm 80 to 200 | YES
6 | field of view between 70° and 100°. | YES
7 | Directions of view 0, 30, or 70 (30 is the most preferred, and pls refer to item description for details) | YES
8 | IRRIGATION FLOW, mL/min 200 to 700 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Length cm: 29 | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
35 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
36 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
37 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
38 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
40 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
42 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
43 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 6
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
6 | 411637814 | 0 | TELESCOPE HOPKINS II 0 DEGREE STRAIGHT FORWARD ENLARGED VIEW DIAMETER 3.3MM LENGTH 25CM AUTOCLAVABLE FIBRE OPTIC LIGHT TRANSMISSION INCORPORATED COLOR CODE: GREEN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Optical design system: rod-lens system and the single-image fiber system. | YES
3 | Diameter, mm : 3.3 MM | YES
4 | working sleeve or sheath Provided | YES
5 | Insertion length with sleeve, mm 80 to 200 | YES
6 | field of view between 70° and 100°. | YES
7 | Directions of view 0, 30, or 70 (30 is the most preferred, and pls refer to item description for details) | YES
8 | IRRIGATION FLOW, mL/min 200 to 700 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Length cm: 25 | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
35 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
36 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
37 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be inf | YES
38 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of productio | YES
40 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
42 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
43 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of th | YES
Company Stamp |  |  |  | Signature

## ITEM 7
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
7 | 453298105 | 0 | TELESCOPE HOPKINS II 3.3MM STRAIGHT FORWARD 30 DEGREE LENGTH 25CM COLOR CODE RED AUTOCLAVABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Optical design system: rod-lens system and the single-image fiber system. | YES
3 | Diameter, mm : 3.3 MM | YES
4 | working sleeve or sheath Provided | YES
5 | Insertion length with sleeve, mm 80 to 200 | YES
6 | field of view between 70° and 100°. | YES
7 | Directions of view 0, 30, or 70 (30 is the most preferred, and pls refer to item description for details) | YES
8 | IRRIGATION FLOW, mL/min 200 to 700 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Length cm: 25 | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
35 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
36 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
37 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
38 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
40 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
42 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
43 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 8
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
8 | 411633221 | 0 | VIDEO GASTROSCOPE SLIM 90K WORKING LENGTH 1050MM OUTER DIAMETER 8.0MM FIELD OF VIEW 140DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Video: HD preferred | YES
3 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
4 | WORKING LENGTH: 1050 MM | YES, Pls Specify
5 | Outer diameter,8 mm | Pls Specify
6 | Graduation marks Preferred, cm | YES
7 | WORKING CHANNELS : ≥1 | YES
8 | Diameter,8 mm | YES , Pls Specify
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Field of view, degrees: 140 | YES, Pls Specify
13 | INSTRUMENT CHANNEL 4.2MM | YES
14 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
15 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
16 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
17 | Aspiration | YES
18 | Air/water delivery, water jet function: Prefered | YES
19 | Compatible cameras: Compatibility with multiple cameras preferred | YES
20 | Compatible light sources: Compatibility with multiple light sources required | YES
21 | Compatible video processors: Compatibility with multiple processors preferred | YES
22 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
23 | Included Accessories
24 | Accessoriey 1
25 | Accessoriey 2
26 | Accessoriey 3
27 | Accessoriey 4
28 | Accessoriey 5
29 | Accessoriey 6
30 | Accessoriey 7
31 | Accessoriey 8
32 | Optional Accessories
33 | Accessoriey 1
34 | Accessoriey 2
35 | Accessoriey 3
36 | Accessoriey 4
37 | Accessoriey 5
38 | Accessoriey 6
39 | Accessoriey 7
40 | Accessoriey 8
41 | Other Specification
42 | Vendor will do upgrade during Warranty FOC | YES
43 | Pre-Installation & Commissioning | Included
44 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
45 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
46 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
47 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
49 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
50 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
51 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
52 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
53 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 9
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
9 | 411633222 | 0 | VIDEO COLONOSCOPE SLIM HD+ WORKING LENGTH 1700MM OUTER DIAMETER 11.6MM FIELD OF VIEW 140DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Applications for Laparoscopy, General Surgery, Ob-Gayne, Pediatric Surgery | YES
3 | Camera controls: four or five programmable buttons | YES
4 | Mechanical buttons: suction and air/water | YES
5 | Control knobs: up/down and left/right control knobs for articulation of the tip of the colonoscope | YES
6 | Video: HD preferred | YES
7 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
8 | WORKING LENGTH: 1700MM | YES
9 | Outer diameter, 11.6 mm | Pls Specify
10 | Graduation marks Preferred, cm | YES
11 | WORKING CHANNELS : ≥1 | YES
12 | Diameter, mm :11.6 | YES Pls Specify
13 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
14 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
15 | Video mounts, Preferred C-mount | YES
16 | Field of view, degrees: 140 | YES
17 | INSTRUMENT CHANNEL 4.2MM | YES
18 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
19 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
20 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
21 | Aspiration | YES
22 | Air/water delivery, water jet function: Prefered | YES
23 | Compatible cameras: Compatibility with multiple cameras preferred | YES
24 | Compatible light sources: Compatibility with multiple light sources required | YES
25 | Compatible video processors: Compatibility with multiple processors preferred | YES
26 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
27 | Included Accessories
28 | Accessoriey 1
29 | Accessoriey 2
30 | Accessoriey 3
31 | Accessoriey 4
32 | Accessoriey 5
33 | Accessoriey 6
34 | Accessoriey 7
35 | Accessoriey 8
36 | Optional Accessories
37 | Accessoriey 1
38 | Accessoriey 2
39 | Accessoriey 3
40 | Accessoriey 4
41 | Accessoriey 5
42 | Accessoriey 6
43 | Accessoriey 7
44 | Accessoriey 8
45 | Other Specification
46 | Vendor will do upgrade during Warranty FOC | YES
47 | Pre-Installation & Commissioning | Included
48 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
49 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
50 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
51 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
52 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
53 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
54 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
55 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
56 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
57 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 10
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
10 | 411633223 | 0 | VIDEO GASTROSCOPE DOUBLE CHANNEL WORKING LENGTH 1050MM OUTER DIAMETER 12.8MM FIELD OF VIEW 140DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Video: HD preferred | YES
3 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
4 | WORKING LENGTH: 1050 MM | YES, Pls Specify
5 | Outer diameter,12.8 mm | Pls Specify
6 | Graduation marks Preferred, cm | YES
7 | WORKING CHANNELS : ≥1 | YES
8 | Diameter,12.8 mm | YES , Pls Specify
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Field of view, degrees: 140 | YES, Pls Specify
13 | INSTRUMENT CHANNEL 4.2MM | YES
14 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
15 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
16 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
17 | Aspiration | YES
18 | Air/water delivery, water jet function: Prefered | YES
19 | Compatible cameras: Compatibility with multiple cameras preferred | YES
20 | Compatible light sources: Compatibility with multiple light sources required | YES
21 | Compatible video processors: Compatibility with multiple processors preferred | YES
22 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
23 | Included Accessories
24 | Accessoriey 1
25 | Accessoriey 2
26 | Accessoriey 3
27 | Accessoriey 4
28 | Accessoriey 5
29 | Accessoriey 6
30 | Accessoriey 7
31 | Accessoriey 8
32 | Optional Accessories
33 | Accessoriey 1
34 | Accessoriey 2
35 | Accessoriey 3
36 | Accessoriey 4
37 | Accessoriey 5
38 | Accessoriey 6
39 | Accessoriey 7
40 | Accessoriey 8
41 | Other Specification
42 | Vendor will do upgrade during Warranty FOC | YES
43 | Pre-Installation & Commissioning | Included
44 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
45 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
46 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
47 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
49 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
50 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
51 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
52 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
53 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 11
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
11 | 411633224 | 0 | VIDEO COLONOSCOPE DOUBLE CHANNEL WORKING LENGTH 1700MM OUTER DIAMETER 13.2MM FIELD OF VIEW 140DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Applications for Laparoscopy, General Surgery, Ob-Gayne, Pediatric Surgery | YES
3 | Camera controls: four or five programmable buttons | YES
4 | Mechanical buttons: suction and air/water | YES
5 | Control knobs: up/down and left/right control knobs for articulation of the tip of the colonoscope | YES
6 | Video: HD preferred | YES
7 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
8 | WORKING LENGTH: 1700 mm | YES
9 | Outer diameter, mm | Pls Specify
10 | Graduation marks Preferred, cm | YES
11 | WORKING CHANNELS : ≥1 | YES
12 | Diameter, mm :13.2 | YES Pls Specify
13 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
14 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
15 | Video mounts, Preferred C-mount | YES
16 | Field of view, degrees: 140 | YES
17 | INSTRUMENT CHANNEL 4.2MM | YES
18 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
19 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
20 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
21 | Aspiration | YES
22 | Air/water delivery, water jet function: Prefered | YES
23 | Compatible cameras: Compatibility with multiple cameras preferred | YES
24 | Compatible light sources: Compatibility with multiple light sources required | YES
25 | Compatible video processors: Compatibility with multiple processors preferred | YES
26 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
27 | Included Accessories
28 | Accessoriey 1
29 | Accessoriey 2
30 | Accessoriey 3
31 | Accessoriey 4
32 | Accessoriey 5
33 | Accessoriey 6
34 | Accessoriey 7
35 | Accessoriey 8
36 | Optional Accessories
37 | Accessoriey 1
38 | Accessoriey 2
39 | Accessoriey 3
40 | Accessoriey 4
41 | Accessoriey 5
42 | Accessoriey 6
43 | Accessoriey 7
44 | Accessoriey 8
45 | Other Specification
46 | Vendor will do upgrade during Warranty FOC | YES
47 | Pre-Installation & Commissioning | Included
48 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
49 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
50 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
51 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
52 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
53 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
54 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
55 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
56 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
57 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 12
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
12 | 413614702 | 0 | VIDEO COLONOSCOPE HGH DEFINITION DEC38- I10L SERIES 13.2MM/3.8MM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Applications for Laparoscopy, General Surgery, Ob-Gayne, Pediatric Surgery | YES
3 | Camera controls: four or five programmable buttons | YES
4 | Mechanical buttons: suction and air/water | YES
5 | Control knobs: up/down and left/right control knobs for articulation of the tip of the colonoscope | YES
6 | Video: HD preferred | YES
7 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
8 | WORKING LENGTH: 1100MM | YES
9 | Outer diameter, mm | Pls Specify
10 | Graduation marks Preferred, cm | YES
11 | WORKING CHANNELS : ≥1 | YES
12 | Diameter, mm :13.2 | YES Pls Specify
13 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
14 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
15 | Video mounts, Preferred C-mount | YES
16 | Field of view, degrees: 140 | YES
17 | INSTRUMENT CHANNEL 4.2MM | YES
18 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
19 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
20 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
21 | Aspiration | YES
22 | Air/water delivery, water jet function: Prefered | YES
23 | Compatible cameras: Compatibility with multiple cameras preferred | YES
24 | Compatible light sources: Compatibility with multiple light sources required | YES
25 | Compatible video processors: Compatibility with multiple processors preferred | YES
26 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
27 | Included Accessories
28 | Accessoriey 1
29 | Accessoriey 2
30 | Accessoriey 3
31 | Accessoriey 4
32 | Accessoriey 5
33 | Accessoriey 6
34 | Accessoriey 7
35 | Accessoriey 8
36 | Optional Accessories
37 | Accessoriey 1
38 | Accessoriey 2
39 | Accessoriey 3
40 | Accessoriey 4
41 | Accessoriey 5
42 | Accessoriey 6
43 | Accessoriey 7
44 | Accessoriey 8
45 | Other Specification
46 | Vendor will do upgrade during Warranty FOC | YES
47 | Pre-Installation & Commissioning | Included
48 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
49 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
50 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
51 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
52 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
53 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
54 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
55 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
56 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
57 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 13
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
13 | 413614701 | 0 | GASTROSCOPE HIGH DEFINITION VIDEO DEG29- I10 SERIES 9.8MM /3.2MM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Video: HD preferred | YES
3 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
4 | WORKING LENGTH: MM | YES, Pls Specify
5 | Outer diameter, mm | Pls Specify
6 | Graduation marks Preferred, cm | YES
7 | WORKING CHANNELS : ≥1 | YES
8 | Diameter, mm | YES , Pls Specify
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Field of view, degrees: 100 | YES, Pls Specify
13 | INSTRUMENT CHANNEL 4.2MM | YES
14 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
15 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
16 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
17 | Aspiration | YES
18 | Air/water delivery, water jet function: Prefered | YES
19 | Compatible cameras: Compatibility with multiple cameras preferred | YES
20 | Compatible light sources: Compatibility with multiple light sources required | YES
21 | Compatible video processors: Compatibility with multiple processors preferred | YES
22 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
23 | Included Accessories
24 | Accessoriey 1
25 | Accessoriey 2
26 | Accessoriey 3
27 | Accessoriey 4
28 | Accessoriey 5
29 | Accessoriey 6
30 | Accessoriey 7
31 | Accessoriey 8
32 | Optional Accessories
33 | Accessoriey 1
34 | Accessoriey 2
35 | Accessoriey 3
36 | Accessoriey 4
37 | Accessoriey 5
38 | Accessoriey 6
39 | Accessoriey 7
40 | Accessoriey 8
41 | Other Specification
42 | Vendor will do upgrade during Warranty FOC | YES
43 | Pre-Installation & Commissioning | Included
44 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
45 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
46 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
47 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
49 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
50 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
51 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
52 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
53 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 14
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
14 | 413614700 | 0 | VIDEO PROCESSOR HIGH DEFINITION DEPK-17000 SERIES |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Camera inputs At least 1 | YES
2 | Hue adjustment Automatic | YES
3 | Illumination Control Automatic
4 | White balance Automatic | YES
5 | Trans illumination | YES
6 | Video gain | YES
7 | Sharpness | YES
8 | Zoom | YES
9 | Other controls: Remote control, digital signal processing, Automatic system check(please specify) | YES
10 | RGB Outputs | YES
11 | Y/C Outputs | YES
12 | Composite Outputs | YES
13 | PAL | YES
14 | HDMI | YES
15 | Digital Outputs Yes, for HD systems | YES
16 | Included Accessories
17 | Accessoriey 1
18 | Accessoriey 2
19 | Accessoriey 3
20 | Accessoriey 4
21 | Accessoriey 5
22 | Accessoriey 6
23 | Accessoriey 7
24 | Accessoriey 8
25 | Optional Accessories
26 | Accessoriey 1
27 | Accessoriey 2
28 | Accessoriey 3
29 | Accessoriey 4
30 | Accessoriey 5
31 | Accessoriey 6
32 | Accessoriey 7
33 | Accessoriey 8
34 | Other Specification
35 | Vendor will do upgrade during Warranty FOC | YES
36 | Pre-Installation & Commissioning | Included
37 | Others HD, HD-SDI, remote-control connectors | YES
38 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
39 | Parts and software updates will be done F.O.C. during warrant period. | YES
40 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
41 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
42 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
43 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
44 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
45 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
46 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
47 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
48 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
49 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
50 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
51 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
52 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
53 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
54 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
55 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
56 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
57 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 15
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
15 | 414249702 | 0 | VIDEO BRONCHOSCOPE 5.2MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.0MM FIELD OF VIEW 120DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATIONS Laparoscopy, General Surgery, Ob-Gayne, Pediatric Surgery | YES
2 | All (with appropriate adapters, with specifications & quantity as indicated above) | YES
3 | SYSTEM INTERFACES Y/C, composite, RBG, IEEE 1394, SDI, DVI | YES
4 | Included Accessories
5 | Accessoriey 1
6 | Accessoriey 2
7 | Accessoriey 3
8 | Accessoriey 4
9 | Accessoriey 5
10 | Accessoriey 6
11 | Accessoriey 7
12 | Accessoriey 8
13 | Optional Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Other Specification
23 | Vendor will do upgrade during Warranty FOC | YES
24 | Pre-Installation & Commissioning | Included
25 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. | YES
26 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
27 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
28 | All Equipment provided according to the terms of this RFP shall be of current production, new, and of first-rate quality. Reconditioned or previously used equipment shall be unacceptable. | YES
29 | . All quality assurance checks, inspection, and testing to be performed by BIDDER before delivery shall be performed at the manufacturer's facility. KFSH reserves the right to witness any or all of these without prior notice to BIDDER. BIDDER will provide KFSH with a schedule of any quality assurance checks, inspections, and testing thirty (30) days before the delivery of the proposed system. However, KFSH shall be under no obligation to witness same, and failure of KFSH to witness any quality assurance checks, inspections, and/or testing shall not in any way relieve BIDDER of the obligation to fully comply with all the requirements, codes, and standards pertaining to the proposed medical equipment. | YES
30 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
31 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
32 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
33 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
34 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
35 | Supplier will provide a similar loaner or replacement of the scope during any failures or repairs | YES
Company Stamp |  |  |  | Signature

## ITEM 16
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
16 | 414249703 | 0 | VIDEO BRONCHOSCOPE 6.4MM TUBE WORKING LENGTH 600MM INSTRUMENT CHANNEL 2.8MM FIELD OF VIEW 120DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | APPLICATIONS Laparoscopy, General Surgery, Ob-Gayne, Pediatric Surgery | YES
2 | All (with appropriate adapters, with specifications & quantity as indicated above) | YES
3 | SYSTEM INTERFACES Y/C, composite, RBG, IEEE 1394, SDI, DVI | YES
4 | Included Accessories
5 | Accessoriey 1
6 | Accessoriey 2
7 | Accessoriey 3
8 | Accessoriey 4
9 | Accessoriey 5
10 | Accessoriey 6
11 | Accessoriey 7
12 | Accessoriey 8
13 | Optional Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Other Specification
23 | Vendor will do upgrade during Warranty FOC | YES
24 | Pre-Installation & Commissioning | Included
25 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. | YES
26 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
27 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
28 | All Equipment provided according to the terms of this RFP shall be of current production, new, and of first-rate quality. Reconditioned or previously used equipment shall be unacceptable. | YES
29 | . All quality assurance checks, inspection, and testing to be performed by BIDDER before delivery shall be performed at the manufacturer's facility. KFSH reserves the right to witness any or all of these without prior notice to BIDDER. BIDDER will provide KFSH with a schedule of any quality assurance checks, inspections, and testing thirty (30) days before the delivery of the proposed system. However, KFSH shall be under no obligation to witness same, and failure of KFSH to witness any quality assurance checks, inspections, and/or testing shall not in any way relieve BIDDER of the obligation to fully comply with all the requirements, codes, and standards pertaining to the proposed medical equipment. | YES
30 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
31 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
32 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
33 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
34 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
35 | Supplier will provide a similar loaner or replacement of the scope during any failures or repairs | YES
Company Stamp |  |  |  | Signature

## ITEM 17
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
17 | 411637910 | 0 | INSUFFLATOR CO2 UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Equipment TYPE: Electronic | YES
2 | Gas Type: CO2 | YES
3 | INSUFFLATION PRESSURE RANGE, mm Hg | Pls Specify
4 | FLOW SETTINGS Range, L/min : 0-30 | YES
5 | Flow Fixed settings, L/min : 1-2 minimum settings | YES
6 | OVERPRESSURE PROTECTION | YES
7 | Overpressure Alarm type: Audible and visual | YES
8 | Overpressure Activation press, mm Hg : >30 (user-selectable) | YES
9 | Overpressure Protection Time delay, sec | Pls Specify
10 | Overpressure Electronic venting | YES
11 | Overpressure Electronic venting Activation press, mm Hg : >30 for >15 sec | YES
12 | PERITONEAL PRESSURE GAUGE/DISPLAY | YES
13 | INSUFFLATION VOLUME GAUGE/DISPLAY | YES
14 | EXTERNAL GAS CYLINDER Gauge/display | YES
15 | EXTERNAL GAS CYLINDER Low-pressure alarm : Visual and Audible | YES
16 | HYDROPHOBIC BACTERIAL FILTER | YES
17 | VIDEO MONITOR CAPABILITY | YES
18 | REMOTE CONTROL OPTION | YES
19 | GAS WARMING/ HUMIDIFYING | YES
20 | The total volume of gas introduced into the patient display | YES
21 | Possibility for direct gas connection to central CO2 pipe line | YES
22 | Interfacable to Integrated OR system option | YES
23 | Included Accessories
24 | Accessoriey 1
25 | Accessoriey 2
26 | Accessoriey 3
27 | Accessoriey 4
28 | Accessoriey 5
29 | Accessoriey 6
30 | Accessoriey 7
31 | Accessoriey 8
32 | Optional Accessories
33 | Accessoriey 1
34 | Accessoriey 2
35 | Accessoriey 3
36 | Accessoriey 4
37 | Accessoriey 5
38 | Accessoriey 6
39 | Accessoriey 7
40 | Accessoriey 8
41 | Other Specification
42 | Vendor will do upgrade during Warranty FOC | YES
43 | Pre-Installation & Commissioning | Included
44 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
45 | Parts and software updates will be done F.O.C. during warrant period. | YES
46 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
47 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
48 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
49 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
50 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
51 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
52 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
53 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
54 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
55 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
56 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
57 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
58 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
59 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
60 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
61 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | Included
62 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
63 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 18
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
18 | 411640201 | 0 | ENDOSCOPIC IRRIGATION PUMP DIMENSIONS 121MM X 197MM X 349MM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Pump Type Peristaltic, electronic regulated control | YES
2 | FLUID PRESSURE Range, mm Hg : 30 | YES
3 | FLOW SETTINGS AS FOLLOWS:
4 | Flow rate, L/min: 0-50 L/MIN | YES
5 | Fixed settings, L/min: Pre-selectable up to 1 | YES
6 | OVERPRESSURE PROTECTION AS FOLLOWS:
7 | Alarm type: Audible alarm and pump stops | YES
8 | Alarm activation pressure, mm Hg | Pls Specify
9 | Time delay, sec | Pls Specify
10 | Electronic venting | YES
11 | Activation pressure, mm Hg : 20% or 10 mm Hg above pressure, whichever is greater | YES
12 | JOINT PRESSURE GAUGE/DISPLAY Required/digital display preferred; accuracy with greater of 20% or 10 mm Hg | YES
13 | REMOTE CONTROL: (optional) | YES
14 | FOOTSWITCH: (optional) | YES
15 | Included Accessories
16 | Accessoriey 1
17 | Accessoriey 2
18 | Accessoriey 3
19 | Accessoriey 4
20 | Accessoriey 5
21 | Accessoriey 6
22 | Accessoriey 7
23 | Accessoriey 8
24 | Optional Accessories
25 | Accessoriey 1
26 | Accessoriey 2
27 | Accessoriey 3
28 | Accessoriey 4
29 | Accessoriey 5
30 | Accessoriey 6
31 | Accessoriey 7
32 | Accessoriey 8
33 | Other Specification
34 | Vendor will do upgrade during Warranty FOC | YES
35 | Pre-Installation & Commissioning | Included
36 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
37 | Parts and software updates will be done F.O.C. during warrant period. | YES
38 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
39 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
40 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
41 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
42 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
43 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
44 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
45 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
46 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
47 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
48 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
49 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
50 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
51 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
52 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
53 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | Included
54 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
55 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 19
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
19 | 411637911 | 0 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 27-INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LCD or TFT Medical/Surgical Grade Monitor Full HD | YES
2 | The System shall Have the Highest – Resolution 27 INCHES | YES
3 | Multi modality display capability , including Picture in Picture (PIP) and Picture out Picture ( POP) with various image size | YES
4 | Widescreen aspect ratio | YES
5 | Adjustable image size . | YES
6 | Attachment adjustable bracket for mounting on video cart or wall mount | YES
7 | 110 to 240 v, mains operation, 50-60 Hz | YES
8 | Medical grade British Plug | YES
9 | HD Transmission Cable to Remote display of 10m length(Alternative electronic Transmission is optional) | YES
10 | Minimum 5 Years full warranty | YES
11 | Included Accessories
12 | Accessoriey 1
13 | Accessoriey 2
14 | Accessoriey 3
15 | Accessoriey 4
16 | Accessoriey 5
17 | Accessoriey 6
18 | Accessoriey 7
19 | Accessoriey 8
20 | Optional Accessories
21 | Accessoriey 1
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Other Specification
30 | Vendor will do upgrade during Warranty FOC | YES
31 | Pre-Installation & Commissioning | Included
32 | The vendor is responsible to provide a minimum of5-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
33 | Parts and software updates will be done F.O.C. during warrant period. | YES
34 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
35 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
36 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
37 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
38 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
39 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
40 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
42 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
43 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
44 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
45 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
46 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
47 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
49 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
50 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
51 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 20
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
20 | 411637912 | 0 | MONITOR MEDICAL GRADE LCD 4K DISPLAY SCREEN 32-INCH |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | LCD or TFT Medical/Surgical Grade Monitor Full HD | YES
2 | The System shall Have the Highest – Resolution 32 INCHES | YES
3 | Multi modality display capability , including Picture in Picture (PIP) and Picture out Picture ( POP) with various image size | YES
4 | Widescreen aspect ratio | YES
5 | Adjustable image size . | YES
6 | Attachment adjustable bracket for mounting on video cart or wall mount | YES
7 | 110 to 240 v, mains operation, 50-60 Hz | YES
8 | Medical grade British Plug | YES
9 | HD Transmission Cable to Remote display of 10m length(Alternative electronic Transmission is optional) | YES
10 | Minimum 5 Years full warranty | YES
11 | Included Accessories
12 | Accessoriey 1
13 | Accessoriey 2
14 | Accessoriey 3
15 | Accessoriey 4
16 | Accessoriey 5
17 | Accessoriey 6
18 | Accessoriey 7
19 | Accessoriey 8
20 | Optional Accessories
21 | Accessoriey 1
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Other Specification
30 | Vendor will do upgrade during Warranty FOC | YES
31 | Pre-Installation & Commissioning | Included
32 | The vendor is responsible to provide a minimum of5-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
33 | Parts and software updates will be done F.O.C. during warrant period. | YES
34 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
35 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
36 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
37 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
38 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
39 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
40 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
41 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
42 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
43 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
44 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
45 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
46 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
47 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
49 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
50 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
51 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 21
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
21 | 411633225 | 0 | VIDEO DUODENOSCOPE 11.6MM TUBE WORKING LENGTH 1250MM INSTRUMENT CHANNEL 4.2MM FIELD OF VIEW 100DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Video: HD preferred | YES
3 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
4 | WORKING LENGTH: 1250MM | YES
5 | Outer diameter, mm | Pls Specify
6 | Graduation marks Preferred, cm | YES
7 | WORKING CHANNELS : ≥1 | YES
8 | Diameter, mm :2 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Field of view, degrees: 100 | YES
13 | INSTRUMENT CHANNEL 4.2MM | YES
14 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
15 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
16 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
17 | Aspiration | YES
18 | Air/water delivery, water jet function: Prefered | YES
19 | Compatible cameras: Compatibility with multiple cameras preferred | YES
20 | Compatible light sources: Compatibility with multiple light sources required | YES
21 | Compatible video processors: Compatibility with multiple processors preferred | YES
22 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
23 | Included Accessories
24 | Accessoriey 1
25 | Accessoriey 2
26 | Accessoriey 3
27 | Accessoriey 4
28 | Accessoriey 5
29 | Accessoriey 6
30 | Accessoriey 7
31 | Accessoriey 8
32 | Optional Accessories
33 | Accessoriey 1
34 | Accessoriey 2
35 | Accessoriey 3
36 | Accessoriey 4
37 | Accessoriey 5
38 | Accessoriey 6
39 | Accessoriey 7
40 | Accessoriey 8
41 | Other Specification
42 | Vendor will do upgrade during Warranty FOC | YES
43 | Pre-Installation & Commissioning | Included
44 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
45 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
46 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
47 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
48 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
49 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
50 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
51 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
52 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
53 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 22
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
22 | 411633226 | 0 | VIDEO DUODENOSCOPE HARD TIP 12.5MM SOFT TIP 10.8MM CHANNEL DIAMETER 3.2MM FIELD OF VIEW 100DEG |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Scope Type: Diagnostic/operating | YES
2 | Video: HD preferred | YES
3 | INSERTION TUBE LENGTH, The length of the insertion tube from the distal tip to the proximal housing | Pls Specify
4 | Working length, mm | Pls Specify
5 | Outer diameter, mm | Pls Specify
6 | Graduation marks Preferred, cm | YES
7 | WORKING CHANNELS : ≥1 | YES
8 | Diameter, mm :2 | YES
9 | COMPATIBLE LIGHT SOURCES All common models (please specify) | YES
10 | RECOMMENDED STERILIZATION METHODS All common methods; autoclavable preferred | YES
11 | Video mounts, Preferred C-mount | YES
12 | Field of view, degrees: 100 | YES
13 | Direction of view, degrees (The angle between the axis of the scope and the center of the field of view.) | Pls Specify
14 | TIP DEFLECTION RANGE, degrees : >90 in all directions | YES
15 | CONTROLS: Up/down and right/left angulation dials, and at least 3 programmable switches | YES
16 | Aspiration | YES
17 | Air/water delivery, water jet function: Prefered | YES
18 | Compatible cameras: Compatibility with multiple cameras preferred | YES
19 | Compatible light sources: Compatibility with multiple light sources required | YES
20 | Compatible video processors: Compatibility with multiple processors preferred | YES
21 | RECOMMENDED REPROCESSING: Chemical method; EtO optional; automatic reprocessor compatibility strongly recommended | YES
22 | Included Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Optional Accessories
32 | Accessoriey 1
33 | Accessoriey 2
34 | Accessoriey 3
35 | Accessoriey 4
36 | Accessoriey 5
37 | Accessoriey 6
38 | Accessoriey 7
39 | Accessoriey 8
40 | Other Specification
41 | Vendor will do upgrade during Warranty FOC | YES
42 | Pre-Installation & Commissioning | Included
43 | The vendor is responsible to provide a minimum of 3-years full unconditional warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel.(extended warranty will be preferred and considered) | YES
44 | Compatibility with existing and future equipment should be taken into account. Scopes should permit connection to all common models of camera heads, and light carriers | YES
45 | Copy of the service/cleaning and disinfection/sterilization documentation, with all cleaning tools (brushes, baste, others) has to be provided at no extra cost | YES
46 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
47 | The vendor to provide regular performance verification and preventive maintenance, including endoscope image quality tests (during warranty period). | YES
48 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
49 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
50 | The user in service training must be provided by qualified clinical instructors who are not sales personnel | YES
51 | List of reference sites both locally and internationally should be submitted with the details of work performed. | YES
52 | Any Repairs should not last more than one week, otherwise a loaner scope of the similar model, or replacement should be provided as a loaner, or replacement of the defective one, during warranty period, and exchange price policy of no more than 30% of the original purchase cost be offered after warranty. | YES
Company Stamp |  |  |  | Signature

## ITEM 23
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
23 | 411638813 | 0 | ELECTROSURGICAL AND ARGON COAGULATION UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | GENERATOR TYPE: Solid-state | YES
2 | FREQUENCY, kHz: Typical: 300-1,000 | YES
3 | Mono-polar Output: Isolated | YES
4 | Hand switch Output | YES
5 | Footswitch Output | YES
6 | Bipolar Output | YES
7 | Mono-polar Modes: Cut (pure, blend), desiccate, fulgurate | YES
8 | Mono-polar Cut Maximum watts at (rated Ω): 300 (300) | YES
9 | Mono-polar Cut Maximum voltage, Vp-p: 2000 | YES
10 | Mono-polar Coagulate Maximum watts at (rated Ω): 120 (500) | YES
11 | Mono-polar Coagulate Maximum voltage, Vp-p: 6000 | YES
12 | Bipolar Modes: Precise, standard, macro | YES
13 | Bipolar Cut Maximum watts at (rated Ω): 70 (100) | YES
14 | Bipolar Cut Maximum voltage, Vp-p: 750 | YES
15 | Coagulate Maximum watts at (rated Ω): 70 (100) | YES
16 | Coagulate Maximum voltage, Vp-p: 320 | YES
17 | PROTECTIVE CIRCUITS: Return electrode contact-quality monitor (RECQM) | YES
18 | SELF-TEST MODE | YES
19 | INDEPENDENT OUTPUT | YES
20 | Visual ACTIVATION INDICATORS | YES
21 | Audible ACTIVATION INDICATORS | YES
22 | Volume control: >45 dBA at 1 meter | YES
23 | Can Not Defeat able-inaudible Volume | YES
24 | POWER SETTING DISPLAYED | YES
25 | COOLING: Convection, or fan (Kindly Specify) | YES
26 | automatic tissue impedance Power Adjustment | YES
27 | Meets requirements of AAMI, CUL, HF-18, IEC 601 and 901, and UL. | YES
28 | ARGON GAS SYSTEM AS FOLLOWS:
29 | Configuration : Modular Type is Preferred | YES
30 | Flow range, L/min: 0-10 | YES
31 | Tank capacity, L (cu ft): >1,200 (42) | YES
32 | Low-pressure alarm | YES
33 | automatic flow adjustment | YES
34 | number of Tanks: one, two is preferred with one active and other standby | YES
35 | Included Accessories
36 | Accessoriey 1
37 | Accessoriey 2
38 | Accessoriey 3
39 | Accessoriey 4
40 | Accessoriey 5
41 | Accessoriey 6
42 | Accessoriey 7
43 | Accessoriey 8
44 | Optional Accessories
45 | Accessoriey 1
46 | Accessoriey 2
47 | Accessoriey 3
48 | Accessoriey 4
49 | Accessoriey 5
50 | Accessoriey 6
51 | Accessoriey 7
52 | Accessoriey 8
53 | Other Specification
54 | Vendor will do upgrade during Warranty FOC | YES
55 | Pre-Installation & Commissioning | Included
56 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of acceptance. Warranty should include all parts, labor, and travel. An extended warranty of at least (2 ) year(s) after the standard year should also be quoted by the vendor as line item for further consideration. The vendor is encouraged to provide such privilege to the hospital without extra cost to improve his quotation position when evaluated by the concerned technical department. | YES
57 | Parts and software updates will be done F.O.C. during warrant period. | YES
58 | The unit should be provided with a line (power) cord of acceptable durability, quality, length, and ampacity and should be secured with adequate strain reliefs. | YES
59 | Copy of the service documentation has to be approved in advance by technical department before issuing PO. Service manual is mandatory for equipment acceptance | YES
60 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
61 | The vendor to provide installation, set up, calibration, performance verification and preventive maintenance (during warranty period). | YES
62 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts and upgrades for a period of five (5) years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
63 | Demo and or loaner equipment might be checked and managed while at hospital before final purchase. | YES
64 | The bidder shall identify all consumable items required to operate the system and provide guaranteed prices for these items for a period of [five (5)] years after acceptance of the systems by hospital. | YES
65 | The user in service training must be provided by qualified clinical instructors who are not sales personnel. | YES
66 | Supplier shall supply software and upgrades to the equipment at no additional charges and upgrades shall include any circuit boards or other parts required. | YES
67 | All specification values should be supported by official original manufacture documentations, certifications or Technical datasheets , to prevent the supplier indicated values from being ignored or not been considered during tender evaluation phase. | YES
68 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
69 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
70 | a) The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
71 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
72 | The vendor to provide: X Set up X Calibration X  Performance verification and Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
73 | One (1) complete service manual (in English language) , including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | Included
74 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
75 | The device manufacturer to confirm connectivity with CERNER – Millennium care-aware platform for data integrity with HIS system, if applicable | YES
Company Stamp |  |  |  | Signature

## ITEM 24
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
24 | 411555406 | 0 | GAMMA DETECTION SYSTEM CONTROL UNIT WITH SOFTWARE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Gamma Detection System GDS MACHINE | YES
2 | To provide individual counts for each isotope as well as a filter indicating whether the user is closer to the injection site or a radio seed, all in real time. | YES
3 | No calibration or preventative maintenance required. | YES
4 | One touch remote count functionality within the sterile field | YES
5 | Individual counts for radioactive tracer and radioactive seed simultaneously. | YES
6 | On screen target count display for convenient reference. | YES
7 | SHOULD HAVE Bluetooth probes deliver highly accurate localization of target tissue. | YES
8 | Outstanding directionality with or without collimation. | YES
9 | SHOULD BE COMPATIBLE WITH DIFFERENT SIZES OF BLUETOUTH PROBES | YES
10 | Custom built for the GDS system and probes | YES
11 | convenient storage and protection for your probes | YES
12 | Medical grade, smooth and durable wheels with individually locking castors. | YES
13 | Included Accessories
14 | PORTABLE Cart | Included
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The supplier is responsible to provide full warranty after Technical acceptance for minimum 3 years (extended warranty would be prefered), and shall provide technical support & the supply of spare parts after end of production announced. | YES
35 | The supplier to provide and include all the necessary labour, parts, PM kits, during the warranty for: x Installation x Set up x Calibration x Performance verification x Preventive maintenance | YES
36 | Equipment downtime during warranty repair shall not exceed annum. 5% calculated based on 44 hours per week X 52 weeks per included response time from official reported time. | YES
37 | If the above If exceeded, Vendor will supply a loaner unit, or replacement, If not applicable the warranty will be extended by 6 months | YES
38 | Response time not to exceed from official reported time 12 Hours | YES
39 | Confirmation letter from the supplier that equipment delivered is/are the latest model with the latest updated software | YES
40 | The manufacturer/vendor/supplier will submit confirmation letter that the equipment is fully approved with relevant documents from SFDA for local supplier or FDA and CE for international msupplier | YES
41 | AC single phase Medical equipment operation will be use in the main hospital KFSH-M and should be 240Vac British style 3-Prongsmedical grade power cable and plug. | YES
42 | Power plug: should be British style 3-Prongs medical grade power cable and plug. | YES
43 | One complete service manual including schematics and spare parts list in hard or soft copies. | YES
44 | If applicable, original software license to be provided, and copy of the operating software and software drivers, passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
45 | Specific test equipment to be included free of charge | YES
Company Stamp |  |  |  | Signature

## ITEM 25
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
25 | 411555405 | 0 | PROBE ELECTRONIC BLUETOOTH STRAIGHT 14MM GAMMA DETECTION SYSTEM CONTROL UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# |  |  | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Gamma Detection PROBE FOR GDS MACHINE, Bluetooth Wireless Probes | YES
2 | To provide individual counts for each isotope as well as a filter indicating whether the user is closer to the injection site or a radio seed, all in real time. | YES
3 | No calibration or preventative maintenance required. | YES
4 | One touch remote count functionality within the sterile field | YES
5 | Individual counts for radioactive tracer and radioactive seed simultaneously. | YES
6 | On screen target count display for convenient reference. | YES
7 | Dual isotope mode for simultaneous scanning of 99mTc, 125I radioisotopes | YES
8 | Included Accessories
9 | 14mm Bluetooth Wireless Probes
10 | Accessoriey 2
11 | Accessoriey 3
12 | Accessoriey 4
13 | Accessoriey 5
14 | Accessoriey 6
15 | Accessoriey 7
16 | Accessoriey 8
17 | Optional Accessories
18 | Accessoriey 1
19 | Accessoriey 2
20 | Accessoriey 3
21 | Accessoriey 4
22 | Accessoriey 5
23 | Accessoriey 6
24 | Accessoriey 7
25 | Accessoriey 8
26 | Other Specification
27 | Vendor will do upgrade during Warranty FOC | YES
28 | Pre-Installation & Commissioning | Included
29 | The supplier is responsible to provide full warranty after Technical acceptance for minimum 3 years (extended warranty would be prefered), and shall provide technical support & the supply of spare parts after end of production announced. | YES
30 | The supplier to provide and include all the necessary labour, parts, PM kits, during the warranty for: x Installation x Set up x Calibration x Performance verification x Preventive maintenance | YES
31 | Equipment downtime during warranty repair shall not exceed annum. 5% calculated based on 44 hours per week X 52 weeks per included response time from official reported time. | YES
32 | If the above If exceeded, Vendor will supply a loaner unit, or replacement, If not applicable the warranty will be extended by 6 months | YES
33 | Response time not to exceed from official reported time 12 Hours | YES
34 | Confirmation letter from the supplier that equipment delivered is/are the latest model with the latest updated software | YES
35 | The manufacturer/vendor/supplier will submit confirmation letter that the equipment is fully approved with relevant documents from SFDA for local supplier or FDA and CE for international msupplier | YES
36 | AC single phase Medical equipment operation will be use in the main hospital KFSH-M and should be 240Vac British style 3-Prongsmedical grade power cable and plug. | YES
37 | Power plug: should be British style 3-Prongs medical grade power cable and plug. | YES
38 | One complete service manual including schematics and spare parts list in hard or soft copies. | YES
39 | If applicable, original software license to be provided, and copy of the operating software and software drivers, passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
40 | Specific test equipment to be included free of charge | YES
Company Stamp |  |  |  | Signature

## ITEM 26
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
26 | 442295207 | 0 | DERMATOME ELECTRIC SLIMLINE MODEL S COMPLETE KIT UNIVERSAL POWER SUPPLY 110-230 VOLTS, # 3539700 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Dermatome for split thickness skin graft complete kit | YES
2 | FDA CLEARANCE | YES
3 | TYPE: Handpiece, power system, 220v, 50/60Hz British Hospital Grade Plug | YES
4 | Speed Control: Hand or foot | YES
5 | Drill operating speed range, rpm (The speed of the handpiece, given in rotations per minute (rpm). A broad speed range is desirable because it permits adjustability for precision during procedures and variability in the types of applications.): Variable | YES
6 | MATERIALS: Sterilizable | YES
7 | STERILIZATION METHOD: A tool that can be sterilized using various methods is preferred | YES
8 | LUBRICATION REQUIRED: Oil Free is prefered | YES
9 | ROUTINE MAINTENANCE (The protocol required to prevent degradation and maintain the handpiece in favorable working condition. A handpiece that requires infrequent refurbishing and little or no lubrication during use is preferred.) Tools that require less maintenance are preferred | YES
10 | SAW CONFIGURATION: Electric Dermatome, for split thickness skin graft complete kit | YES
11 | MOTOR TYPE: Electric | YES
12 | Included Accessories
13 | power supply, power supply wall cord, 3, 4, 5, 6 inch (5, 8, 10, 13, 15 cm) wide grafts (0.05 to 0.76mm)
14 | Accessoriey 2
15 | Accessoriey 3
16 | Accessoriey 4
17 | Accessoriey 5
18 | Accessoriey 6
19 | Accessoriey 7
20 | Accessoriey 8
21 | Optional Accessories
22 | Accessoriey 1
23 | Accessoriey 2
24 | Accessoriey 3
25 | Accessoriey 4
26 | Accessoriey 5
27 | Accessoriey 6
28 | Accessoriey 7
29 | Accessoriey 8
30 | Other Specification
31 | Vendor will do upgrade during Warranty FOC | YES
32 | Pre-Installation & Commissioning | Included
33 | the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and the equipment meets the standard of electromagnetic compatibility | YES
34 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, SFDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall. | YES
35 | All equipment (and powered accessories) delivered to the must use 220VAC/60 Hz power and power cable with British hospital medical grade outlet 3-prong plug. | YES
36 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license | YES
37 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
38 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models | YES
39 | Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released | YES
40 | All prices must include delivery, installation, warranty and any related technical, specialized test equipment and calibrations, preventive maintenance kits, battery replacements and both technical certified service training to the clinical engineering and end users application training by qualified certified trainer to all related end users covering their shifts and would be repeated regularly during warranty period as refresher training sessions, as requested and agreed between both parties.. | YES
41 | Vendors shall guarantee system upgradeability of all equipment supplied for any obligatory Upgrade\update during and after the warranty period at no additional | YES
42 | Vendors must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operates the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities. | YES
43 | The minimum warranty period provided by the vendor that is acceptable to this hospital shall be three (3) years, and higher preferences will be given to any other offered longer warranty periods (unlimited warranty should be applied to all purchased equipment/systems, with an agreed consumables purchase of periods covered by the same agreements) | YES
44 | The warranty period is to commence after the official hospital preliminary acceptance document has been completed, and successful clinical operations of the supplied equipment. | YES
45 | Two (2) complete operational manuals for hardware and software. | YES
46 | Two (2) service technical manuals | YES
47 | Two (2) spare part lists (if applicable). | YES
48 | Completed schematic diagrams with documentation of onsite modifications (if required). | YES
49 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
50 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
51 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
52 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
53 | Service response time for routine calls should be one (1) hour. | YES
54 | The vendor should provide the manufacturers country of origin for each offered system\equipment | YES
55 | Dermatome Should accept a variety of tools for ENT, and dermatomes | YES
Company Stamp |  |  |  | Signature

## ITEM 27
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
27 | 414256513 | 0 | OPHTHALMIC FUNDUS IMAGING PHOTO AND ANGIOGRAPHY |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | True Color (with Red, Green and Blue channel split) | YES
2 | Fluorescein Angiography | YES
3 | Autofluorescence-Green | YES
4 | Autofluorescence-Blue | YES
5 | Infrared reflectance | YES
6 | External eye image (ocular surface) | YES
7 | Stereo | YES
8 | Field of View (measured from the center of the eye):Widefield (one image) 133˚ | YES
9 | Ultra-widefield (two images) : 200˚ | YES
10 | Montage (up to six images) : up to 267˚ | YES
11 | Resolution:Optical 7.3 μm | YES
12 | Minimum Pupil Diameter : 2.5 mm | YES
13 | Working Distance:25 mm (patient’s eye to front lens) | YES
14 | Compensation for ametropia:-24 D to + 20 D continuous | YES
15 | Light Sources as follows:
16 | Red LED: 585 - 640 nm | YES
17 | Green LED: 500 - 585 nm | YES
18 | Blue LED:435 - 500 nm | YES
19 | Infrared laser diode : 785 nm | YES
20 | Automatic Operations as follows:
21 | Auto-focus | YES
22 | Auto-gain | YES
23 | Auto Montage | YES
24 | Auto-laterality | YES
25 | Aquisition Speed: Live IR Preview 10 frames/second | YES
26 | Image Capture: ≤ 0.2 seconds | YES
27 | Instrument Specifications as follows:
28 | Instrument Weight: ≤ 23 kg | YES
29 | Instrument Dimensions (W x D x H) | Pls Specify
30 | Instrument Table as follows:
31 | Wheelchair accessible | YES
32 | electronic lift Table | YES
33 | Dimensions recommended:37” (94 cm) x 27.5” (70 cm) | YES
34 | Weight:≤ 40 kg | YES
35 | Instrument Voltage and Mains Frequency: 100-240VAC, 50/60 Hz | YES
36 | electric cable Standard: British UK (medical grade) | YES
37 | Monitor size: 22” Full HD | YES
38 | Monitor type:LCD with LED technology with Capacitive Multi-Touch Screen | YES
39 | RAM: min 32GB | YES
40 | Processor: Intel® 6th Generation Core i5-6500TE or better | YES
41 | Input/Output: USB 3.0 x 3 | YES
42 | Hard Drive: ≤ 2 TB (minimum 200,000 images) | YES
43 | Operating System: Windows 10 or later | SYNCHRONIZED
44 | Included Accessories
45 | Accessoriey 1
46 | Accessoriey 2
47 | Accessoriey 3
48 | Accessoriey 4
49 | Accessoriey 5
50 | Accessoriey 6
51 | Accessoriey 7
52 | Accessoriey 8
53 | Optional Accessories
54 | Accessoriey 1
55 | Accessoriey 2
56 | Accessoriey 3
57 | Accessoriey 4
58 | Accessoriey 5
59 | Accessoriey 6
60 | Accessoriey 7
61 | Accessoriey 8
62 | Other Specification
63 | Vendor will do upgrade during Warranty FOC | YES
64 | Pre-Installation & Commissioning | Included
65 | The vendor is responsible to provide a minimum of Three (3) year full warranty commencing from the date of first case. | YES
66 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
67 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
68 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
69 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
70 | The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period. Software and upgrades | YES
71 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
72 | Vendor to provide connectivity with MUSE server and integration | YES
73 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
74 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
75 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent). | YES
76 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
77 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | Included
78 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. The Frequency must be 60Hz. Power plug should be British style medical grade. | YES
79 | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies. b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 28
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
28 | 442183003 | 0 | OPTICAL BIOMETRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Detect unusual eye geometries | YES
2 | Detect poor fixation | YES
3 | Visually measurement | YES
4 | telecentric keratometry | YES
5 | Measurement range As Follows, OR better, please specify:
6 | Axial length 15 – 35 mm, White-to-white 0.1 mm, | Yes Please Specify
7 | Corneal radii 6 – 10 mm | Yes Please Specify
8 | Anterior chamber depth 0.9 – 7.5 mm | Yes Please Specify
9 | Lens thickness | Yes Please Specify
10 | Central corneal thickness 0.2 – 1.2 mm | Yes Please Specify
11 | White-to-white 8 – 16 mm | Yes Please Specify
12 | Display scaling As Follows OR better, please specify:
13 | Axial length 0.01 mm | Yes please specify
14 | Corneal radii 0.01 mm | Yes please specify
15 | Anterior chamber depth 0.01 mm | Yes please specify
16 | Lens thickness 0.01 mm | Yes please specify
17 | Central corneal thickness 1 µm | Yes please specify
18 | White-to-white 0.1 mm | Yes please specify
19 | SD of repeatability As Follows OR better, please specify:
20 | Axial length 8 µm | Yes please specify
21 | Corneal radii 0.09 D | Yes please specify
22 | Cylinder > 2.0 D axis 1.7° | Yes please specify
23 | Anterior chamber depth 11 µm | Yes please specify
24 | Lens thickness 12 µm | Yes please specify
25 | Central corneal thickness 2 µm | Yes please specify
26 | White-to-white 77 µm | Yes please specify
27 | IOL calculation formulas As Follows Or Better, Pls Specify:
28 | SRK/ T, Holladay 2, Hoffer Q, Haigis Suite (includes Haigis, Haigis-L calculation for eyes following myopic / hyperopic LASIK / PRK / LASEK surgery and Haigis-T for toric IOL power calculation) | Yes Please Specify
29 | Interfaces As Follows, OR better, Pls Specify
30 | Data management system | Yes Please Specify
31 | Computer-assisted cataract surgery system (via USB & FORUM) | Yes Please Specify
32 | Data interface for electronic medical record / patient management systems | Yes Please Specify
33 | Data export to USB storage media | Yes Please Specify
34 | Ethernet port for network connection and network printer | Yes Please Specify
35 | Printer Color | Yes
36 | Mobile trolley | YES
37 | Included Accessories
38 | Accessoriey 1
39 | Accessoriey 2
40 | Accessoriey 3
41 | Accessoriey 4
42 | Accessoriey 5
43 | Accessoriey 6
44 | Accessoriey 7
45 | Accessoriey 8
46 | Optional Accessories
47 | Accessoriey 1
48 | Accessoriey 2
49 | Accessoriey 3
50 | Accessoriey 4
51 | Accessoriey 5
52 | Accessoriey 6
53 | Accessoriey 7
54 | Accessoriey 8
55 | Other Specification
56 | Vendor will do upgrade during Warranty FOC | YES
57 | Pre-Installation & Commissioning | Included
58 | The vendor is responsible to provide a minimum of 3-years full warranty commencing from the date of Technical acceptance. | YES
59 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
60 | The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment . | YES
61 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
62 | Voltage: 220V, 60 Hz, UK Hospital Grade | YES
63 | The vendor to provide installation, set up, calibration, performance verification, preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
64 | Site preparations, and pre-installtion works (civil, electrical, mechanical, networking, etc), are the responsibility of the HOSPITAL, unless, otherwise agreed upon/quoted by the SUPPLIER. | YES
65 | One (1) complete operations manual, to be provided in both hard and soft copies if available. | YES
66 | One (1) complete service manual (in English language) including schematics and spare parts list (if applicable). The manual is to be provided in both hard and soft copies if available. | YES
67 | Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge. | YES
68 | Full local technical service training to be provided, by a certified trainer, at the time of installation (or at a mutually agreed upon time free of charge) covering the following areas as applicable: a) Theory of Operation including circuit description and schematics b) Hardware installation, assembly, and disassembly c) Software installation and upgrade procedures d) Troubleshooting procedures e) Preventive Maintenance and Calibration procedures f) Overview of consumables, supplies, and accessories | YES
69 | If the above local User training , service training, manuals, and/or software are not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled. | YES
70 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
71 | Equipment Downtime: maximum allowable time down is 3% of the total operational time, per annum (calculated based on 44 hrs/week X 52 weeks per annum). In case downtime exceeds this percentage, then the warranty will be extended automatically to twice of the actual down time. | YES
72 | The Equipment’s should have SFDA certificate | YES
73 | To vendor,:the software should be compatible with our hospital net work, and to cooperate with HITA dept. to ensure the performance of the system with the net work | YES
Company Stamp |  |  |  | Signature

## ITEM 29
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
29 | 414256512 | 0 | SCANNER OPHTHALMIC VISUAL FIELD ANALAYZER 3 |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Intended Use: The Humphrey Field Analyzer is an automatic perimeter which is intended to be used to measure the visual field of the eye | YES
2 | Power (dedicated power outlet.):90 to 264Vac, 47Hz to 63Hz | YES
3 | License: two Windows® operating system licenses: an embedded version and a consumer version. | YES
4 | Complies with 93/42/EEC Medical Device Directive. The product is RoHS-compliant according to Directive 2011/65/EU | YES
5 | Class I Equipment – Protection against electrical shock | YES
6 | User replicable parts: projection lamp, Trial Lens holder and Liquid Trial Lens , patient response button, external keyboard, and fan filter | YES
7 | Software: Provided with free of charge with 10 user license | YES
8 | Free Software Upgrades for at least 5 years after warranty | YES
9 | Ports: Display, USB, Ethernet, two seven segment for operation code, 12 Volt Output | YES
10 | Front Panel Include Adjustable chin rest inside the testing bowl, visor handle, Forehead rest. | YES
11 | Manual Kinetic Custom Kinetic Patterns Custom Static Patterns | YES
12 | DICOM OPV (Ophthalmic Visual Field) | YES
13 | IOD (Information Object Definition) | YES
14 | OPV IOD Advanced Indices | YES
15 | Remote Diagnostics and Software Loading | YES
16 | SITA , STATPAC | YES
17 | SITA-SWAP | YES
18 | Stimulus Size (I-V) | YES
19 | Auto Pupil Measurement | YES
20 | Stimulus Color | YES
21 | Foveal Threshold | YES
22 | Gaze Tracking | YES
23 | Head Tracking | YES
24 | Vertex Monitor | YES
25 | GPA | YES
26 | Liquid Trial Lens | YES
27 | RelEY Monitor | YES
28 | HD/USB format support:NTFS | YES
29 | Power Protection:UPS with automatic shutdown | YES
30 | Printer Recommended: PCL 6 compliant color LaserJet 600 dpi | YES
31 | Main Screen: Touch LCD LED | YES
32 | Test profiles: Pre-Programed test profiles | YES
33 | Lense Adjustment: Auto | YES
34 | Screen Volume, Brightness: Adjustable | YES
35 | Password protection | YES
36 | Integration with hospital HIS | YES
37 | Self diagnostic test < 5 min | YES
38 | Emergency login | YES
39 | DICOM and FTP Compatible | YES
40 | Local DB: Automatically managed by the device | YES
41 | Reports Automatically generated on command | YES
42 | Debug and error recording: Log file automatically recorded | YES
43 | import patient tests from a shared network folder or USB storage device | YES
44 | Duplicate patient record: Flag notification, Marge capability | YES
45 | simulation mode On/Off | YES
46 | Default Eye Sequence: Programmable | YES
47 | Head Tracking | YES
48 | Vertex Monitor | YES
49 | Auto Pupil Measurement | YES
50 | DHCP | YES
51 | Network Drive Configuration: ability to store and retrieve data from a mapped drive | YES
52 | Auto logoff time | YES
53 | Kinetic Test | YES
54 | Blind Spot Map Mode: Manual Selection | YES
55 | Adjustable table | YES
56 | GPA Summary Report | YES
57 | Stimulus: Maximum Intensity: 10,000 ASB •Duration: 200 msec, 500 msec for Esterman tests, continuous for kinetic tests •Wave Length: Broadband visible light | YES
58 | Visual Field Testing Distance: 30 cm | YES
59 | Bowl Illumination: 31.5 ASB | YES
60 | Maximum Temporal Range: 90 degree | YES
61 | Dynamic Range 50 dB | YES
62 | Accessoriey 1
63 | Accessoriey 2
64 | Accessoriey 3
65 | Accessoriey 4
66 | Accessoriey 5
67 | Accessoriey 6
68 | Accessoriey 7
69 | Accessoriey 8
70 | Other Specification
71 | Vendor will do upgrade during Warranty FOC | YES
72 | Pre-Installation & Commissioning | YES
73 | The vendor is responsible to provide a minimum of Three (_3_) year full warranty commencing from the date of Technical acceptance | YES
74 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
75 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
76 | Vendor or supplier will guarantee the availability of a loaner unit in Riyadh whenever requested by Hospital during the repair of any defective equipment. | YES
77 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment. | YES
78 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
79 | Voltage: a) ☒ Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted). b) ☒ Frequency: 60Hz. c) ☒ Power plug should be British style medical grade. | YES
80 | If applicable, the device manufacturer to confirm the medical device connectivity with HIS / DICOM system. If software/hardware components are required for such connectivity, it should be provided by vendor at no cost to the hospital, otherwise to be quoted accordingly. | YES
81 | The vendor to provide: a) ☒ Installation. b) ☐ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☐ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. h) ☒ UPS: if applicable | YES
82 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
83 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
84 | Full local User training to be provided, by a qualified application trainer, at the time of installation (or at a mutually agreed upon time) to be provided by the vendor free of charge. | YES
85 | Full local technical service training to be provided, by a certified trainer, at the time of installation (or at a mutually agreed upon time free of charge) covering, but not limited to, the following areas as applicable: i) ☒ Theory of Operation including circuit description and schematics ii) ☐ Hardware installation, assembly, and disassembly iii) ☐ Software installation and upgrade procedures iv) ☒ Troubleshooting procedures v) ☒ Preventive Maintenance and Calibration procedures vi) ☐ Overview of consumables, supplies, and accessories | YES
86 | If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. | YES
87 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
88 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 30
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
30 | 411640503 | 0 | MICROSCOPE OPHTHALMOLOGY SURGICAL ON FLOORSTAND WITHOUT ILLUMINATION UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | MICROSCOPE OPHTHALMOLOGY SURGICAL ON FLOOR STAND | YES
2 | Personalized settings to support precise surgical actions and uninterrupted workflow | YES
3 | Individual procedures for each user and surgery type with Combination Mode | YES
4 | Low light, high contrast with adjustable field of illumination via footswitch | YES
5 | Same view for surgeon, assistant, and camera | YES
6 | Built-in 3CMOS HD camera – 4K-ready | YES
7 | Easy integration of recording and documentation systems | YES
8 | Ready for innovative digital imaging solutions such as IOL guidance and Optical Coherence Tomography (OCT) | YES
9 | Magnification: 6:1 zoom, motorized | YES
10 | Total magnification 4.1× to 24.5× with 10× eyepiece 5.1× to 30.7× with 12.5× eyepiece | YES
11 | Focus range 75 mm | YES
12 | Objective / working distance WD 175 mm/f = 200 mm WD 200 mm/f = 225 mm WD 225 mm/f = 250 mm WD: Working distance, f: Focal length | YES
13 | Field of view 51.4–8.6 mm Ø with 10× eyepiece | YES
14 | Eyepieces Wide-field eyepieces for persons wearing glasses 8.3×, 10× and 12.5× dioptric adjustment, ± 5 diopter settings, adjustable eyecup | YES
15 | Direct illumination with 2 LED lamps Main light Integrated LED illumination system for intensive uniform illumination of the field of view Continuously adjustable brightness with halogen-like color temperature CoAx 4 coaxial illumination Illumination unit for generating a clear and stable Red Reflex, decreasing stray light through the sclera and increasing the image contrast Integrated keratascope and slit illumination Conversion filters allow surgeon to select preferred color temperature of main illumination | YES
16 | Adjustable CoAx 4 Diameter of coaxial illumination is adjustable between 4 and 23 mm via footswitch | YES
17 | Fine focus Available for assistant and integrated camera or external 1/3 camera with C-mount interface | YES
18 | OpenArchitecture Prepared for integration of video camera systems, digital recording and imaging systems such as IOLcompass, EnFocus OCT, and monitors | YES
19 | Connectors Numerous built-in connectors for video and control data transfer Internal power supply 12 VDC, 19 VDC, 24 VDC and AC terminals | YES
20 | 2D/3D HD Video Optional fully integrated 2D HD and/or 3D HD video and recording | YES
21 | Optics > 360° rotation > 15° /+ 105° motorized inclination tilt | YES
22 | XY speed Zoom linked XY speed | YES
23 | XY range 62 × 62 mm | YES
24 | Balancing Adjustable gas spring via balancing knob | YES
25 | Brakes Floor stand with 4 electromagnetic brakes | YES
26 | Monitor arm 860 mm flexible arm with 4 axis for rotation and inclination, max. weight 15 kg and up to 32” | YES
27 | Control unit > User-friendly, individually programmable touch-screen (up to 30 surgeons) for control of motor functions and light intensity > Menu selection based on unique software for user-specific configuration > Built-in electronic auto-diagnosis and user support > Software independent hard keys and indicator for illumination > Data shown by means of LCD | YES
28 | Control elements > Rotary handles > 14-function and 12-function wireless footswitch with optional back-up cable | YES
29 | IR sensor Remote control of the HDR recorder | YES
30 | Indicators > LED for video record status > Surgeon information panel for setting status | YES
31 | Included Accessories
32 | Accessoriey 1
33 | Accessoriey 2
34 | Accessoriey 3
35 | Accessoriey 4
36 | Accessoriey 5
37 | Accessoriey 6
38 | Accessoriey 7
39 | Accessoriey 8
40 | Optional Accessories
41 | Accessoriey 1
42 | Accessoriey 2
43 | Accessoriey 3
44 | Accessoriey 4
45 | Accessoriey 5
46 | Accessoriey 6
47 | Accessoriey 7
48 | Accessoriey 8
49 | Other Specification
50 | Vendor will do upgrade during Warranty FOC | YES
51 | Pre-Installation & Commissioning | Included
52 | The vendor is responsible to provide a minimum of Three (3) year full warranty commencing from the date of first case. | YES
53 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
54 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
55 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
56 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
57 | The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period. Software and upgrades | YES
58 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
59 | Vendor to provide connectivity with MUSE server and integration | YES
60 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
61 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
62 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent). | YES
63 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
64 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | Included
65 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. The Frequency must be 60Hz. Power plug should be British style medical grade. | YES
66 | a) One (1) complete operations manual, to be provided preferably in both hard and soft copies. b) One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 31
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
31 | 414253300 | 0 | AUTOREFRACTOR KERATOMETER UNIT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Autotracking / Auto shot X,7,Z direction | YES
2 | Measurement recording Built In Printer (up to 10 measurements of each eye memory) | YES
3 | Vision Comparison Available with VA Chart | YES
4 | Energy Saving Automatically after 10 minutes if not in use | YES
5 | Auto / Manual Mode | YES
6 | Curvature Radius 5.00 to 13.00 mm (0.01 mm increments) | YES
7 | Refractive Power 25.96 to 67.50 D (n = 1.3375) (0.01 / 0.12 / 0.25 D increments) | YES
8 | Cylindrical Power 0D to 12D , 0.01 /0.12/ 0.25 increments | YES
9 | Axis 0° to 180°, 1°/ 5° increments | YES
10 | Auto / Manual Measuring Mode | YES
11 | Corneal measurement range 10.0 to 14.0 mm (0.1 mm increments) | YES
12 | Screen Display Alignment | YES
13 | Included Accessories
14 | Motorized Table | YES
15 | printer rolls | YES
16 | dust cover | YES
17 | spare fuses | YES
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | YES
34 | The vendor is responsible to provide a minimum of three (_3_) years full warranty commencing from the date of Technical acceptance | YES
35 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime | YES
36 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
37 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
38 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
39 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
40 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
41 | Voltage Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted). | YES
42 | Frequency: 60Hz. | YES
43 | Power plug should be British style medical grade. | YES
44 | The vendor to provide: a) ☒ Installation. b) ☒ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☒ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. | YES
45 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
46 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
47 | a) If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. ☐ b) If the factory-level service training is not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled up to one year. If training still not provided at the end of the warranty extension, then the retention will be deducted and closed. 19. ☒ The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
48 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 32
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
32 | 442183007 | 0 | EXAMINATION CHAIR MOTORIZED RECLINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Complete Examination Unit with Patient Chair, Digital Slit Lamp & Imaging System | YES
2 | Type Diagnostic | YES
3 | FDA, CE, ISO approved Yes | YES
4 | Input Ratings 220VAC, 60Hz | YES
5 | Accessories as required | YES
6 | 2 Instrument plate | YES
7 | 2 Charging wheels | YES
8 | Control Panel | YES
9 | Head light control | YES
10 | Room light control | YES
11 | Main ON/OFF Switcgh | YES
12 | Control Switch for peripheral | YES
13 | Arms and Plates to OPTHALMIC CHAIR | YES
14 | Seat Height | YES
15 | Tower Type 5.00 to 13.00 mm (0.01 mm increments) | YES
16 | Backrest Reclining Angle 90° | YES
17 | Headrest, lock lever | YES
18 | Armrest type, swing up | YES
19 | Built in Switch for Up and Down Control | YES
20 | Upholstery: Vinyl | YES
21 | Load Capacity, Kg. | as required
22 | SLIT LAMP TYPE HS | YES
23 | Illumination System, as required | YES
24 | Slit Width, mm ≥ 0 -10, adjustable | Yes
25 | Slit Length, mm ≥ 0 -10, adjustable | YES
26 | Diaphragm Sizes, mm 8mm | YES
27 | Rotation ≥ 0 - 90 | YES
28 | Slit Tilt as required | YES
29 | Bulb Type LED Technology | YES
30 | Brightness Adjustable | YES
31 | Intensity, Lux | SPECIFY
32 | Filters Heat Absorption | YES
33 | Microscope Type: Binocular | YES
34 | Objectives (Focal length) mm as required | SPECIFY
35 | Eyepiece Magnification 6X, 10X, 16X, 25X, 40X or as required | YES
36 | Total Magnification 5 | YES
37 | Field of View, mm | SPECIFY
38 | Pupillary Distance, mm | SPECIFY
39 | Diopter Adjustment | as required
40 | Range of Movement AS FOLLOWS:
41 | Side Shift, mm | SPECIFY
42 | Depth Shift, mm | SPECIFY
43 | Height Shift, mm | SPECIFY
44 | Fine Adjustment by Joystick, mm | SPECIFY
45 | Chinrest Height Adjustment, mm | SPECIFY
46 | Chinrest AS FOLLOWS:
47 | Height Adjustable | YES
48 | Fixation light | YES
49 | Digital Camera Yes, 3MP | YES
50 | TANOMETER of HS Swivel with accessories | YES
51 | LENSES 20D, 30D, 60D, 78D, 90D & 3 Mirror Lenses | YES
52 | PRISME SER | YES
53 | PHOROPTER | YES
54 | TRIAL LENS SETWITH TRIAL FRAMES | YES
55 | COLOR TEST | YES
56 | ACCESSORIES | as required
58 | SYNCHRONIZED
59 | Included Accessories
60 | Accessoriey 1
61 | Accessoriey 2
62 | Accessoriey 3
63 | Accessoriey 4
64 | Accessoriey 5
65 | Accessoriey 6
66 | Accessoriey 7
67 | Accessoriey 8
68 | Optional Accessories
69 | Accessoriey 1
70 | Accessoriey 2
71 | Accessoriey 3
72 | Accessoriey 4
73 | Accessoriey 5
74 | Accessoriey 6
75 | Accessoriey 7
76 | Accessoriey 8
77 | Other Specification
78 | Vendor will do upgrade during Warranty FOC | YES
79 | Pre-Installation & Commissioning | Included
80 | Vendor will do upgrade during Warranty FOC | YES
81 | Pre-Installation & Commissioning | YES
82 | The vendor is responsible to provide a minimum of three (_3_) years full warranty commencing from the date of Technical acceptance | YES
83 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime | YES
84 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
85 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
86 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
87 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
88 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
89 | Voltage Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted). | YES
90 | Frequency: 60Hz. | YES
91 | Power plug should be British style medical grade. | YES
92 | The vendor to provide: a) ☒ Installation. b) ☒ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☒ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. | YES
93 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
94 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
95 | a) If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. ☐ b) If the factory-level service training is not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled up to one year. If training still not provided at the end of the warranty extension, then the retention will be deducted and closed. 19. ☒ The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
96 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 33
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
33 | 442183005 | 0 | ADVANCED SLIT LAMP SOPHISTICATED BIO MICROSCOPY AND LED POWER, INCLINED ADAPTER AND STEREO VARIATOR WITH COMPLETE SYSTEM ACCESSORIES SET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CLINICAL SPECIFICATIONS :
2 | ADVANCED SLIT LAMP SOPHISTICATED BIO MICROSCOPY AND LED POWER, INCLINED ADAPTER AND STEREO VARIATOR WITH COMPLETE SYSTEM ACCESSORIES SET
3 | · SLIT LAMP
4 | · Type, HS
5 | · Illumination System
6 | · Slit Width, mm : > 0 – 10, adjustable
7 | · Slit Length, mm : > 0 – 10, adjustable
8 | · Diaphragm Sizes, mm: 8mm
9 | · Rotation: > 0 – 90
10 | · Bulb Type: LED technology,
11 | · Brightness: Adjustable
12 | · Intensity, lux
13 | · Filters : Heat absorption
14 | · Microscope
15 | · Type of Microscope: Binocular
16 | · Objectives (Focal length) mm
17 | · Eyepiece Magnification : 6X, 10X, 16X, 25X, 40X
18 | · Total Magnification : 5
19 | · Field of View, mm
20 | · Method of Changing Magnification
21 | · Binocular Angle
22 | · Pupillary Distance, mm
23 | · Diopter Adjustment Range, Diopters
24 | Range Of Movement
25 | · Side Shift, mm: Yes
26 | · Depth Shift, mm: Yes
27 | · Height Shift, mm: Yes
28 | · Fine Adjustment by Joystick, mm: Yes
29 | · Chinrest Height Adjustment, mm: Yes
30 | · DIGITAL CAMERA: yes 3 MP
31 | · TONOMETER of HS: Swivel with accessories
32 | · PHOROPTER: YES, AUTO
33 | ACCESSORES (Itemized Price)
34 | · POWER SUPPLY: 220 Vac., 60 Hz.
35 | · OTHER SPECIFICATIONS : FDA, CE, ISO approved
36 | · the electric cable should be British standard
Company Stamp |  |  |  | Signature

## ITEM 34
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
34 | 414283905 | 0 | VISUAL ACUITY SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CLINICAL SPECIFICATIONS:
2 | Digital chart system for use in the refraction clinic with remote control with Contrast acuity and sinewave grating, contrast sensitivity testing
3 | IT SHOULD HAVE
4 | wall mountable, self-contained digital eye chart system for use in the refraction clinic without the need for a conventional projector
5 | Easy to use, infrared remote control to allow access of all tests from anywhere in the exam room.
6 | Large range of optotypes and tests Random sequencing
7 | Contrast acuity and sinewave grating, contrast sensitivity testing
8 | Adjustable Red/Green for compatibility with any refraction system
9 | Max Illumination: 220 cd/m2 or close equivalent Photopic Illumination 85 cd/m2 or close equivalent
10 | Mesopic Illumination 3 cd/m2 or close equivalent Refraction Range: 1.8 m to 9 m
11 | Electric cable should be British standard if any
Company Stamp |  |  |  | Signature

## ITEM 35
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
35 | 411541200 | 0 | DETECTOR AUTO-LENSMASTER LENS TYPE DTECTION UV MEASUREMENT 5.7-INCH COLOR LCD WITH TOUCH SCREEN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CLINICAL SPECIFICATIONS:
2 | Automatic Lensmeter With Built in Thermal Printer
3 | the electric cable should be British standard
4 | MEASUREMENT RANGE
5 | SPHERICAL POWER | specify
6 | Range, diopters | 0 to ± 25
7 | CYLINDER POWER | specify
8 | Range, diopters | 0 to ± 9.9 or more
9 | CYLINDER AXIS | specify
10 | Range, degrees | 0 to 180 (1° step increment)
11 | ADD POWER | Specify
12 | Range, diopters | 0 to ± 9.9 or more
13 | PRISMATIC POWER | specify
14 | Range, prism | 0 to ± 9.9 or more
15 | LENS DIAMETER, mm | 27 to 90 or wider
16 | BUILT IN THERMAL PRINTER | yes
17 | POWER REQUIREMENTS | 220 volts/60 Hz
18 | ACCESSORIES | 3 Rolls printer paper, Contact lens holder , Dust Cover, Spare fuses, 2 Spare lamps, Other accessories, (Itemized prize)
Company Stamp |  |  |  | Signature

## ITEM 36
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
36 | 414256510 | 0 | SCANNER OPHTHALMIC DIAGNOSTIC WITH INTEGRATED AXIAL LENGTH MEASURMENT |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | A/B Scan Ultrasound Diagnostic Machine with UBM Probe and Simulteneous A/B Scan Display | YES
2 | Application Diagnostic | YES
3 | Input Ratings 220VAC, 60Hz | YES
4 | FDA, CE certified Yes | YES
5 | Display Colored, LCD , at least 17in, 1024 x 480 pixels | YES
6 | System Control Unit 2 Button mouse, Dual foot pedal | YES
7 | Simulteanous Display of Images at least up to 4 | YES
8 | A and B-scan display Simultaneous | YES
9 | A-scan vector Selectable | YES
10 | Zoom at least x0.5 to x 1.5 | YES
11 | Pre and Post Processing Capabilities | YES
12 | Real time clock | YES
13 | Monitor Colored, LCD , at least 17in, 1024 x 480 pixels | YES
14 | Image Storage Computer Hard Disk | YES
15 | Image Storage Flush Memory, USB Capability at least up to 4 | YES
16 | Color Video Monitor output 1 RGB | YES
17 | Data outputs: 1 Firewire | YES
18 | Computer Hardware, Latest Technology | YES
19 | Printer Port 1 USB | YES
20 | Input ports 4 USB , 1 for keyboard & 1 for mouse | YES
21 | Transducer Frequency as follows:
22 | A-scan probe approximately 13 MHz | YES
23 | B-scan probe approximately 10 MHz | YES
24 | UBM probe approximately 50 MHz | YES
25 | B Scan Probe as follows:
26 | Sealed Probe | approximately 10 MHz
27 | Window Diamater | approximately 16 MHz
28 | Lateral Resolution | approximately 650 micron
29 | Axial Resolution | approximately 120 micron
30 | Biometry Resolution | +/- 0.3mm
31 | Scan Characteristics as follows:
32 | B-Scan Rate | approximately 10 or 20 MHz
33 | Sampling Resolution | approximately 5 micron
34 | Scan Angle | Selectable
35 | Scan Depth | approximately 0-60mm
36 | Receiver Characteristics as follows:
37 | Log, Linear, S-curve Amplify, Log Amp | 80 dB
38 | Selectable Dynamic Ranges | at least two
39 | Scan Conversion | 384 Scan lines x 512 samples / line 64 levvels of grayscale 256 levels (A-Scan) in vectore mode
40 | Included Accessories
41 | Accessoriey 1
42 | Accessoriey 2
43 | Accessoriey 3
44 | Accessoriey 4
45 | Accessoriey 5
46 | Accessoriey 6
47 | Accessoriey 7
48 | Accessoriey 8
49 | Optional Accessories
50 | Accessoriey 1
51 | Accessoriey 2
52 | Accessoriey 3
53 | Accessoriey 4
54 | Accessoriey 5
55 | Accessoriey 6
56 | Accessoriey 7
57 | Accessoriey 8
58 | Other Specification
59 | Vendor will do upgrade during Warranty FOC | YES
60 | Pre-Installation & Commissioning | Included
61 | Vendor will do upgrade during Warranty FOC | YES
62 | Pre-Installation & Commissioning | YES
63 | The vendor is responsible to provide a minimum of three (_3_) years full warranty commencing from the date of Technical acceptance | YES
64 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime | YES
65 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
66 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
67 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
68 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
69 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
70 | Voltage Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted). | YES
71 | Frequency: 60Hz. | YES
72 | Power plug should be British style medical grade. | YES
73 | The vendor to provide: a) ☒ Installation. b) ☒ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☒ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. | YES
74 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
75 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
76 | a) If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. ☐ b) If the factory-level service training is not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled up to one year. If training still not provided at the end of the warranty extension, then the retention will be deducted and closed. 19. ☒ The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
77 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 37
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
37 | 410442902 | 0 | TONOMETER NON-CONTACT WITH ADJUSTABLE LCD SCREEN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Air Puff Non-contact Tonometer for Ocular Pressure measurement | YES
2 | Air puff non-contact | YES
3 | FDA, CE certified | YES
4 | Display: Colored, LCD , at least 8.5 in, touchsreen | YES
5 | Carrying case Included | YES
6 | Pressure Measurements as follows:
7 | Ocular Pressure:1 to 60 mmHg (1 mmHg step) | YES
8 | Measureng range average indication: 1 mmHg / 0.1 mmHg steps changeable | YES
9 | Measuring mode: 1-30 mmHg or 1-60 mmHg | YES
10 | Working distance: at least 11cm | YES
11 | Included Accessories
12 | Accessoriey 1
13 | Accessoriey 2
14 | Accessoriey 3
15 | Accessoriey 4
16 | Accessoriey 5
17 | Accessoriey 6
18 | Accessoriey 7
19 | Accessoriey 8
20 | Optional Accessories
21 | Accessoriey 1
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Other Specification
30 | Vendor will do upgrade during Warranty FOC | YES
31 | Pre-Installation & Commissioning | YES
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | YES
34 | The vendor is responsible to provide a minimum of three (_3_) years full warranty commencing from the date of Technical acceptance | YES
35 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime | YES
36 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
37 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
38 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
39 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
40 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
41 | Voltage Dual (To be operated by 220V as well as 110V as applicable, otherwise 220V is accepted). | YES
42 | Frequency: 60Hz. | YES
43 | Power plug should be British style medical grade. | YES
44 | The vendor to provide: a) ☒ Installation. b) ☒ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☒ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. | YES
45 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
46 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
47 | a) If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. ☐ b) If the factory-level service training is not provided by the end of the warranty period, then the warranty will automatically be extended until satisfactorily fulfilled up to one year. If training still not provided at the end of the warranty extension, then the retention will be deducted and closed. 19. ☒ The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
48 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 38
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
38 | 411639800 | 0 | PHACOEMULSIFICATION MACHINE CENTURION VISION SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Phacoemulsification Machine with IOP Control Technology w/ Phaco and I/A Handpieces | YES
2 | FDA, CE certified | YES
3 | Display: Touchscreen, Colored Display | YES
4 | Patient Eye Level indicators | YES
5 | Incision Size: 2.2/2.4 | YES
6 | Fluidics Module: Active | YES
7 | Handpiece: Piezoelectric, lightweight | YES
8 | Irrigation Type: Active Fluidics with IOP Control and I.V. Pole | YES
9 | Construction: Single Piece, Titanium | YES
10 | Frequency Longtudinal: 43.5 KHz +/- 3.0 KHz | YES
11 | Frequency Torsional: 32 KHz +/- 2.0 KHz | YES
12 | Pulse Mode | YES
13 | Burst | YES
14 | Autoclavable | YES
15 | Irrigation Handpiece | YES
16 | Aspiration Handpiece | YES
17 | Pump Type: Peristaltic Pump | YES
18 | Tubing: Dual Segment Fluid Management System | YES
19 | Tubing Dimentsion: 1.21mm internal diameter | YES
20 | Maximum Vacuum: 0-700 mmHg max | YES
21 | Reflux Method | YES
22 | Vents to atmosphere | YES
23 | Bottle | YES
24 | I/A System | YES
25 | /A Handpiece | YES
26 | Port Diameter, mm: 0.3mm | YES
27 | Autoclavable | YES
28 | Collection Container Size: 250cc or more | YES
29 | Anterior Vitrectomy | YES
30 | Variable Speed: 1-5000 cpm | YES
31 | Disposable | YES
32 | Operational Modes As Follows:
33 | Irrigation | YES
34 | Ultrasound | YES
35 | I/A | YES
36 | Diathermy: Bipolar | YES
37 | Displays the Following:
38 | Preset/Actual Vacuum limits | YES
39 | Power/ Time | YES
40 | Cutting rate | YES
41 | Footswitch position | YES
42 | Included Accessories
43 | Accessoriey 1
44 | Accessoriey 2
45 | Accessoriey 3
46 | Accessoriey 4
47 | Accessoriey 5
48 | Accessoriey 6
49 | Accessoriey 7
50 | Accessoriey 8
51 | Optional Accessories
52 | Accessoriey 1
53 | Accessoriey 2
54 | Accessoriey 3
55 | Accessoriey 4
56 | Accessoriey 5
57 | Accessoriey 6
58 | Accessoriey 7
59 | Accessoriey 8
60 | Other Specification
61 | Vendor will do upgrade during Warranty FOC | YES
62 | Pre-Installation & Commissioning | Included
63 | Remote Control | YES
64 | Wireless Footswitch | YES
65 | Footswitch Hangar | YES
66 | Motorized IV Pole: Yes, 20 to 110cm | YES
67 | Adjustable tray | YES
68 | Storage Drawer | YES
69 | Lockable Casters | YES
Company Stamp |  |  |  | Signature

## ITEM 39
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
39 | 411639801 | 0 | VITRECTOMY MACHINE VISION SYSTEM LXT INCLUDES EMBEDDED LASER |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Vitrectomy Machine with emebedded Green Laser Coagulation for the retina | YES
2 | Input Ratings 220-230VAC, 60Hz | YES
3 | FDA, CE certified | YES
4 | Display at least 17in, Touchscreen, Colored Display | YES
5 | Patient Eye Level indicators | YES
6 | Fluidic Module As Follows:
7 | Dual Aspiration Mode: Venturi, up to 650 mmHg | YES
8 | Flow Control: for anterior and posterior procedures | YES
9 | Vacuum Control: 0-650 mmHg | YES
10 | Flow Control | YES
11 | IOP control | YES
12 | Auto infusion valve | YES
13 | Pneumatic Module As Follows:
14 | Beveled High Speed Victrectomy Probe: from 10,000 until 20,000 CPM | YES
15 | VFC Injection: 80psi | YES
16 | VFC Extrusion: up to 650 mmHg | YES
17 | Pneumatic Scissors / Forceps | YES
18 | Surgeon Controlled Duty Cycle: at least 3 different duty cycle | YES
19 | Illumination Source As Follows:
20 | Advanced Xenon Illuminator | YES
21 | Light Probes Types: Straight, Wide-Angle | YES
22 | Bulb Type and Lifetime: Xenon, 800hrs | YES
23 | Laser and Filters as follows:
24 | 532 nm Embedded Green Endolaser | YES
25 | Laser Probe Types: Articulating, Curved, Straight and Flexible | YES
26 | Laser Indirect Opthalmoscope | YES
27 | Ultrasound Module As Follows:
28 | Automated Ultrasound Injector | YES
29 | Diathermy Capability | YES
30 | Fragmentation Capability | YES
31 | Anterior Segment Features As Follows:
32 | Ultrasonic Technology | YES
33 | Phacoemulsification handpiece Type: Pizoelectric, Titanium, Lightweight | YES
34 | Frequency: Torsional , 32 KHz | YES
35 | Autoclavable | YES
36 | I/A handpiece: port diameter at least 3mm | YES
37 | Max Vacuum: up to 650 mmHg | YES
38 | Reflux: Linear and Fixed | YES
39 | Bottle Hanger | YES
40 | Drainage Bag: 500cc | YES
41 | Speed: 0-5000 cpm | YES
42 | Posterior Segment Features As Follows:
43 | Cutter Type: Dual, Pneumatic | YES
44 | Cut Rate Range: 0-20,000 cpm | YES
45 | Micro Incisional Vitreoretinal Surgery | YES
46 | Light Source: Xenon | YES
47 | Micro Incisional Vitreoretinal Surgery | YES
48 | Micro Incisional Vitreoretinal Surgery | YES
49 | Included Accessories
50 | Accessoriey 1
51 | Accessoriey 2
52 | Accessoriey 3
53 | Accessoriey 4
54 | Accessoriey 5
55 | Accessoriey 6
56 | Accessoriey 7
57 | Accessoriey 8
58 | Optional Accessories
59 | Accessoriey 1
60 | Accessoriey 2
61 | Accessoriey 3
62 | Accessoriey 4
63 | Accessoriey 5
64 | Accessoriey 6
65 | Accessoriey 7
66 | Accessoriey 8
67 | Other Specification
68 | Vendor will do upgrade during Warranty FOC | YES
69 | Pre-Installation & Commissioning | Included
70 | Procedure Metrics Report | YES
71 | Wireless remote control | YES
72 | Articulating Tray Arm | YES
73 | Multi Function Footswitch | YES
74 | Proportional Diathermy and Reflux | YES
75 | Auto Gas Fill | YES
76 | Pneumatic Scissors & Forceps | YES
77 | Barcode Scanner | YES
78 | The vendor is responsible to provide a minimum of three (_3_) years full warranty commencing from the date of Technical acceptance | YES
79 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
80 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
81 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
82 | a) ☒ The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). ☒ b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
83 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
84 | Power plug should be British style medical grade. | YES
85 | The vendor to provide: a) ☒ Installation. b) ☒ Supplier is responsible for travel and accommodation expenses (applicable for overseas purchase). c) ☒ Set up. d) ☒ Calibration. e) ☒ Performance verification. f) ☒ Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. g) ☒ Pre-installation & site preparations (mechanical & electrical, except civil), is the responsibility of the equipment’s vendor. The vendor has the right to inspect the site prior to quotation and he must coordinate with Hospital’s Clinical Engineering, Facility Projects Management and the main contractor for any work at site before starting. | YES
86 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
87 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
88 | If the vendor fails to provide user training, manuals, software, local service training and/or schedule for factory-level service training as stated above, then equipment acceptance will be put on hold until requirements provided. | YES
89 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
90 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
Company Stamp |  |  |  | Signature

## ITEM 40
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
40 | 414256509 | 0 | OPHTHALMIC MULTISPOT LASER SYSTEM FOR RETINAL DISORDERS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application:  Glaucoma: Cyclophotocoagulation. And laser trabeculoplasty a tissue-sparing laser therapy intended to reduce intraocular pressure  Retina: conventional photocoagulation. DME, CSR, and PDR | YES
2 | Type MicroPulse Laser System | YES
3 | Wavelength: 577 nm Yellow | YES
4 | Dimensions | Pls Specify
5 | Connector Type RFID ¦ Resistor | YES
6 | Cooling: Air/TEC cooled | YES
7 | Exposure Duration: CW-Pulse: 10 ms – 3000 ms or CW to 60 seconds | YES
8 | Exposure Interval: CW-Pulse: 10 ms – 3000 ms or single pulse | YES
9 | MicroPulse Duration: MicroPulse: 0.05–1.00 ms | YES
10 | MicroPulse Interval: MicroPulse: 1.00–10.00 ms | YES
11 | Aiming Laser: Diode laser, 635 nm nominal | YES
12 | Laser Energy Source: Frequency-doubled solid-state and direct diode | YES
13 | Maximum Power: 2000 mW | YES
14 | Exposure Duration CW-Pulse: 10–3000 ms | YES
15 | Exposure Interval: CW-Pulse: 10–3000 ms | YES
16 | MicroPulse Duration: MicroPulse: 0.05–1.00 ms
17 | MicroPulse Interval: MicroPulse: 1.00–10.00 ms | YES
18 | MicroPulse Duty Cycle Preset selections of 5%, 10%, and 15% | YES
19 | Aiming Beam: Diode laser, 635 nm nominal | YES
20 | Patterns: Grid (2x2 - 7x7), Circle, Triple Arc | YES
21 | Pattern Spacing Confluent (zero), 1-, 2-, 3-spot spacing in 0.25 diameter increments | YES
22 | User Interface Touch-screen & knobs | YES
23 | Spot Sizes: Single spot: 50 μm, 100 μm, 200 μm, 300 μm, 500 μm Multi-spot: 100 μm, 200 μm, 300 μm, 500 μm | YES
24 | Delivery Device Slit Lamp 0-2000mW | YES
25 | Included Accessories
26 | Accessoriey 1
27 | Accessoriey 2
28 | Accessoriey 3
29 | Accessoriey 4
30 | Accessoriey 5
31 | Accessoriey 6
32 | Accessoriey 7
33 | Accessoriey 8
34 | Optional Accessories
35 | Accessoriey 1
36 | Accessoriey 2
37 | Accessoriey 3
38 | Accessoriey 4
39 | Accessoriey 5
40 | Accessoriey 6
41 | Accessoriey 7
42 | Accessoriey 8
43 | Other Specification
44 | Vendor will do upgrade during Warranty FOC | YES
45 | Pre-Installation & Commissioning | Included
46 | • the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and • the equipment meets the standard of electromagnetic compatibility | YES
47 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, SFDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall. | YES
48 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
49 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
50 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
51 | Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released. | YES
52 | All prices must include delivery, installation, warranty and any related technical, specialized test equipment and calibrations, preventive maintenance kits, battery replacements and both technical certified service training to the clinical engineering and end users application training by qualified certified trainer to all related end users covering their shifts and would be repeated regularly during warranty period as refresher training sessions, as requested and agreed between both parties.. | YES
53 | Vendors shall submit price quotation for any upgrade\update that will add functional capabilities to the system during and after the warranty period. | YES
54 | Vendors must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operates the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities. | YES
55 | The vendor must supply complete warranty details for all quoted equipment and attached Accessories. | YES
56 | The vendor shall make suitable arrangements with the hospital to examine the room(s) accommodating the installation | YES
57 | DOCUMENTATION 1.Two (2) complete operational manuals for hardware and software. 2.Two (2) service technical manuals. 3.Two (2) spare part lists (if applicable). 4.Completed schematic diagrams with documentation of onsite modifications (if required). | YES
58 | All above documentation 2,3 and4) must be provided in English, and technically approved by Clinical Engineering, for the suitability to fully maintain the offered equipment and their components. | YES
59 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
60 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
61 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
62 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
63 | Service response time for routine calls should be one (1) hour. | YES
64 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
65 | The vendor should provide the manufacturers country of origin for each offered system\equipment. | YES
66 | The vendor must provide sufficient manufacture documentation to support the data quoted. | YES
Company Stamp |  |  |  | Signature

## ITEM 41
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
41 | 414256514 | 0 | OPHTHALMIC COMBINED LASER MACHINE ND:YAG PLUS ARGON-LASER TABLETOP |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | SFDA APPROVED | YES
2 | FDA CLEARANCE | YES
3 | TYPE OF LASER AS FOLLOWS:
4 | Argon | YES
5 | Dye | YES
6 | Krypton | YES
7 | Nd:YAG, frequency-doubled | YES
8 | Nd:YAG, Q-switched | YES
9 | Other (Other types of ophthalmic lasers include Er:YAG, Yb:YAG, and Nd:YLF.) | Pls Specify
10 | WAVELENGTH, nm 530-540 | YES
11 | INDICATION(S): Photoablation, Photocoagulation and Photodisruption | YES
12 | APPLICATION(S): Capsulotomy, Cataract treatment, Glaucoma treatment, Iridectomy Iridotomy, LASEK, LASIK, Macular degeneration treatment, PRK, PTK, Retinopathy treatment and Trabeculoplasty | YES
13 | Other (Other ophthalmic applications not listed above.) | Pls Specify
14 | Power Output at tissue, W (Maximum therapeutic energy delivered by the laser system. The greater the maximum output power, the greater the range of tissue effects and surgical applications.): Argon blue-green 3, argon green 1, dye 1, krypton green 1.5, krypton yellow 1.5, krypton red 1, Nd:YAG 1 | YES
15 | Power Output display type Surgeon control screen / surgeon assistant screen | YES
16 | TREATMENT MODES/BEAM Single, repeat | YES
17 | Continuous (duration [sec], repetition frequency [Hz]) | Pls Specify
18 | Pulsed (energy delivered [mJ], pulse duration/width [nsec], repeat time [sec], repetition frequency [Hz]) | Pls Specify
19 | Micropulse | YES
20 | Longpulse | YES
21 | Single pulse | YES
22 | Double pulse | YES
23 | Triple pulse | YES
24 | Super pulse | YES
25 | Pattern scanning | YES
26 | BEAM DIAMETER, mm As Follows:
27 | Multi-spot capable | YES
28 | Spot size, µm: 50-1,000 | YES
29 | Cone angle, deg (Focal angle produced by the exit optics for the beam.): YAG: 16, SLT: <3 | YES
30 | AIMING BEAM AS FOLLOWS:
31 | Type/color: Diode | YES
32 | Wavelength, nm 630 | YES
33 | DELIVERY SYSTEM TYPE AS FOLLOWS:
34 | Slit lamp | YES
35 | Intraocular probe | YES
36 | Handpiece(s) | YES
37 | Indirect ophthalmoscope | YES
38 | Other | Pls Specify
39 | COMPUTER SYSTEM Windows embedded | YES
40 | SAFETY FEATURES: Safety interlocks; key operation; emergency off switch, Optional laser warning light outside the laser therapy room | YES
41 | POWER REQUIREMENTS AS FOLLOWS:
42 | VAC: 110-240V | YES
43 | Hz: 50/60 | YES
44 | Power Plug: 3 Prong Medical Grade British Type | YES
45 | DISPLAYS/CONTROLS AS FOLLOWS:
46 | Beam actuator (The means by which the therapeutic beam is activated. The Nd:YAG beam is generally activated by a footswitch or a joystick push button.) | YES
47 | Selected energy | YES
48 | Shot selection | YES
49 | Power output | YES
50 | Delivered energy counter | YES
51 | Pulse counter | YES
52 | AUTO-CALIBRATION | YES
53 | COOLING METHOD AS FOLLOWS:
54 | Air: Preferred | YES
55 | Thermoelectric | YES
56 | LASER GAS CYLINDERS | Pls Specify
57 | H x W x D, cm (in) | Pls Specify
58 | WEIGHT, kg (lb) | Pls Specify
59 | Warranty: 3 Yrs, Prefered 5yrs | YES
60 | TRAINING: Certified Required Training for both End Users and Technical Service Training to Clinical Engineering, All deemed tools and Test Clibration Equipment, included | YES
61 | Hospital Information System Integration | YES
63 | Included Accessories
64 | Standard Accessories: Laser console, footpedal, owner's manual, power cord, Eye Safety, laser goggles, Surgeon Chair, chinrest paper, head belt, grips, focusing rod, attachment plates, arm rest, dust cover, DANGER labe…etc. | YES
65 | Optional Accessories: Swiveling patient bed, integrated surgical microscope, swiveling slit lamp, wireless footswitch | YES
66 | Accessoriey 3
67 | Accessoriey 4
68 | Accessoriey 5
69 | Accessoriey 6
70 | Accessoriey 7
71 | Accessoriey 8
72 | Optional Accessories
73 | Accessoriey 1
74 | Accessoriey 2
75 | Accessoriey 3
76 | Accessoriey 4
77 | Accessoriey 5
78 | Accessoriey 6
79 | Accessoriey 7
80 | Accessoriey 8
81 | Other Specification
82 | Vendor will do upgrade during Warranty FOC | YES
83 | Pre-Installation & Commissioning | Included
84 | • the equipment and its accessories is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building code (i.e.: NFPA or equivalent); and • the equipment meets the standard of electromagnetic compatibility | YES
85 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, SFDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equivalent on any recall. | YES
86 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
87 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
88 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
89 | Any components not proposed, but made available to the marketplace as new technology post proposal deadline, shall be priced and proposed as released. | YES
90 | All prices must include delivery, installation, warranty and any related technical, specialized test equipment and calibrations, preventive maintenance kits, battery replacements and both technical certified service training to the clinical engineering and end users application training by qualified certified trainer to all related end users covering their shifts and would be repeated regularly during warranty period as refresher training sessions, as requested and agreed between both parties.. | YES
91 | Vendors shall submit price quotation for any upgrade\update that will add functional capabilities to the system during and after the warranty period. | YES
92 | Vendors must provide at least three (3) clients (name of contact person, address, phone number and email) within the Middle East, Europe and/or North America that operates the same or similar equipment as the system profiled in the proposal. This listing may be used to assess satisfaction with equipment and service at these facilities. | YES
93 | The vendor must supply complete warranty details for all quoted equipment and attached Accessories. | YES
94 | The vendor shall make suitable arrangements with the hospital to examine the room(s) accommodating the installation | YES
95 | DOCUMENTATION 1.Two (2) complete operational manuals for hardware and software. 2.Two (2) service technical manuals. 3.Two (2) spare part lists (if applicable). 4.Completed schematic diagrams with documentation of onsite modifications (if required). | YES
96 | All above documentation 2,3 and4) must be provided in English, and technically approved by Clinical Engineering, for the suitability to fully maintain the offered equipment and their components. | YES
97 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
98 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
99 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
100 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
101 | Service response time for routine calls should be one (1) hour. | YES
102 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
103 | The vendor should provide the manufacturers country of origin for each offered system\equipment. | YES
104 | The vendor must provide sufficient manufacture documentation to support the data quoted. | YES
Company Stamp |  |  |  | Signature

## ITEM 42
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
42 | 414256508 | 0 | ULTRASOUND OPHTHALMIC A/B SCAN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | ULTRASOUND BIO-MICROSCOPY FOR EYE U.B.M, combinatioN, B-scan. UBM. A-Scan | YES
2 | B-SCAN As Follows:
3 | Unparalleled image quality with proprietary Enhanced Focus Rendering | YES
4 | 12 and 20 MHz frequency transducers available for posterior B-scan | YES
5 | 256 ray scan with 2048 sample points for each ray (over a half-million sample points per transducer sweep) | YES
6 | Capture 50-frame video clips at up to 20 frames per second | YES
7 | Store unlimited images and up to 12 video clips in each exam | YES
8 | A-SCAN As Follows:
9 | Immersion or direct contact A-scan | YES
10 | Manual or automatic capture (cataract, dense cataract, aphakic, silicone oil, and pseudophakic modes) | YES
11 | Auto calculation of axial length, anterior chamber depth, lens thickness, and vitreous, with axial length average and standard deviation provided for up to 10 scans per exam | YES
12 | Standard IOL formulas: Binkhorst, Regression-II, Theoretic/T, Holladay, Hoffer-Q, and Haigis Post-refractive IOL formulas include Latkany Myopic Regression, Latkany Hyperopic, and Aramberri Double-K | YES
13 | UBM As Follows:
14 | Unparalleled image quality with proprietary Enhanced Focus Rendering | YES
15 | 35 and 50 MHz frequency transducers available for UBM | YES
16 | 256 ray scan with 2048 sample points for each ray (over a half-million sample points per transducer sweep) | YES
17 | Capture 50-frame video clips at up to 20 frames per second | YES
18 | Store unlimited images and up to 12 video clips in each exam | YES
19 | UBM scan cap kit compatible with ClearScandevice for direct contact UBM scanning, Angle analysis Eye Tracking alignment feature | YES
20 | Transportable on the Special Cart | YES
21 | Ultrasound Screen Size : 21.0 Inch OR better | YES
22 | Image Resolution : 0.03 mm or better | YES
23 | Gain : 40 TO 130 | YES
24 | T.G.C : 0 to 60 db | YES
25 | UBM Probe frequency : 35MHz to 50 MHZ | YES
26 | OPEN/Closed Probe : , with Cap Protection | YES
27 | Measurment method : Contact Clearscan Or Window immersion | YES
28 | Transducer ultrasounf Movement : Sector Technology | YES
29 | Removable Protection Window for 50 MHz Probe | YES
30 | UltraSound Resolution,AXiaL : 35 micrometer Lateral : 60 micrometer | YES
31 | Advanced Software tools for Glaucoma Diagnosis : AOD , TIA , IT | YES
32 | Color Filter for post proccessing images,Noise Reduction , Average Filter, Vitreuos enhance | YES
33 | Possibility adding Multiple Probes in Future ( Option ), 25 MHZ , 20 MHZ, 10 MHZ, Biometery :10MHz, 12MHZ, 20MHz | YES
34 | STS , WTW , ATA Measurments | YES
35 | Focal Distance: 10to 13 mm | YES
36 | Depth of Exploration : 21 mm | YES
37 | Tissue Penetration : 5 to 6 mm | YES
38 | Image Freez: Footpedal Or Mouse | YES
39 | Video recording & Images Store & Patient Report , Database | YES
40 | Desktop Computer connected to System | YES
41 | Full Cerner Integration HL7 | YES
42 | Probe Holder: Functional , 180 Degree Orientation | YES
43 | Calibration Tools Included | YES
44 | Printer | YES
45 | Connection to medical server: DICOM and EMR | YES
46 | SFDA, CE, FDA | YES
47 | Power Requirements, 230V with British Hospital Grade Plugs | YES
48 | Included Accessories
49 | Accessoriey 1
50 | Accessoriey 2
51 | Accessoriey 3
52 | Accessoriey 4
53 | Accessoriey 5
54 | Accessoriey 6
55 | Accessoriey 7
56 | Accessoriey 8
57 | Optional Accessories
58 | Accessoriey 1
59 | Accessoriey 2
60 | Accessoriey 3
61 | Accessoriey 4
62 | Accessoriey 5
63 | Accessoriey 6
64 | Accessoriey 7
65 | Accessoriey 8
66 | Other Specification
67 | Vendor will do upgrade during Warranty FOC | YES
68 | Pre-Installation & Commissioning | YES
69 | All equipment (and powered accessories) delivered to the must use 220VAC/60 Hz power and power cable with British hospital medical grade outlet 3-prong plug. | YES
70 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
71 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
72 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
73 | The minimum warranty period provided by the vendor that is acceptable to this hospital shall be three (3) years, and higher preferences will be given to any other offered longer warranty periods | YES
74 | Two (2) complete operational manuals for hardware and software. | YES
75 | Two (2) service technical manuals. | YES
76 | Two (2) spare part lists | YES
77 | Completed schematic diagrams | YES
78 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
79 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
80 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
81 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
82 | Service response time for routine calls should be one (1) hour. | YES
83 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
84 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
85 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
Company Stamp |  |  |  | Signature

## ITEM 43
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
43 | 411548111 | 0 | OPHTHALMIC CHAIR PROCEDURE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | 4 Movements with independent engine as follows: | YES
2 | Up and Down | YES
3 | Backrest | YES
4 | Legrest | YES
5 | Trendelemburg | YES
6 | Full adjustable Headrest | YES
7 | Rechargeable Battery System | YES
8 | 6” Caster with Multiple Break System | YES
9 | Min. & Max: 21’’- 30’’ | YES
10 | Retractable and Lockable Side rails | YES
11 | Remote Control and Safety Key | YES
12 | Footswitch | YES
13 | Full Length with Headrest: 6 FT | YES
14 | 2 Programable Position (Sitting and laying) | YES
15 | Standby System | YES
16 | Upholstery Width:26’’ | YES
17 | Total Width: 30’’ | YES
18 | Headrest Stroke: 8’’ | YES
19 | Weight Support: 400 lbs | YES
20 | Included Accessories
21 | Accessoriey 1
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Optional Accessories
30 | Accessoriey 1
31 | Accessoriey 2
32 | Accessoriey 3
33 | Accessoriey 4
34 | Accessoriey 5
35 | Accessoriey 6
36 | Accessoriey 7
37 | Accessoriey 8
38 | Other Specification
39 | Vendor will do upgrade during Warranty FOC | YES
40 | Pre-Installation & Commissioning | Included
41 | All equipment (and powered accessories) delivered to the must use 220VAC/60 Hz power and power cable with British hospital medical grade outlet 3-prong plug. | YES
42 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
43 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
44 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
45 | The minimum warranty period provided by the vendor that is acceptable to this hospital shall be three (3) years, and higher preferences will be given to any other offered longer warranty periods | YES
46 | Two (2) complete operational manuals for hardware and software. | YES
47 | Two (2) service technical manuals. | YES
48 | Two (2) spare part lists | YES
49 | Completed schematic diagrams | YES
50 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
51 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
52 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
53 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
54 | Service response time for routine calls should be one (1) hour. | YES
55 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
56 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
57 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
Company Stamp |  |  |  | Signature

## ITEM 44
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
44 | 414256511 | 0 | OPTICAL COHERENCE TOMOGRAPHY A-SCAN WITH FUNDUS CAMERA |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | OCT-ANTERIOR & POSTERIOR with Angiography & Fundus Camera | YES
2 | Expandable diagnostic imaging platform which combines scanning laser fundus imaging with high-resolution OCT. | YES
3 | Eye Tracking technology to recognize structures in the cSLO image in real time and uses the information to actively control the position of the SD-OCT scan and only saves scans that were taken at the correct location of the retina. | YES
4 | Automatically repeats scans at the exact same location during follow-up exams (Auto Rescan function). | YES
5 | Eye Tracking and Auto Rescan for the precise and reliable measurement of nerve fiber layer thickness for the early detection and monitoring of glaucoma. | YES
6 | Automated segmentation of the GCL to enable the identification of retinal layers that are best suited for differentiating glaucomatous from healthy eyes. | YES
7 | Infrared Reflectance Imaging: Using long wavelengths of laser light, infrared imaging to capture fundus images in fine detail even without pupil dilation. | YES
8 | Deeper penetration of light, combined with the confocal principle, to provide more distinct details of intra retinal lesions such as CME or CSCR. | YES
9 | Lower brightness for elderly patients and children who may experience discomfort with flash photography. | YES
10 | Images to be acquired with a minimal pupil size of 1.5 mm. | YES
11 | Spectral-domain OCT with 85 kHz scanning rate to provide high spatial resolution cross sections of the retina and the anterior segment. | YES
12 | Smallest change in retinal thickness measurable should be 1 micron | YES
13 | Application: For anterior & posterior segment analysis | YES
14 | Spectral Domain OCTA with Fundus Camera Technology | YES
15 | Light Source: SLED, 830nm | YES
16 | Scanning Speed :80,000 Ascans/Sec | YES
17 | Axial Resolution5µm | YES
18 | Transverse Resolution: 12µm | YES
19 | Bandwidth 50nm half bandwidth | YES
20 | Overall Scan Depth 2.4mm | YES
21 | Minimum Pupil Size 3mm | YES
22 | Focus Adjustment .-25D to +25D | YES
23 | Scan Range Posterior 5mm to 12mm | YES
24 | Scan Range Anterior 3mm to 16mm | YES
25 | Scan Range Angiography 3mm to 9mm | YES
26 | Scan Type Posterior 3D, Radial, B-Scan, Raster, Cross, Angio | YES
27 | Scan Type Anterior 3D, Radial, Cross, B-Scan | YES
28 | IOP Correction | YES
29 | Dye Free Retinal Angiography with Mosaic View | YES
30 | Operator Voice Guided | YES
31 | Full automatic, semi & manual alignment & Scanning mode | YES
32 | Fundus Image Non Mydratic Funud Camera | YES
33 | Angle of View 45°-50° | YES
34 | Bulitin Camera 12megapixel CCD Camera | YES
35 | Retinal thickness analysis, Inner Retinal Thickness, outer retinal Thickness, RNFL+GCL+IPL thickness, GCL+IPL Thickness, RNFL Thickness, RPE Deformation, IS/OS Thickness | YES
36 | Angio view of Superficial Plexus, Deep Plexus, Outer Retina, Choriocapilaries, Depth Coded, Custom, Enface, Thickness Map | YES
37 | Angiography Acquistion methods Auto & Manual Mode upto 12 images | YES
38 | Angiography predefined Modes 7x7mm, 10x6mm, 10x10mm, 12x5mm | YES
39 | Angiography correlation images upto 12 images | YES
40 | Glaucoma Analysis, RNFL, ONH  morphology, DDLS, OU & Hemisphere asymmtery, Gangalion analysis as RNFL+GCL+IP and GCL+IPL | YES
41 | Progression Morphology Upto 6 Exams | YES
42 | Progression Quantification Upto 6 Exams | YES
43 | Wide field Scannning 12mm | YES
44 | Peripheral Scanning | YES
45 | Disk Scanning | YES
46 | Comparision & Follow Up | YES
47 | Anterior Analysis, Pachymetery, lasik Flap assesment,AIOP, Angle Assessment, AOD 500/750, TISA 500/750 | YES
48 | Anterior Wide Scan Angle to Angle measurement | YES
49 | Network & Storege compatibility DICOM Storage SCU, Networking | YES
50 | Fixation Target OLED Internal & External | YES
51 | Minimum System requirements "Core i7, 3.0GHz, 4 Physical Cores,  HDD 1TB, 16GB RAM, Color Printer | YES
52 | Image Export option | YES
53 | 3D Mode View | YES
54 | Network Printer | YES
55 | Power Requirement: 240V-50-60 Hz with British Hospital Grade Plug | YES
56 | Optional Accessories
57 | Accessoriey 1
58 | Accessoriey 2
59 | Accessoriey 3
60 | Accessoriey 4
61 | Accessoriey 5
62 | Accessoriey 6
63 | Accessoriey 7
64 | Accessoriey 8
65 | Other Specification
66 | Vendor will do upgrade during Warranty FOC | YES
67 | Pre-Installation & Commissioning | Included
68 | All equipment (and powered accessories) delivered to the must use 220VAC/60 Hz power and power cable with British hospital medical grade outlet 3-prong plug. | YES
69 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
70 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
71 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
72 | The minimum warranty period provided by the vendor that is acceptable to this hospital shall be three (3) years, and higher preferences will be given to any other offered longer warranty periods | YES
73 | Two (2) complete operational manuals for hardware and software. | YES
74 | Two (2) service technical manuals. | YES
75 | Two (2) spare part lists | YES
76 | Completed schematic diagrams | YES
77 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
78 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
79 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
80 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
81 | Service response time for routine calls should be one (1) hour. | YES
82 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
83 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
84 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
Company Stamp |  |  |  | Signature

## ITEM 45
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
45 | 410442903 | 0 | TONOMETER BATTERY OPERATED MEASURES INTRAOCULAR PRESSURE IOP PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Air Puff Non-contact Tonometer for Ocular Pressure measurement | YES
2 | FDA, CE certified | YES
3 | Display, Colored, LCD , at least 8.5 in, touchsreen | YES
4 | Pressure Measurements, Ocular Pressure: 1 to 60 mmHg (1 mmHg step) | YES
5 | Pressure Measurements, Measureng range average indication: 1 mmHg / 0.1 mmHg steps changeable | YES
6 | Pressure Measurements, Measuring mode: 1-30 mmHg or 1-60 mmHg | YES
7 | Working distance: at least 11cm | YES
8 | Power supply: Rechargeable internal lithium-ion battery | YES
9 | Position of the person during the measurement, Standing, sitting or lying down in the supine position | YES
10 | Data storage | YES
11 | Display Type, organic light-emitting diode (OLED) display | YES
12 | Included Accessories
13 | Carrying case | YES
14 | Other Needed Consumables valid for 3 months operation | YES
15 | Optional Accessories
16 | Accessoriey 1
17 | Accessoriey 2
18 | Accessoriey 3
19 | Accessoriey 4
20 | Accessoriey 5
21 | Accessoriey 6
22 | Accessoriey 7
23 | Accessoriey 8
24 | Other Specification
25 | Vendor will do upgrade during Warranty FOC | YES
26 | Pre-Installation & Commissioning | Included
27 | Warranty 3 Years | Included
28 | Operation and Field Service Manuals Included | Included
Company Stamp |  |  |  | Signature

## ITEM 46
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
46 | 410442702 | 0 | OPHTHALMIC CLASSIC SLIT LAMP PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CLINICAL SPECIFICATIONS :
2 | PORTABLE SLIT LAMP
3 | Versatility and portability are the key benefits of the Classic Portable Slit Lamp. This handheld slit lamp is designed and manufactured with the strength to withstand the daily traumas of the demanding ophthalmic environment, being lightweight, durable, and robust.
4 | Advanced Optics
5 | Top-quality multi-layered glass is crafted into multi-element and multi-layer coated optical system to give maximum performance with guaranteed clarity. 10x and 16x magnification.
6 | Controllable Illumination
7 | With a proven halogen illumination system, illumination levels are controlled by a rheostst, giving you control of your examination parameters. From zero to maximum, you choose.
8 | Big Slit Lamp Features, Portable Usability Unique features such as the fixation targets, and the 1mm square light patch for assessing anterior chamber flare set the new standard for portable slit lamps.
9 | Features:
10 | · 10x and 16x magnification
11 | · LED illumination
12 | · Filters: neutral density, red-free, diffuser, and cobalt blue
13 | · Slit Wheel - 1x1, 0.15, 0.5, 0.8, 1.6, 12
14 | · Blue and white
15 | · Charging base included and direct charging capability
16 | the electric cable should be British standard if any
Company Stamp |  |  |  | Signature

## ITEM 47
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
47 | 411548112 | 0 | SEAT BALANCE ADVANCE BRAKE WITH HEIGHT ADJUSTMENT VIA FOOT ACTIVATION AND A BRAKE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | CLINICAL SPECIFICATIONS :
2 | SURGEON STOOL
3 | High sitting comfort for your surgical practice The specialist in OPHTAHLMOLOGY . Achieve the perfect sitting position - with the three- dimensionally adjustable armrests and the patented seat cushion - for a healthy back. Stay sterile and simply adjust the seat height using the foot release.
4 | · Three-dimensionally adjustable armrests
5 | · Firm fixing of the armrest position free of play
6 | · Flexible height adjustment from 590 to 790 mm
7 | · Foot release to adjust the seat height and actuate the central brake
8 | · Adjustable back section with lordosis support
9 | · Patented seat section with adjustable sitting angle
10 | · Comfort undercarriage with antistatic castors
11 | · High quality finish with top quality material
12 | · Outstanding design
13 | TECHNICAL DATA:
14 | Five ball bearing castors made of polyurethane, one of them conductive
15 | · Height adjustable Ø 65 m
16 | · Backrest angle 480 - 680 mm
17 | · Backrest padding angle +6° / -16°
18 | · Adjustable backrest height 25°
19 | · Seat section angle 100 mm
20 | · Seat surface dimensions (W x H x D) :+5° / -8°
21 | · Maximum load :460 x 60 x 420 mm 120 kg
22 | the electric cable should be British standard if any
Company Stamp |  |  |  | Signature

## ITEM 48
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
48 | 411548113 | 0 | PATIENT OPHTHALMOLOGY SURGICAL OPERATING TABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Application: Ophthalmic surgery | YES
2 | FDA CLEARANCE | YES
3 | TABLE POSITIONS AS FOLLOWS:
4 | Degrees from horizontal
5 | Trendelenburg: ≥25 | YES
6 | Reverse Trendelenburg: ≥25 | YES
7 | Lateral tilt: ≥18 | YES
8 | Vertical range, cm (in): About 58.1-108 (22.8-42.5) | YES
9 | TABLE SECTIONS: ≥4 | YES
10 | Degrees from horizontal
11 | Back section: +55 to -25 minimum | YES
12 | Foot/leg section: +20 to -40 minimum | YES
13 | Head section: +45 to -90 minimum | YES
14 | ACCESSORY SIDERAIL | YES
15 | MAXIMUM PATIENT CAPACITY, kg (lb)
16 | Static: ≥270 (600); 450 (1,000) for bariatric patients | YES
17 | Articulated: 300 (660) | YES
18 | CONTROLS: Hand or foot | YES
19 | CONTROLS Type: IR remote, corded hand, corded foot, override panel | YES
20 | Manual override | YES
21 | ACCESSORIES: Arm boards, shoulder braces, stirrups, leg holders, counter supports and Complete line of table accessories for Ophthlmic Surgery | YES
22 | Radiolucent tabletop | YES
23 | Prefered C-arm accessible | YES
24 | BASE ATTACHMENT TYPE: Mobile | YES
25 | Column housing: Stainless steel | YES
26 | Conductive casters | YES
27 | Caster lock | YES
28 | Prefered Sliding top, cm (in) | YES
29 | Interchangeable top: Head and foot | YES
30 | POWER REQUIREMENTS: 220-240V/50/60 Hz | YES
31 | Battery: Rechargeable | YES
32 | Electric cable should be British standard HG Plug | YES
33 | Included Accessories
34 | Ergonomic head rest for the patient
35 | patient arm rest with belt | Qty : 2
36 | Mountining brackets: Yes
37 | Microsurgery ring
38 | Hourse shoes head rest with telescopic head section
39 | Accessoriey 6
40 | Accessoriey 7
41 | Accessoriey 8
42 | Optional Accessories
43 | Other Specification
44 | Vendor will do upgrade during Warranty FOC | YES
45 | Pre-Installation & Commissioning | YES
46 | All equipment (and powered accessories) delivered to the must use 220VAC/60 Hz power and power cable with British hospital medical grade outlet 3-prong plug. | YES
47 | All equipment must conform connectivity with ORACLE/CERNER hospital information system with all needed hardware and software requirements and it license. | YES
48 | The equipment proposed shall incorporate any “latest generation” hardware and software Technology enhancements available from the manufacturer in the marketplace at the time of equipment installation. | YES
49 | The vendor will guarantee an official notification to the hospital of any equipment Updates or model changes occurring immediately before the processing of this order, or the delivery of the equipment. The Vendor shall notify the hospital accordingly and grant the hospital the option to Purchase the upgrade or revised models. | YES
50 | The minimum warranty period provided by the vendor that is acceptable to this hospital shall be three (3) years, and higher preferences will be given to any other offered longer warranty periods | YES
51 | Two (2) complete operational manuals for hardware and software. | YES
52 | Two (2) service technical manuals. | YES
53 | Two (2) spare part lists | YES
54 | Completed schematic diagrams | YES
55 | All specialized, dedicated needed test equipment, dummy loads, functional verifications and calibrations, should be provided and included at no extra charges, to be handed over to the concerned clinical engineering, during the delivery of the purchased equipment or immediately before the end of warranty period. | YES
56 | Provide listing of qualified service field engineers who will be routinely dispatched from the local office to service the quoted equipment. Indicate their years of experience and whether they have, or will be factory trained on the specified equipment. | YES
57 | The vendor shall guarantee to provide to this hospital with technical support and the supply of spare parts for a period of five (5) years after the end of production. The vendor shall also inform the hospital the end of production dates when they are announced. | YES
58 | Equipment downtime shall not exceed 2% of the total operation time per annum; Based on 52 weeks per year X 44 hours per week. If the downtime exceeds the above two percent (2%), The contract warranty periods will be extended to cover 3 TIMES the actual downtime. | YES
59 | Service response time for routine calls should be one (1) hour. | YES
60 | Service engineer must provide a service report each time they evaluate and/or work on the equipment. | YES
61 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
62 | 17. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
Company Stamp |  |  |  | Signature

## ITEM 49
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
49 | 411689102 | 0 | ANALYZER WITH THE CAPACITY FOR ION-SELECTIVE ELECTRODE (ISE) DETERMINATION OF SODIUM POTASSIUM, AND CHLORIDE IN SERUM, PLASMA, AND URINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Analyzer is capable of performing clinical chemistry, special protein and therapeutic drugs mentoring. Analyzer with capacity for ION-SELECTIVE ELECTRODE (ISE) determination of Sodium, Potassium and Chloride in serum, plasma and urine samples. | YES
2 | Configuration: Fully automated Analyzer | YES
3 | Processing modes: Random access, stat | YES
4 | Sample Type: ION-SELECTIVE ELECTRODE (ISE) DETERMINATION OF SODIUM POTASSIUM, AND CHLORIDE IN SERUM, PLASMA, AND URINE | YES
5 | Liquid-level sensing | YES
6 | Auto dilution | YES
7 | Clot detection | YES
8 | Abnormal values flag : Visual & audible | YES
9 | Auto Calibration | YES
10 | DISPLAY TYPE : High resolution touchscreen color | YES
11 | LIS | YES
12 | LINE POWER, VAC 120-240 V | YES
13 | Frequency 50/60 Hz | YES
14 | UPS | YES
15 | Printer | YES
16 | Barcode Compatibility | YES
17 | Included Accessories
18 | Accessoriey 1
19 | Accessoriey 2
20 | Accessoriey 3
21 | Accessoriey 4
22 | Accessoriey 5
23 | Accessoriey 6
24 | Accessoriey 7
25 | Accessoriey 8
26 | Optional Accessories
27 | Accessoriey 1
28 | Accessoriey 2
29 | Accessoriey 3
30 | Accessoriey 4
31 | Accessoriey 5
32 | Accessoriey 6
33 | Accessoriey 7
34 | Accessoriey 8
35 | Other Specification
36 | Vendor will do upgrade during Warranty FOC | YES
37 | Pre-Installation & Commissioning | YES
38 | Vendor will do upgrade during Warranty FOC | YES
39 | Pre-Installation & Commissioning | YES
40 | The vendor is responsible to provide a minimum of 3 year full warranty commencing from the date of first case. | YES
41 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
42 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
43 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
44 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
45 | The contractor will provide all the essential tools and test equipment required to maintain the equipment free of charge. | YES
46 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
47 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
48 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
49 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
50 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
51 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. | YES
52 | The Frequency must be 60Hz. | YES
53 | Power plug should be British style medical grade. | YES
54 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
55 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 50
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
50 | 411689103 | 0 | ANALYZER WITH ELECTROCHEMILUMINESCENCE (ECL) TECHNOLOGY FOR IMMUNOASSAY ANALYSIS |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | analyzer that uses ElectroChemiLuminescence (ECL) technology for immunoassay analysis. Capable of performing both quantitative and qualitative in vitro assay determinations for a broad range of applications (including anemia; bone, cardiac and tumor markers; critical care; fertility/hormones; maternal care; and infectious diseases) | YES
2 | Configuration: Fully Automated Analyzer | YES
3 | Processing modes: Random access, stat | YES
4 | Method used: ElectroChemiLuminescence (ECL) | YES
5 | Liquid-level sensing | YES
6 | Auto dilution | YES
7 | Clot detection | YES
8 | Abnormal values flag | YES
9 | Autowash | YES
10 | Auto calibration | YES
11 | LIS | YES
12 | LINE POWER, VAC 120-220 V | YES
13 | Frequency: 50/60 Hz | YES
14 | UPS | YES
15 | Printer | YES
16 | Barcode Compatibility | YES
17 | Included Accessories
18 | Accessoriey 1
19 | Accessoriey 2
20 | Accessoriey 3
21 | Accessoriey 4
22 | Accessoriey 5
23 | Accessoriey 6
24 | Accessoriey 7
25 | Accessoriey 8
26 | Optional Accessories
27 | Accessoriey 1
28 | Accessoriey 2
29 | Accessoriey 3
30 | Accessoriey 4
31 | Accessoriey 5
32 | Accessoriey 6
33 | Accessoriey 7
34 | Accessoriey 8
35 | Other Specification
36 | Vendor will do upgrade during Warranty FOC | YES
37 | Pre-Installation & Commissioning | YES
38 | The vendor is responsible to provide a minimum of 3 year full warranty commencing from the date of first case. | YES
39 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
40 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
41 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
42 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
43 | The contractor will provide all the essential tools and test equipment required to maintain the equipment free of charge. | YES
44 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
45 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
46 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
47 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
48 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
49 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. | YES
50 | The Frequency must be 60Hz. | YES
51 | Power plug should be British style medical grade. | YES
52 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
53 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 51
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
51 | 411689104 | 0 | ANALYZER EVOLUTION COAGULATION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Automated benchtop analyzer for coagulation laboratory tests. Capable of performing STAT testing with no impact on routine laboratory test time. | YES
2 | Configuration: Fully Automated benchtop analyzer | YES
3 | Method used: Clotting, immunologic, chromogenic | YES
4 | Sample type: Plasma | YES
5 | Reagents refrigerated onboard | YES
6 | LIS | YES
7 | LINE POWER, VAC :120-220 V | YES
8 | Frequency: 50/60 Hz | YES
9 | Barcode Compatibility | YES
10 | UPS | YES
11 | Printer | YES
12 | Included Accessories
13 | Accessoriey 1
14 | Accessoriey 2
15 | Accessoriey 3
16 | Accessoriey 4
17 | Accessoriey 5
18 | Accessoriey 6
19 | Accessoriey 7
20 | Accessoriey 8
21 | Optional Accessories
22 | Accessoriey 1
23 | Accessoriey 2
24 | Accessoriey 3
25 | Accessoriey 4
26 | Accessoriey 5
27 | Accessoriey 6
28 | Accessoriey 7
29 | Accessoriey 8
30 | Other Specification
31 | Vendor will do upgrade during Warranty FOC | YES
32 | Pre-Installation & Commissioning | YES
33 | The vendor is responsible to provide a minimum of 3 year full warranty commencing from the date of first case. | YES
34 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
35 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
36 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
37 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
38 | The contractor will provide all the essential tools and test equipment required to maintain the equipment free of charge. | YES
39 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
40 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
41 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
42 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
43 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
44 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. | YES
45 | The Frequency must be 60Hz. | YES
46 | Power plug should be British style medical grade. | YES
47 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
48 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 52
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
52 | 411689105 | 0 | ANALYZER DIFFERENTIAL FULLY INTEGRATED IPU INCLUDING LCD COLOUR TOUCHSCREEN |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Automated hematology analyzer for performing the differential and CBC Tests. Capable of performing STAT testing with no impact on routine laboratory test time. | YES
2 | Fully Automated benchtop analyzer | YES
3 | Method used: Basic hematology | YES
4 | Auto dilution | YES
5 | Autosampler | YES
6 | Automated Calibration | YES
7 | Slide staining | YES
8 | LIS | YES
9 | LINE POWER, VAC :120-220 V | YES
10 | Frequency: 50/60 Hz | YES
11 | UPS | YES
12 | Barcode Compatibility | YES
13 | Printer | YES
14 | Included Accessories :
15 | Optional Accessories
16 | Accessoriey 1
17 | Accessoriey 2
18 | Accessoriey 3
19 | Accessoriey 4
20 | Accessoriey 5
21 | Accessoriey 6
22 | Accessoriey 7
23 | Accessoriey 8
24 | Other Specification
25 | Vendor will do upgrade during Warranty FOC | YES
26 | Pre-Installation & Commissioning | YES
27 | The vendor is responsible to provide a minimum of 3 year full warranty commencing from the date of first case. | YES
28 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
29 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
30 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
31 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
32 | The contractor will provide all the essential tools and test equipment required to maintain the equipment free of charge. | YES
33 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
34 | If applicable, original software license to be provided, and a back-up copy of the software for software driver equipment. In addition, any software or passwords required for maintenance and troubleshooting are to be provided free of charge. | YES
35 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
36 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
37 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
38 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. | YES
39 | The Frequency must be 60Hz. | YES
40 | Power plug should be British style medical grade. | YES
41 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
42 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
Company Stamp |  |  |  | Signature

## ITEM 53
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
53 | 438793100 | 0 | FLOW HOOD LAMINAR |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | 1- Configuration(free standing), 4 feet size. | YES
2 | 2- To be used for IV admixture preparation and drug compounding | YES
3 | 3-Horizontal flow laminar flow clean bench. | YES
4 | 4-Ensuring at least ISO class 5 (Class 100) clean air flow | YES
5 | 5- Mag guage | YES
6 | 6- Adjustable leg risers | YES
7 | 7- Additional duplex outlet | YES
8 | 8- Plastic storage pin. | YES
9 | 9- Stainless steel IV bar | YES
10 | 10- Frame material | Stainless steel.
11 | 11- Instrument size. | 4 FEET
12 | 12- UV light. | YES
13 | 13-. The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
14 | 14-The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
15 | 15-Voltage: 220V Frequency: 50/60Hz | YES
16 | Included Accessories
17 | Accessoriey 1
18 | Accessoriey 2
19 | Accessoriey 3
20 | Accessoriey 4
21 | Accessoriey 5
22 | Accessoriey 6
23 | Accessoriey 7
24 | Accessoriey 8
25 | Optional Accessories
26 | Accessoriey 1
27 | Accessoriey 2
28 | Accessoriey 3
29 | Accessoriey 4
30 | Accessoriey 5
31 | Accessoriey 6
32 | Accessoriey 7
33 | Accessoriey 8
34 | Other Specification
35 | 1-Vendor will do upgrade during Warranty FOC | YES
36 | 2-Pre-Installation & Commissioning | Included
37 | 3 – Pharmacy diffuser | Included
38 | 4- Drip pan. | Included
39 | 5- Adjustable speed. | Included
Company Stamp |  |  |  | Signature

## ITEM 54
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
54 | 414255002 | 0 | TABLET COUNTER MACHINE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA, SFDA and CE Approved | YES
2 | Configuration(Bench top) | YES
3 | Automatic pharmaceutical tablet/pill counter for use in filling outpatient pharmaceutical prescriptions, with the following features. | YES
4 | Convenient operation with no calibration or adjustments needed for different products. | YES
5 | Counting rate / throughput: > 15 tablets / second | YES
6 | Accuracy: 99.9% or better. | YES
7 | Tablet size: from 3 mm diameter pills up to 20 mm long capsules. | YES
8 | Approximate tray capacity > 500 cc. | YES
9 | Digital control panel with digital readout display of count. | YES
10 | Have an inventory mode, with last count register. | YES
11 | Ergonomic, easy to use and simple to clean. | YES
12 | Compact design | YES
13 | Included Accessories
14 | Accessoriey 1
15 | Accessoriey 2
16 | Accessoriey 3
17 | Accessoriey 4
18 | Accessoriey 5
19 | Accessoriey 6
20 | Accessoriey 7
21 | Accessoriey 8
22 | Optional Accessories
23 | Accessoriey 1
24 | Accessoriey 2
25 | Accessoriey 3
26 | Accessoriey 4
27 | Accessoriey 5
28 | Accessoriey 6
29 | Accessoriey 7
30 | Accessoriey 8
31 | Other Specification
32 | Vendor will do upgrade during Warranty FOC | YES
33 | Pre-Installation & Commissioning | Included
34 | The companies are fully responsible to evaluate the site drawings, available services and dimensions before submitting their quotations and should clearly verify in their quotation the full adaptability of their quoted items to the site where it will be installed. Such verifications will include the guarantee of the transportation of the delivered items up to the destination site without any barriers due to door openings or hallways dimensions. Also the quantities, weight and dimensions of the delivered items will fit or installed into the designated areas as per standards and safety regulations. | YES
35 | During Warranty period, the supplier is required to visit at each consignee’s site at least once in 6 months commencing from the date of the installation for preventive maintenance of the goods supplied. | YES
36 | The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
37 | The manufacturer will confirm that the equipment met the standard of electromagnetic compatibility. | YES
38 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
39 | 17. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
40 | Voltage: 220V Frequency: 50/60Hz, With Hospital Grade British Plugs | YES
Company Stamp |  |  |  | Signature

## ITEM 55
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
55 | 441827101 | 0 | MACHINE FLUIDOSE SERIES 5 MPI UNIT DOSE PACKAGING SYSTEM |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA/CE Approved | YES
2 | Configuration (free standing/Bench top). | YES
3 | Unit-Dose Packaging System with built-in pump mechanism: Syringe / peristaltic pump | YES
4 | Automated, barcoding packaging solution for unit-dose oral liquid medication. | YES
5 | Capacity Dose Per Minute:>15 dose per minute | YES
6 | Automated cup dispensing | YES
7 | Should have FDA-grade disposable tubing sets. | YES
8 | Class A packaging | YES
9 | Support Different Packages Cup Sizes | YES
10 | Color coding of controlled or refrigerated medications | YES
11 | Custom colors availability Option | YES
12 | Dispensing oral liquid medication, Includes UD Barcode Labeling Software. | YES
13 | Bar code verifier | YES
14 | Prefered Touchscreen controls | YES
15 | Fully-customizable pump settings | YES
16 | Settings Parameters: Volume, seal temperature, and limits, draw and dispense rates. | YES
17 | Automatic low fluid detection | YES
18 | Disposable FDA-grade tubing | YES
19 | Safety Features, pauses packaging when fluid level reaches preset limit. | YES
20 | Included Accessories
21 | Accessoriey 1
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Optional Accessories
30 | Accessoriey 1
31 | Accessoriey 2
32 | Accessoriey 3
33 | Accessoriey 4
34 | Accessoriey 5
35 | Accessoriey 6
36 | Accessoriey 7
37 | Accessoriey 8
38 | Other Specification
39 | Vendor will do upgrade during Warranty FOC | YES
40 | Pre-Installation & Commissioning | Included
41 | The companies are fully responsible to evaluate the site drawings, available services and dimensions before submitting their quotations and should clearly verify in their quotation the full adaptability of their quoted items to the site where it will be installed. Such verifications will include the guarantee of the transportation of the delivered items up to the destination site without any barriers due to door openings or hallways dimensions. Also the quantities, weight and dimensions of the delivered items will fit or installed into the designated areas as per standards and safety regulations. | YES
42 | During Warranty period, the supplier is required to visit at each consignee’s site at least once in 6 months commencing from the date of the installation for preventive maintenance of the goods supplied. | YES
43 | The manufacturer will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). | YES
44 | The manufacturer will confirm that the equipment met the standard of electromagnetic compatibility. | YES
45 | The manufacturer will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the hospital should be informed officially by the manufacturer throughout life expectancy of the equipment on any recall. | YES
46 | 17. The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
47 | Voltage: 220V Frequency: 50/60Hz, With Hospital Grade British Plugs | YES
Company Stamp |  |  |  | Signature

## ITEM 56
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
56 | 447190100 | 0 | SUCTION UNIT PORTABLE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Safety Standards: EN60601, IEC 60601 | YES
2 | Device Classification: Class II Equipment, Type BF | YES
3 | Vacuum, MAX: 350 mmHg or more | YES
4 | Vacuum Range: 80 – 350 mmHg | YES
5 | Battery Run Time: Minimum 45 mins | YES
6 | Battery Life : Two Years minimum | YES
7 | Battery Type: Not SLA, Preferred Lithium Ion | YES
8 | Weight: Not More than 4Kg | YES
9 | Noise Level: Less than 60dBA | YES
10 | Pressure Guage: Analog or Digital | YES
11 | Audible alarm: Low Battery, Occlusion Pressure | YES
12 | Power requirements: 100-240V; 50/60Hz | YES
13 | Air Flow : 25LPM or more | YES
14 | Canister Capacity: 800ml to 1200ml | YES
15 | IPX7 | YES
16 | bacteria filters | YES
17 | collection canisters | YES
18 | aspirator tubing kits | YES
19 | Accessoriey 4
20 | Accessoriey 5
21 | Accessoriey 6
22 | Accessoriey 7
23 | Accessoriey 8
24 | Optional Accessories
25 | Accessoriey 1
26 | Accessoriey 2
27 | Accessoriey 3
28 | Accessoriey 4
29 | Accessoriey 5
30 | Accessoriey 6
31 | Accessoriey 7
32 | Accessoriey 8
33 | Other Specification
34 | Vendor will do upgrade during Warranty FOC | YES
35 | Pre-Installation & Commissioning | Included
36 | The vendor is responsible to provide a minimum of (Three) year full warranty commencing from the date of first case. | Included
37 | Equipment downtime during warranty shall not exceed three percent (3%) per annum, excluding time for preventive maintenance, calculated based on 44 hours per week X 52 weeks per year. If exceeded, then the warranty will be extended by twice (2X) the actual downtime. | YES
38 | Response time during warranty should not exceed 24 hours, for working days, or 48 hours, for official holidays and weekends, from when notification of failure or problem is received by vendor. | YES
39 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
40 | The vendor to provide: Installation Set up Calibration Performance verification Preventive maintenance during warranty period including all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
41 | The vendor should shoulder the periodic maintenance with pm kits and parts cost during warranty period. | YES
42 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this tender. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
43 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical equipment interface with appropriate building codes (i.e. NFPA or equivalent). | YES
44 | The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
45 | The manufacturer/vendor will confirm that the equipment and its accessories are not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall. | YES
46 | Equipment Must be auto Voltage (To be operated by 220 V as well as 120 V) otherwise 220 V is accepted. | YES
47 | The Frequency must be 60Hz. | YES
48 | Power plug should be British style medical grade | YES
49 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
50 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies | YES
Company Stamp |  |  |  | Signature

## ITEM 57
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
57 | 417098109 | 0 | WARMER UNIT BLOOD FLUID |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | FDA CLEARANCE | YES
2 | HEAT EXCHANGER TECHNOLOGY: Dry heat | YES
3 | Max temperature setting, °C (°F): 43 (109.4) | YES
4 | Increments, °C (°F): 0.1 | YES
5 | Warm-up time, min: <2 | YES
6 | MAXIMUM FLOW, mL/min, 10°C input to ≥37°C output: ≥250 | YES
7 | PRIMING VOLUME, Ml:20-150 | YES
8 | Display Temperature range, °C (°F):33-41 (91.4-105.8) | YES
9 | Display Temperature Increments, °C (°F): 0.5 | YES
10 | HIGH-TEMPERATURE CUTOFF, °C (°F): 43 (109.4) | YES
11 | Alarms: Overtemperature, undertemperature, and independent backup safety alarms | YES
12 | Alarms Audible/Visual | YES
13 | Alarm test | YES
14 | AIR VENT OR TRAP: Manual bubble trap on standard flow sets, autoventing bubble trap on all high-flow sets | YES
15 | IV-pole attachment: Preferred | YES
16 | MAINTENANCE REQUIREMENTS: Free Maintenance | YES
17 | POWER REQUIREMENTS: 220-240 50/60 Hz | YES
18 | Warranty: MIN 5 Yrs Unlimited | YES
19 | Open Unlimited Warranty and Free Model Upgrade against Consumable Deal | YES
20 | Included Accessories
21 | DISPOSABLE PARTS: Fluid warming sets: pediatric (0-100 mL/min), neonate (0-100 mL/min), standard (0-150 mL/min), standard with air aspiration port (0-150 mL/min), standard with extension (0-150 mL/min), high flow (0-500 mL/min), hybrid high flow (0-500 mL/min), hybrid high flow with patient line (0-500 mL/min) | YES
22 | Accessoriey 2
23 | Accessoriey 3
24 | Accessoriey 4
25 | Accessoriey 5
26 | Accessoriey 6
27 | Accessoriey 7
28 | Accessoriey 8
29 | Optional Accessories
30 | Accessoriey 1
31 | Accessoriey 2
32 | Accessoriey 3
33 | Accessoriey 4
34 | Accessoriey 5
35 | Accessoriey 6
36 | Accessoriey 7
37 | Accessoriey 8
38 | Other Specification
39 | Vendor will do upgrade during Warranty FOC | YES
40 | Pre-Installation & Commissioning | Included
41 | Rapid infusor; built-in volumetric pump; high-efficiency warmer; line pressure monitor; screen displays true fluid temperature, line pressure, flow rate, total volume infused; precise flow rates from 2.5 to 1,000 mL/min; automatic air removal; redundant air detectors; stop pump and clamp line at any alarm; reaches temperature within 20 sec of priming | Prefered
42 | The vendor will guarantee an official notification to the Hospital of any equipment updates or model changes occurring immediately before the processing of this order. Said notification shall grant the Hospital the option to purchase the upgrade or revised model. However, any model changes, hardware or software upgrades occurring during the warranty period shall be provided to the Hospital free of charge. | YES
43 | The manufacturer/vendor will confirm that the equipment is approved by UL544 or IEC 601/1 or equivalent for medical and dental equipment & UL1262 for laboratory equipment or equivalent and will interface with appropriate building codes (i.e. NFPA or equivalent). b) The manufacturer/vendor will confirm that the equipment met the standard of electromagnetic compatibility. | YES
44 | The manufacturer/vendor will confirm that the equipment and its accessories are approved and not presently listed on any recall action or alerts by an authorized testing or regulating agency, such as SFDA, ECRI, FDA, BSI or equivalent. In addition, the Hospital should be informed officially by the manufacturer/vendor throughout life expectancy of the equipment on any recall and/or alert. | YES
45 | The vendor to provide: a) Installation. b) Set up. c) Performance verification. d) Preventive maintenance (during warranty period) to include all the necessary labor, parts, PM kits, and supporting service reports free of charge. | YES
46 | One (1) complete operations manual, to be provided preferably in both hard and soft copies. | YES
47 | One (1) complete service manual (in English language) including schematics and spare parts list with pricing fixed for three (3) years. The manual is to be provided preferably in both hard and soft copies. | YES
48 | The supplier shall guarantee to provide the Hospital with the technical support & the supply of the spare parts for a period of five years after the date of end of production. The supplier shall inform the Hospital of the end of production dates when announced. | YES
49 | The Equipment must be registered /certified by Saudi Food & Drug Authority (SFDA/MDMA). | YES
50 | One temperature dongle for PM free of charge against every 5 provided warmer units | YES
Company Stamp |  |  |  | Signature

## ITEM 58
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
58 | 411661003 | 0 | MEDICATION FRIDGE 5.5 CUBIC FEE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | High-Performance Undercounter Lab Refrigerators, Solid Door
2 | COLOR | SILVER OR WIGHT
3 | Operating Temperature Rang | 3°C to 7°C
4 | Capacity (English) | 5.5 cu. ft.
5 | Capacity (Metric) | 156 L
6 | Dimensions | (L x W x H) Exterior26 x 23.6 x 31.8 in.(66 x 60 x 81 cm), (D x W x H) Interior20.5 x 19.5 x 23 in. (52 x 49.5 x 58.4 cm)
7 | Weight (Metric) | 53.98 kg
8 | Shipping Weight (Metric) | 71 kg
9 | Display Full color | TFT-LCD; 4.3”; 480 x 272
10 | Door Style | Single Solid
11 | Shelves | 3
12 | Defrost | Automatic
13 | Numbers of Doors | 1
Company Stamp |  |  |  | Signature

## ITEM 59
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
59 | 417843000 | 0 | MEDICATION FRIDGE LARGE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | high-performance pharmacy refrigerators feature adjustable basket drawers and glass doors for the easy identification and access of vaccines, pharmaceuticals, and chemotherapeutic material, SHOULD BE quiet, sustainable choice for medical and pharmacy-grade refrigeration. SHOULD HAVE V-drive technology is designed to detect usage patterns such as door openings when a higher compressor speed is needed and periods of stability where the compressor runs at a lower speed, saving energy without compromising protection
2 | COLOR | SILVER OR WIGHT
3 | Operating Temperature Rang | 2° to 8°C, factory setpoint at 5°C Convenient
4 | Capacity (English) | 11.5 cu. ft.
5 | Capacity (Metric) | 326 L
6 | Dimensions | Height (Metric) Exterior 185.4 cm, Width (Metric) Exterior 61.9 cm.
7 | Door Style | Glass
8 | Drawers | sliding drawer baskets for simplified inventory and access Easy to use microprocessor control with interior light, setpoint security and alarms SHOULD BE Compatible with wireless monitoring systems and remote alarm contacts.
9 | Door locks | Standard
Company Stamp |  |  |  | Signature

## ITEM 60
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
60 | 93196 | 0 | MEDICATION FRIDGE 11.5 CUBIC FEET |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Operating Temperature Rang | +2°C to +8°C
2 | Capacity (English) | 11.5 cu. ft.
3 | Dimensions | 73"H x 24"W x 30.5"D
4 | Weight (Metric) | 363 lbs
5 | Shipping Weight (Metric) | 388 lbs
6 | Electrical | 208-230 V / 60 Hz Need a Different Voltage? See 115V TSX1205PA
7 | Door Style | Single Glass
8 | Interior | 52.4"H x 20"W x 21.8"D
9 | Defrost | Automatic
10 | Certifications | UL, cUL, RoHS, Energy Star
11 | Accessories | Locking Casters, 5 Pull-Out Basket Drawers, Key Lock
Company Stamp |  |  |  | Signature

## ITEM 61
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
61 | 411661007 | 0 | REFRIGERATOR SMALL MEDICAL-GRADE |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | Type | REFRIGERATOR SMALL MEDICAL-GRADE WITH PYXIS LOCK ADAPTER KIT, Access Control provides data capture of unit access
2 | LED lighting | Energy-efficient LED lighting illuminates interior and includes on/off switch.
3 | Casters | 2" (50mm)
4 | Shelves | Full size
5 | Drawers | Locking drawers
Company Stamp |  |  |  | Signature

## ITEM 62
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
62 | 432359100 | 0 | WHEELCHAIR MULTIFUNCTION |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Wheelchair for patient transfer use in hospital setting. lightweight frame, weighing around 36 lb. a semi-pneumatic caster to accommodate a weight capacity of 300 lb. Durable, low-maintenance, triple chrome-plated carbon steel frame is long-lasting Dual axle positions allow variation of seat-tofloor heights 14-gauge cross braces add strength and durability Urethane rear tires mounted on “no flex” wheels offer Sectional back should be 16" to 18"heights Heavy-duty inner liners keep seat and back
2 | Overall Dimensions | Folded width Overall depth w/o riggings 34" Overall height Weight w/o riggings 36 lb. Weight capacity 250 lb.
3 | Chair Dimensions | Seat width 14” Seat depth 16” Back height 16” Seat-to-floor height 17.5”
4 | Casters Dimensions | Caster height 6"
5 | Wheel Dimensions | Wheel height 20" to 24"
6 | ADDITIONAL FEATURES | IV pole and oxygen tank holder. Seat belts. Educational logo to fasten seat belt when use the wheelchair. Hospital logo
7 | Color | Black vinyl upholstery /lather material
Company Stamp |  |  |  | Signature

## ITEM 63
ORIGINAL QUOTATION
SN | CODE | GROUP NUMBER | ITEM DESCRIPTION |  | MOH MODEL OFFERED: ID
63 | 410470200 | 0 | WHEELCHAIR BARIATRIC |  | 0
Quantity Quoted | SUPPLIER NAME |  | Unit Price including the VAT | MANUFACTURER | COUNTRY OF ORIGIN
0 | 0 |  | 0 | 0 | 0
MAN. CATALOG # | UNIT PRICE | UNIT PRICE IN WRITING |  | Total Price (SR) including Vat for quoted quantity
0 | 0 | 0 |  | 0
delivery 1st Shipment Quantity ( not less than 50 % of offered QTY) within maximum 30 days of PO date |  | 0 | delivery 2nd Shipment (remaining quantity ) within maximun 60 days of PO date |  | 0
COMPANY COMMENTS/ REMARKS: |  | 0
TECHNICAL & PERFORMANCE SPECIFICATION FOR MEDICAL EQUIPMENT
# | Technical Parameters | Specified | Yes/No | Catalogue/Brochure PAGE NUMBER where specification is mentioned | Supplier's Confirmation/ Remarks
1 | TYPE | Wheelchair for Bariatric patient transfer use in hospital setting lightweight frame, weighing less than 58 lb. makes this chair perfect for rental and long- or Short-term use. This wheelchair offers multiple seat widths, arm styles, sectional back and seat to- floor heights by utilizing the dual axle position Frame base. The 18" to 26" widths of frame base is the optional A semi-pneumatic caster to accommodate a weight capacity of 250 to 550 lb. Durable, low-maintenance, triple chrome-plated carbon steel frame is long-lasting Dual axle positions allow variation of seat-tofloor heights cross braces add strength and durability. Urethane rear tires mounted on “no flex” wheels offer Heavy-duty inner liners keep seat and back from stretching
2 | Overall Dimensions | Folded width Overall depth w/o riggings 38" Overall height Weight w/o riggings 58 lb. Weight capacity 550 lb.
3 | Chair Dimensions | Seat width 24” Seat depth 20” Back height 18” Seat-to-floor height 20”
4 | Casters Dimensions | height 8"
5 | Wheel Dimensions | height 24"
6 | ADDITIONAL FEATURES | IV pole and oxygen tank holder. Seat belts. Educational logo to fasten seat belt when use the wheelchair. Hospital logo
7 | Color | Black vinyl upholstery /lather material
Company Stamp |  |  |  | Signature